## AnnexD_NewFormatTerms.csv_Final
Annex Name: Annex D
Annex Title: Cause Investigation – Investigation Conclusion
Annex Version: 3.0
Annex Description: Terms/codes for describing the conclusion of the device involved in the reported event.
Annex Instructions: Terms/Codes in this annex should be used to answer the question What is the conclusion regarding the root cause? (based on the results noted in the Investigation Findings and the way they were obtained noted in the Type of Investigation.)
Annex Approval Date: 21 January 2021
Level 1 Term | Level 2 Term | Code | Definition | Non-IMDRF Code/Term | Status | Status Description | CodeHierarchy
Cause Traced to Device Design |  | D01 | Problems traced to the design specifications (e.g. in the requirements, testing processes, hazard analysis, implementation strategy). |  |  |  | D01
| Design Inadequate for Purpose | D0101 | Problems traced to design/design features of the device that do not support or interfere with the intended purpose of the device. |  |  |  | D01¦D0101
| Human Factors Engineering - Device Difficult to Operate | D0102 | Problems traced to inappropriate and/or inadequate assessment and engineering design of the device to accommodate how or where the device will be used. |  |  |  | D01¦D0102
| Human Factors Engineering - Device Difficult to Assemble | D0103 | Problems traced to inadequate design of the component parts and/or assembly steps resulting in the device not being able to be assembled correctly. |  |  |  | D01¦D0103
| Human Factors Engineering - Device Difficult to Reprocess | D0104 | Problems traced to inadequate design of the reprocessing steps and/or the device resulting in the device remaining unclean. |  |  |  | D01¦D0104
| Missing or Inadequate Safety Measures | D0105 | Problems traced to inadequate design or complete lack of safety measures leading to device malfunction or unintended properties of the device including possible hazards for persons using the device. |  |  |  | D01¦D0105
| Design Change Validation Inadequate | D0106 | Problems traced to inadequate or lack of validation of design changes of the device leading to malfunction or unintended properties of the device including possible hazards for persons using the device. |  |  |  | D01¦D0106
| Inadequate Cybersecurity | D0107 | Problems traced to inadequate protection of computers, servers, mobile devices, electronic systems, networks, programs, and data from malicious attacks, damage, or unauthorized access. |  | New | New term/definition added on 21 January 2021. for details, refer to comment No. 26 of the Change Log. | D01¦D0107
| Inadequate Software Design. | D0108 | Cause traced to inadequate or inappropriate design of the software component of a medical device, or Software as a Medical Device. |  | New | New term/definition added on 21 January 2021. for details, refer to comment No. 27 of the Change Log. | D01¦D0108
Cause Traced to Component Failure |  | D02 | Expected or random component failure without any design or manufacturing issue. |  |  |  | D02
Cause Traced to Manufacturing |  | D03 | A defect in the processes or systems used in the manufacture of the device. Examples include problems within the change control, production, or quality control processes. |  |  |  | D03
| Manufacturing Deficiency | D0301 | Problems traced to manufacturing process. |  |  |  | D03¦D0301
| Quality Control Deficiency | D0302 | Problems traced to the failure to maintain or establish techniques for controlling and verifying the product specifications (including materials used) identified by the manufacturer himself. |  |  |  | D03¦D0302
Cause Traced to Transport/Storage |  | D04 | Problems traced to the inappropriate transport or storage of the device. |  |  |  | D04
Cause Traced to Infrastructure |  | D05 | Problems traced to underlying framework, systems, and processes at the healthcare facility or other point of use (e.g. as building power supply, network, oxygen systems). |  |  |  | D05
Cause Traced to Environment |  | D06 | Problems caused by exposure to environmental conditions outside the expected range. |  |  |  | D06
Cause Traced to Maintenance |  | D07 | Problems traced to improper routine or preventative maintenance. |  |  |  | D07
Cause Traced to Training |  | D08 | Problems caused by inadequate training. |  |  |  | D08
Cause Traced to Labeling |  | D09 | Problems that occur as the result of problems with the labeling (including package inserts, instruction manuals, instructions for use). |  |  |  | D09
Cause Traced to Non-Device Related Factors |  | D10 | The adverse event that occurred is confirmed to have been caused by something other than the device or usage. |  | Modified (technical) | The term and definition were modified on 21 January 2021. (Old term/definition: Cause Cannot be Traced to Device/The adverse event that occurred is not attributable to a device.). For details, refer to comment No. 28 of the Change Log. | D10
| Adverse Event Related to Patient Condition | D1001 | An existing condition or disease is demonstrably responsible for the adverse event and use of the device has neither caused nor otherwise influenced this condition/disease-related adverse event. |  |  |  | D10¦D1001
| Adverse Event Related to Procedure | D1002 | The adverse event occurred during the procedure and the device had no influence on event. |  |  |  | D10¦D1002
| Adverse Event Related to Commutability | D1003 | The adverse event occurred because the material being used to calibrate or assess IVD performance did not have similar properties to those of human samples, leading to inappropriate bias and erroneous results. Examples of materials include: reference materials, calibrators, proficiency testing samples. |  |  |  | D10¦D1003
Cause Traced to User |  | D11 | The adverse event caused partially or wholly by the user of the device including sample handling. |  |  |  | D11
| Failure To Follow Instructions | D1101 | Problems traced to the user not following the manufacturer's instructions. |  |  |  | D11¦D1101
| Unintended Use Error Caused or Contributed to Event | D1102 | The interaction between the user and device, or sample, caused or contributed to the error. This includes unintended inappropriate use of the device and incorrect sample preparation. |  |  |  | D11¦D1102
| Cause Traced to Intentional Off-Label, Unapproved, or Contraindicated Use | D1103 | Problems traced to the intentional use of the device in an unapproved procedure, for an unapproved patient, or for which it is contraindicated, or not listed on the label. |  |  |  | D11¦D1103
| Shelf Life/Expiration Date Exceeded | D1104 | A device that has exceeded the period of time/date recommended by the manufacturer for storing the device without a degradation in quality. |  |  |  | D11¦D1104
| End of Life Problem Identified | D1105 | Problems traced to the device reaching the end of its useful life. |  |  |  | D11¦D1105
| Reuse of Single Use Device | D1106 | Problems traced to the use of the device more than once when it is designed for only one use. |  |  |  | D11¦D1106
| Sabotage or Intentional Misuse | D1107 | Problems occurred because of willful damage or reckless misuse. This term indicates that there was an apparent intent to cause damage or injury. |  | Modified (editorial) | Editorial change to the definition on 16 February 2021. | D11¦D1107
Known Inherent Risk of Device |  | D12 | Reported adverse event known and documented in the labeling and all reasonable mitigation steps have been taken (including both short or long term known complications or adverse reactions). |  | Modified (technical) | The definition was modified on 21 January 2021 (addition of phrase "and all reasonable mitigation steps have been taken."). For details, refer to comment No. 31 of the Change Log. | D12
Falsified Device |  | D13 | Devices that deliberately and/or fraudulently misrepresent their identity, composition or source. |  |  |  | D13
No Problem Detected |  | D14 | The device complaint or problem cannot be confirmed. |  |  |  | D14
Cause Not Established |  | D15 | The investigation findings do not lead to a clear conclusion about the cause of the reported adverse event. |  |  |  | D15
| Cause Linked to Device but Unable to Trace More Specifically | D1501 | Problems traced to the device, but the investigation could not identify the actual root cause (e.g. Design, Manufacturing, Component). |  | New | New term/definition added on 21 January 2021. for details, refer to comment No. 30 of the Change Log. | D15¦D1501
Conclusion Not Yet Available |  | D16 | A conclusion has yet to be established as the investigation is incomplete. Do not use this code if the investigation is complete. |  |  |  | D16
Appropriate Term/Code Not Available |  | D17 | The concluded cause is not adequately described by any other term. Note: This code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. |  |  |  | D17
Cause Traced to Software Coding |  | D18 | Problems traced to an error, flaw or fault in a computer program or system that causes it to produce an incorrect or unexpected result, or to behave in unintended ways. |  | New | New term/definition added on 21 January 2021. for details, refer to comment No. 27 of the Change Log. | D18