## B. Change Request Log
Appendix B: Change Log
| Requester information |  | Change Proposal Information | [10 empty] | IMDRF Decision
|  |  | Identification of code / term for which proposal is made | [5 empty] | Proposal of change
| Date submitted (DD/MM/YYYY) | Submitter (organisation name) | Terminology (Annex A, B, C, D, E, F, G) | Version of Annex | Code | Term | Location in the hierarchy | Definition | Category of change (Select: Add, delete, modify) | Description of change (e.g. "modification of definition …") | Rationale for change (e.g. "the change is necessary to accommodate a new type of device…") | Impact on other existing terms | Example of an incident which would be coded using the proposed term | Outcome of change request (ACCEPTED or REJECTED) | Justification | New code if applicable | New Term and/or Definition
1 | 2025-08-29 | IMDRF AET WG | B | 2025 | B13 | Communications/interviews | Level 1 | The investigation involved communication/interviews (either interpersonal or through technical means, e.g. phone, e-mail) with persons close to the adverse event, e.g. healthcare professionals (doctors, nurses etc.), the affected patient(s) or other users including, where appropriate, relatives or others engaged in caring for the affected patient. | Delete | Delete - use B15 instead |  | Retire B15 |  | ACCEPT | The description of B13 overlaps with B15. Communication/interviews can be captured by the 'Analysis of Information Provided by User/Third Party'.
2 | 2025-08-29 | IMDRF AET WG | B | 2025 | B12 | Trend Analysis | Level 1 | The investigation involved trend analysis of adverse event of the actual device involved in the adverse event and/or of products from the same and/or different batches/lots. It should be noted that trend analysis typically is not considered sufficient as a stand-alone method, but should be used in conjunction with other investigation methods for providing for instance complementary information. | Delete | Delete- use b11 instead - Historical Data Analysis |  | Retire B12 |  | ACCEPT | The description of B12 overlaps with B11. 'Trend Analysis' is captured by 'Historical Data Analysis'.
3 | 2025-08-29 | IMDRF AET WG | D | 2025 | D1401 | Unable to exclude device problem | Level 2 | The investigation findings do not clearly confirm the presence or absence of the reported problem with the device. | Modify | D14 L2 terms to be moved to C19 | This should be captured as a finding not a concluded root cause. D14 L2 terms to be moved to C19 | C19 |  | ACCEPT | Agreed that this should be captured as a finding not a concluded root cause. D14 L2 terms to be moved to C19 | C1901
4 | 2025-08-29 | IMDRF AET WG | D | 2025 | D1402 | Device problem excluded | Level 2 | The investigation findings clearly confirm that there was no problem with the device. | Modify | D14 L2 terms to be moved to C19 | This should be captured as a finding not a concluded root cause. D14 L2 terms to be moved to C19 | C19 |  | ACCEPT | Agreed that this should be captured as a finding not a concluded root cause. D14 L2 terms to be moved to C19 | C1902
5 | 2025-08-29 | IMDRF AET WG | F | 2025 | F2201 | Biopsy | Level 2 | Patient required removal of tissue or fluid specimen to establish a diagnosis. | Modify | Biopsy (Patient required removal of tissue or fluid specimen to establish a diagnosis.) Either redefine to exclude blood sample or revise term to include blood sample. |  | None | None | MODIFY | Agreed it was important to capture both blood samples and biopsys. Agreed to make term more generic. |  | New Term: Patient Sample New Definition: Patient required removal of tissue or fluid specimen to establish a diagnosis, including biopsy and blood samples.
6 |  | AdvaMed | F | 2025 | F11 | Non-serious Injury/Illness/Impairment | Level 1 | A non-serious injury, illness or impairment which can be treated with minimal or no intervention, including monitoring only. | Modify | For discussion, if we want to add that F11 and F12 should only be used if there are no other more appropriate F terms. | Modify definition | NA - Editorial |  | REJECT | Feedback provided on the guidance document which was taken into account during the maintenance cycle. Further information has been added to the guidance document to resolve this issue.
7 |  | AdvaMed | F | 2025 | F12 | Serious Injury/Illness/Impairment | Level 1 | A severe injury, illness or impairment which requires hospitalization or medical intervention. | Modify | For discussion, if we want to add that F11 and F12 should only be used if there are no other more appropriate F terms. | Modify definition | NA - Editorial |  | REJECT | Feedback provided on the guidance document which was taken into account during the maintenance cycle. Further information has been added to the guidance document to resolve this issue.
8 | 2025-08-29 | IMDRF AET WG | A | 2025 | A17 | Compatibility Problem | Level 1 | Problem associated with compatibility between device, patients or substances (medication, body fluid, etc.) | Modify | Remove the reference to patient | IMDRF will discuss internally if they want this split into specific incompatability types. Noting that definition here includes 'patients' but there is Level 2 term of Patient-Device Incompatability under A01 and there's a separate Level 1 term for 'Environmental Compatibility Problem' If term will ramain as is then MSSO Action: change request for new MedDRA term 'Device incompatibility issue' for review/consensus by IMOs - Remove the reference to patient | None |  | ACCEPT | Agree to remove reference to patient from the definition. |  | Problem associated with compatibility between device or substances (medication, body fluid, etc.)
9 | 2025-08-29 | IMDRF AET WG | A | 2025 | A0402 | Burst Container or Vessel | Level 2 | Problem associated with the pressure inside a vessel or container rising to such a degree that the container ruptures. | Modify | should the word vessel be included | should the word vessel be included | None |  | MODIFY | Editorial change - remove vessel, add brackets for tubing and bottle. or lines |  | New Term: Burst Container New Definition: Problem associated with the pressure inside a vessel or container (tubing, bottle or lines) rising to such a degree that the container ruptures
10 |  | IMDRF AET WG | C | 2025 | C20 | No Findings Available | Level 1 | Use when no investigation can be performed and therefore no results will be obtained. | Modify | Consider adding statement to use only with B16 E(or very vague reports where literally no investigation can be carried out.) |  |  |  | ACCEPT | Agreed to amend definition |  | New Definition: Use when no investigation is performed or when the investigation is insufficient to reasonably exclude a device malfunction with adequate certainty. This code should typically be used in conjunction with B16 and D13.
11 | 2025-08-29 | IMDRF AET WG | B | 2025 | BXX | Incorrect Manufacturer Attribution | Level 1 | Further information was obtained which established that the manufacturer of the device involved was not the one to which it was initially attributed. This includes mistaken identity or falsified device. Use in conjunction with C20 and D13 or D14. | Add | Further information was obtained which established that the manufacturer of the device involved was not the one to which it was initially attributed. This includes mistaken identity or falsified device. Use in conjunction with C20 and D13 or D14. |  |  |  | ACCEPT | New code required to describe where the manufacturer attribution was incorrect, either due to mistaken identity or intentional falsification. | B33 | Incorrect Manufacturer Attribution Further information was obtained which established that the manufacturer of the device involved was not the one to which it was initially attributed. This includes mistaken identity or falsified device. Use in conjunction with C20 and D13 or D14.
12 | 29/08/2025 | PMDA, Japan | F | 2025 | F07 | Exacerbation of Existing Condition | Level 1 | Use of the device has lead to worsening of the existing disease or condition. | Modify | It would be preferable to add the explanation of examples described in the column "Example of an incident which would be coded using the proposed term" to the definition. | Because of adding F30, the difference of F07 and F30 is slightly difficult to understand. | None | To be discussed if the concepts and examples below are proper. 1. F07 Basic concept: Impact on the patient condition or the planned treatment which is not directly associated with the intended use of the subject device. Example: Worsening of pneumonia after implantation of a pacemaker. 2. F30 Basic concept: The disease or condition associated with the intended use of the subject device is worsening and potential relevance between the worsening and device performance is considerd. Example: To be discussed. It is preferable to give the examples relatively obvious adverse event of the device rather than the events that seem to be it's perfomance limits. How about the cases IVD's false negative result leads to missing and subsequent progression of the disease. | MODIFY | Agreed to add wording to definition to clarify difference between the codes, instead of adding direct device examples |  | Use of the device has lead to worsening of the existing disease or condition, as a direct consequence of device use,
13 | 29/08/2025 | PMDA, Japan | F | 2025 | F30 | Disease Progression | Level 1 | The worsening of a disease over time related to device performance. | Modify | It would be preferable to add the explanation of examples described in the column "Example of an incident which would be coded using the proposed term" to the definition. | Because of adding F30, the difference of F07 and F30 is slightly difficult to understand. | None | To be discussed if the concepts and examples below are proper. 1. F07 Basic concept: Impact on the patient condition or the planned treatment which is not directly associated with the intended use of the subject device. Example: Worsening of pneumonia after implantation of a pacemaker. 2. F30 Basic concept: The disease or condition associated with the intended use of the subject device is worsening and potential relevance between the worsening and device performance is considerd. Example: To be discussed. It is preferable to give the examples relatively obvious adverse event rather than the events that seem to be perfomance limits of the device. How about the cases IVD's false negative result leads to missing and subsequent progression of the disease. | MODIFY | Agreed to add wording to definition to clarify difference between the codes, instead of adding direct device examples |  | The worsening of a disease over time related to device performance (e.g. insufficient therapeutic effect, delayed action).
14 | 2025-08-29 | Health Canada | E | 2025 | E08xx | Hyporeflexia (Ocular) or Diminished Ocular Reflexes | Level 2 | Diminished or absent reflex responses in ocular structures, such as pupillary, corneal, or oculo-vestibular reflexes. | Add | New; primary in eyes, secondary in nervous system? | No appropriate term in eye injuries, none of the terms describe loss of reflexes. We also have terms for nerve damage, but nothing specific to reflexes or the cranial nerves (save the optic nerve). In this instance we could potentially just code E2113 "oversedation", but if we ever need to capture diminished/abesnt reflexes when not a result of oversedation it wouldn't work. | No real overlap | " Patient became overly sedated/burst suppressed. Changes in assessment findings pre and post medical device incident include: loss of corneal reflexes, levophed requirements increased from 0 to 6 mcg/min to meet MAP goal." | MODIFY | Agreed to add but to make more general under nervous system, rather than eye. | E0146 | New Term: Hyporeflexia New Definition: Diminished or absent reflex responses
15 | 2025-08-29 | Health Canada | E | 2025 | E06XX | Abnormal Precordial Findings | Level 2 | Unusual physical findings on precordial examination, including palpable vibrations, thrills, or ventricular heaves. | Add | New | No way to report these findings, which may have clinical significance. | None | It was reported ...post an index procedure completion using a drug-coated balloon catheter...the subject developed a serious adverse event of arteriovenous fistula clot at venous anastomosis due to decreased access blood flow, access thrombosis, diminished or abnormal thrill, and difficult puncture which required an arteriovenous access circuit reintervention | ACCEPT | Agree to add | E0630 | New Term: Abnormal Precordial Findings NewDefinition: Unusual physical findings on precordial examination, including palpable vibrations, thrills, or ventricular heaves.
16 | 2025-08-29 | Health Canada | E | 2025 | E09XX | Subcutaneous Emphysema | Level 2 | Presence of gas or air within subcutaneous or deeper soft tissues; may be described as soft tissue emphysema or subdermal gas accumulation | Add | New | Some similar conditions have codes (pneumothorax, injection site pooling) in the terms, but gas can be a result of other internal inuries (e.g. esophageal perforation), external sources (penetrating trauma) or even infection. | Some overlap with existing terms describing gas in the wrong spots, but this is genrally describing subcutanous or gas spreading along muscle fascia. | regarding a chest drainage system - "Upon assesment by new RN the airleak and subcutaneous emphysema noted to upper chest." | MODIFY | Agree to add but change mapping to E17 rather than E09. | E1731 | New Term: Subcutaneous Emphysema New Definition: Presence of gas or air within subcutaneous or deeper soft tissues; may be described as soft tissue emphysema or subdermal gas accumulation
17 | 2025-08-29 | Health Canada | E | 2025 | E05XX | Vasoplegia | Level 2 | A condition characterized by abnormally low systemic vascular resistance, often presenting with hypotension and preserved or elevated cardiac output; may be referred to as vasodilatory shock or vascular tone loss | Add | New | While low blood pressure can be captured, the combination of vascular dilation and normal cardiac output isn't. | None | "the patient had septicemia perioperatively post Left Ventricular Assist Device (LVAD) and right heart failure post operatively with severe vasoplegia and tamponade " | ACCEPT | Agree to add | E0524 | New Term: Vasoplegia New Definition: A condition characterized by abnormally low systemic vascular resistance, often presenting with hypotension and preserved or elevated cardiac output; may be referred to as vasodilatory shock or vascular tone loss
18 | 2025-08-29 | Health Canada | E | 2025 | E1804XX | Device-Associated Sarcoma? Bone sarcoma? Osteosarcoma? | Level 3 | Malignant mesenchymal neoplasm, such as osteosarcoma, chondrosarcoma, or soft tissue sarcoma, occurring in proximity to or in association with implanted orthopedic devices or hardware. | Add | New | As a type of sarcoma associated to devices, may be worth capturing. We had 17 cases of bone sarcoma. Chondrosarcoma is also possible, hence suggesting a possible broader term. | Under solid tumour? | Several reports of bone sarcomas arising from the article https://pubmed.ncbi.nlm.nih.gov/39243359/ | REJECT | Sarcoma is covered by solid tumour.
19 | 2025-08-29 | Health Canada | E | 2025 | E16XX | Patella Maltracking | Level 2 | Abnormal alignment or movement of the patella relative to the trochlear groove during knee flexion; may be described as patellar instability or lateral patellar displacement | Add | New | Similar term of patella/glenoid erosion suggested in 2022-23 as a device problem and rejected as it should be a patient problem. We have several refeences to patella maltracking in our reports | None | "<less than a year> years after <TKA left revision>, the patient sustained a patella maltracking that required a TKA re-revision. It is unknown which devices were explanted/exchanged during this procedure." | MODIFY | Modified to align with MedDRA | E1653 | New Term: Patellar Tracking Disorder New Definition: Abnormal alignment or movement of the patella relative to the trochlear groove during knee flexion; may be described as patellar instability or lateral patellar displacement
20 | 2025-08-29 | Health Canada | E | 2025 | E10XX | Ileus | Level 2 | Reduced or absent intestinal motility resulting in impaired transit of intestinal contents; may be referred to as paralytic ileus, non-obstructive ileus, or functional bowel obstruction | Add | New | We have no way to capture teither obstructive ileus (e.g., caused by adhesions to a bad staple line) or post operative or paralytic ileus. We had 15 examples in the databae. | None | "<Stapler> used for a laparoscopic therapeutic treatment of ileocecal resection & anastomosis. It was reported that after the implant, the patient experienced abscess, ileus, intrabdominal hemorrhage, sudden deterioration, decreased oxygen saturations, pleural effusions, sepsis, foul-smelling bloody fluid, hole in staple line of anastomosis ..." | ACCEPT | Agree to add | E1037 | New Term: Ileus New Definition: Reduced or absent intestinal motility resulting in impaired transit of intestinal contents; may be referred to as paralytic ileus, non-obstructive ileus, or functional bowel obstruction
21 | 2025-08-29 | Health Canada | E | 2025 | E10XX | Enteritis | Level 2 | Inflammation of the small intestine, including duodenitis, jejunitis, or ileitis; may be associated with mucosal edema or intestinal wall thickening | Add | New | Gastritis and Colitis already exist, but ileitis, jejunitis and duodenitis arern't captured under those terms. | None if added, but one could group gastritis, colitis, and enteritis all together under a term for any inflammations of the GI tract like IBD. | "<Stapler> was used for therapeutic treatment of Hartmann's procedure. It was reported that after the implant, the patient experienced labs abnormal for WBC; leucocyte, fluid collection, hematoma, bowel obstruction, adhesions, discharge, ileitis, torsion of bowel, weakness, fever, somnolence, hypotension, hallucinations, purulent drainage from wound, infarcted sigmoid colon..." | ACCEPT | Agree to add | E1038 | New Term: Enteritis New Definition: Inflammation of the small intestine, including duodenitis, jejunitis, or ileitis; may be associated with mucosal edema or intestinal wall thickening
22 | 2025-08-29 | Health Canada | E | 2025 | E2327XX | Infarction | Level 3 | Irreversible tissue necrosis resulting from inadequate blood supply; commonly referred to as ischemic necrosis or tissue infarct | Add | New | Infarction could be captured as necrosis, or could be a subterm of necrosis if we want to be specific about the reason for tissue death. If rejected, could perhaps edit the definition of necrosis to include mention of infarction to aid in term selection. | Possibly locate under necrosis | "<Stapler> was used for therapeutic treatment of Hartmann's procedure. It was reported that after the implant, the patient experienced labs abnormal for WBC; leucocyte, fluid collection, hematoma, bowel obstruction, adhesions, discharge, ileitis, torsion of bowel, weakness, fever, somnolence, hypotension, hallucinations, purulent drainage from wound, infarcted sigmoid colon..." | REJECT | Reject. This can also be coded by using Ischemia together with Necrosis
23 | 2025-08-29 | Health Canada | E | 2025 | E21xx | Local Anesthetic Systemic Toxicity (LAST) | Level 2 | A systemic adverse reaction resulting from elevated plasma concentrations of local anesthetics | Add | New | Since this could be related to volume or speed of delivery, as well as delivery into vascular spaces, it seems like the kind of reaction that could be device-related and isn't well captured elsewhere |  | " the pump emptied at least a day prior to what was expected. One patient experienced symptoms of Local anesthetic systemic toxicity (LAST)." | ACCEPT | Agree to add | E2133 | New Term: Local Anesthetic Systemic Toxicity (LAST) New Definition: A systemic adverse reaction resulting from elevated plasma concentrations of local anesthetics
24 | 2025-08-29 | Health Canada | E | 2025 | E23XX | Oral Soft Tissue Recession | Level 2 | Retraction or loss of gingival tissue along the buccal or lingual surfaces of the teeth | Add | New | Some devices could cause this effect, like tooth alignment devices worn for long periods. It's not really well captured under any of the other major categories, as it's not truly musculoskeletal, skin and subcutaneous, or gastrointestinal. |  | In one example a clear orthodontic aligner was implicated in causing buccal and lingual recession. | MODIFY | Modified to align with MedDRA. Mapped to E21 rather than E23. | E2136 | New Term: Gingival Recession New Definition: Retraction or loss of gingival tissue resulting in exposure of the root surface
25 | 2025-08-29 | Health Canada | E | 2025 | E23XX | Periodontal Disease | Level 2 | A chronic inflammatory condition affecting the supporting structures of the teeth, including the gingiva, periodontal ligament, and alveolar bone | Add | New | With a variety of dental devices being reported on, there are cases where gingivitis/periodontitis and such are alleged to have been caused by the device. |  | "patient's tooth 4.1 had to be extracted due to periodontal disease while using the <device> and stage 7 treatment" | ACCEPT | Agree to add | E2354 | New Term: Periodontal Disease New Definition: A chronic inflammatory condition affecting the supporting structures of the teeth, including the gingiva, periodontal ligament, and alveolar bone
26 | 2025-08-29 | Health Canada | E | 2025 | E05XX | Abnormal Central Venous Pressure (CVP) | Level 2 | A deviation from normal central venous pressure values | Add | New | High CVP can be due to too much fluid in the system(and thus be a sign of overinfusion) or can reflect a problem with the ability of the heart to pump blood. |  | "The anesthesia assistant ... called at 11:43 reporting 'low fluid alarm' and high CVP of 26" | MODIFY | Modified to align with MedDRA | E0525 | New Term: Central venous pressure abnormal New Definition: A deviation from normal central venous pressure values
27 | 2025-08-29 | Health Canada | E | 2025 | E06XX | Cardiac Vegetation | Level 2 | An abnormal growth or mass, typically composed of thrombotic material and microorganisms, attached to cardiac valves or endocardial surfaces | Add | New | While we could use a general code like "infection" to perhaps capture vegetations, the association of these with implanted devices may make it worth adding to capture as a distinct symptom |  | Four years and four and a half months post implant ... a small echogenic mass with attachment to the left ventricular outflow tract (LVOT) was noted. Endocarditis and vegetation was suspected. " | REJECT | Reject. This is covered by E2123
28 | 2025-08-29 | Health Canada | E | 2025 | E10XX | Gagging/Retching | Level 2 | Spasmodic and periodic contractions of the respiratory muscles (diaphragm and muscles of the chest and abdominal wall) without expulsion of the stomach contents | Add | New | We have a code for "vomit" and for nausea, but if a report simply says that "it made me gag" neither term captures it. |  | " the patient had been hospitalized with Diabetic Ketoacidosis (DKA) ... Symptoms reported include hyperglycemia, shortness of breath, dry mouth, gagging, and vomiting. | MODIFY | Modified to align with MedDRA | E1039 | New Term: Retching New Definition: Spasmodic and periodic contractions of the respiratory muscles (diaphragm and muscles of the chest and abdominal wall) without expulsion of the stomach contents
29 | 2025-08-29 | Health Canada | E | 2025 | E21xx | Periimplantitis | Level 2 | Periimplantitis is a pathological condition occurring in the tissues around dental implants, characterized by inflammation in the peri-implant mucosa and progressive loss of supporting bone | Add | New | This term was previously suggested by MedTech Europe and the suggestion was to use E2123 - Medical device site infection. There are however, a number of reports with this condition listed, and it may be worth considering eitehr explicitly adding this to the definitionof medical device site infection, or to adding a sub-code? It was requested 27 times by HC staff. |  | "It was reported that the implant at tooth site #36 lost integration due periimplantitis and was removed." | REJECT | Reject. This is covered by E2123
30 | 2025-08-29 | Health Canada | E | 2025 | E12XX | Malnutrition | Level 2 | A state of nutritional deficiency or imbalance | Add | New | Issues with gastric surgery can result in malnutrition. There is no good code for this in the AET. Not the same as E1203 - Feeding Problem. |  | "uneventful laparoscopic gastric bypass surgery … Approximately five months following surgery, patient presented 'with malnutrition without eating difficulties.'" | ACCEPT | Agree to add | E1216 | New Term: Malnutrition New Definition: A state of nutritional deficiency or imbalance
31 | 2025-08-29 | Health Canada | E | 2025 | E1207 | Polydipsia | Level 2 | Extreme thirst accompanied by chronic excessive intake of water. | Modify | Suggested change - Extreme thirst, often accompanied by chronic excessive intake of water" | It's generally accepted that the excessive thirst in symptoms of daibetic ketoacidosis are polydpisia, but many observations simply indicate thrist, failing to indicate drinking of exceessibve quantities of water. |  | "Patient's BG at time of incident: 23 mmol/l. Patient's Current BG Value (at time of call): 24 mmol/L. Is customer currently reporting any symptoms related to their High BGs: Yes Difficulty breathing, shaking, thirst, ... | ACCEPT | Agree to modify |  | Extreme thirst, often accompanied by chronic excessive intake of water
32 | 2025-08-29 | Health Canada | E | 2025 | E1008 | Diarrhea | Level 2 | Frequent, watery bowel movements. | Modify | Suggested change - Passage of loose or watery stools, often occurring with increased frequency. | We get mentions of short prhases like "watery stools" in our reports, but the requirement of increased frequency mmakes term selection difficult. The lack of frequency doesn't necessarily signal the absence of it, and such a change may capture more true cases of diarrhea than it would introduce false positives. Although frequency is not specified, the description of watery stool in the context of significant gastrointestinal pathology and systemic symptoms supports classification as diarrhea for adverse event coding purposes. |  | "It was reported that after the implant, the patient experienced fluid collections, ileus, anastomotic leak, dehiscence, inflammatory adhesions, descending colon necrotic, perforation, fever, diaphoresis, chills, watery stool …" | ACCEPT | Agree to modify |  | New Definition: Passage of loose or watery stools, often occurring with increased frequency.
33 | 2025-08-29 | Health Canada | F | 2025 | F01XX | Delay in Therapeutic Response | Level 2 | An increase in the time taken to achieve the desirable and beneficial effects resulting from a medical treatment | Add | New | It certainly falls under F01 as a change in therapeutic response, and thus could simply be coded that way, but as we have an "increase" and a "decrease" referencing the magnitude or degree of response, we could similarly speak on the timing of a response. | None | There was a delayed patient response to the bolus administered following the under-infusion.' | ACCEPT | Agreed to add new code to capture this issue as it cannot be covered by current codes. | F0104 | New Term: Delay in Therapeutic Response New Definition: An increase in the time taken to achieve the desirable and beneficial effects resulting from a medical treatment
34 | 2025-08-29 | Health Canada | A | 2025 | A0510 | Retraction Problem | Level 2 | Problem associated with drawing back the device to an intended location. | Modify | Problem associated with drawing back the device or part of the device to an intended location. | This definition was causing problems with some staff arguing that it only applies if the device is moved, and that parts of the device are not covered by this code. However, many codes make reference to the device, and it may be easier to simply add that codes can refer tot he devices or the components of the device. Similar issues were brought up for C codes for Interoperability problems or Communciations problems. |  | Examples would be needles not fully drawing back into the device; this could be captured under an activation or positioning problem perhaaps, but the mecahnical nature of it and the prsence of a code for retraction problems suggests it might fit well here. | ACCEPT | Minor editorial change |  | New Definition: Problem associated with drawing back the device or part of the device to an intended location.
35 | 2025-08-29 | BfArM | F | 2025 | F2301 | Additional device required | Level 2 | Use of additional or alternative device required to achieve optimal outcome | add childs | F2301 Replacement device required (Use of an additional, identical device); F2302 Additional device of different type required (use of an additional device required to prevent further harm to the patient) | current code doesnt allow for a distinction. Both scenarios pose different risks and are interesting to track | None | F2301 a non-functioning device had to be replaced during the implant procedure F2302 A external defibrillator was needed to save the patient when the implanted ICD was not able to convert the rhythm | ACCEPT | add 2 new child code for alternative device required or similar kind | F230101 / F230102 | New: The situation where an additional device of the same type and intended purpose as the original device is required to continue or complete the intended procedure or therapy. New: The situation where an additional device of a different type or intended purpose is required to mitigate risk or prevent harm to the patient following the use or failure of the original device. This may involve introducing a complementary or alternative device to address complications
36 | 2025-08-29 | BfArM | A | 2025 | new | leakage current | Level 3 | an unintended electric current that flows through insulation, the enclosure, grounding, or patient-connected components back to the mains or earth. | Add | add | can not be coded by any of the existing terms if it doesnt manifest to the user but is measured by the manufacturer during maintenance service |  | It was reported that the programmer failed the safety test for device leakage current values as they were too high. | REJECT | This can be coded as A072101
37 | 2025-08-29 | BfArM | B | 2025 | BXX | Software Code Analysis | Level 1 | systematic review and testing of source code, through manual and automated methods, to detect errors, vulnerabilities, and deviations from requirements | Add | add | there is no existing code to cover the systematic review of a device software |  | It was determined the cause is due to an incorrect implementation in a specific section of the code where a parameter variable was not maintained | ACCEPT | there is no existing code to cover the systematic review of a device software | B34 | New Term: Software Code Analysis New Definition: Systematic review and testing of source code, through manual and automated methods, to detect errors, vulnerabilities, and deviations from requirements.
38 | 2025-08-29 | BfArM | D | 2025 | D1401 | Unable to exclude device problem | Level 2 | The investigation findings do not clearly confirm the presence or absence of the reported problem with the device. | Delete | delete or provide additional guidance | there is no scenario that can be coded with D1401 that could not also be coded with D15. |  |  | ACCEPT | Fixed per items 3 and 4.
39 | 2025-08-29 | BfArM | D | 2025 | new | Cause Traced to Installation |  | Cause Traced to non-conform installation of device | Add | add | currently no other code available; would go together with C1603 - Installation problem |  | wrong configuration during installation or in general non-conform installation of device | REJECT | This scenario can be coded using C16 and D17
40 | 2025-08-29 | BfArM | G | 2025 | G0405211 | Retainer | Level 3 | A physical component designed to hold items in place. | Modify | combine? | clarify differentiation from G04068 - Holder A component designed to hold another object. |  | fixation of the monitor ceiling suspension the securing segment was not installed correctly | MODIFY | a physical component designed to hold a fastener in place. AMEND holder definition to include the term bracket |  | New Definition: A physical component designed to hold a fastener in place.
41 | 2025-08-22 | AbbVie | A | 2025 | A020501 | Difficult to Open or Remove Packaging Material | Level 3 | Current - Problem associated with difficulty for users to operate the device, specifically as it relates to the opening or removal of the outer wrapping. Proposed - Problem associated with difficulty for users to operate the device, specifically as it relates to the opening or removal of the packaging materials. | Modify | Modification of definition | Remove "outer wrapping" and Add "packaging materials" for clarity. | None | For Breast implants | ACCEPT | Minor editorial change |  | New Definition: Problem associated with difficulty for users to operate the device, specifically as it relates to the opening or removal of the packaging materials
42 | 2025-08-22 | AbbVie | A | 2025 | A020502 | Incomplete or Missing Packaging | Level 3 | Current - Problem associated with the nonconformance to the device specifications due to incomplete or missing packaging that may compromise the device operation as intended. Proposed - Problem associated with incomplete or missing packaging that may compromise the device operation as intended. | Modify | Modification of definition | Remove "the nonconformance to the device specifications due to" for clarity. | None | Applicable for Keller Funnel | ACCEPT | Minor editorial change |  | New Definition: Problem associated with incomplete or missing packaging that may compromise the device operation as intended.
43 | 2025-08-22 | AbbVie | A | 2025 | A020505 | Device packaging look-alike | Level 3 | Current - Problem associated with medical device packaging that looks similar to another product leading to confusion. The other product may be another device or a different product such as a medicine or cleaning agent. Use A210107 for Device label look-alike. Proposed - Problem associated with medical device packaging that looks similar to another product leading to confusion. The other product may be another device or a different product altogether. Use A210107 for Device label look-alike. | Modify | Modification of definition | Remove "such as a medicine or cleaning agent" and add "altogether" for consistency. | None | Applicable for Keller Funnel | REJECT | The examples in the definition provide clarity. IMDRF WG thinks this is important to keep.
44 | 2025-08-22 | AbbVie | A | 2025 | A020701 | Delivered as Unsterile Product | Level 3 | Current - Problem associated with a device being received in such a manner to indicate that its sterility has been compromised (e.g. sterile packaging breached, visible contaminate present) Proposed - Problem associated with a device being received in such a manner to indicate that its sterility has been compromised (e.g. sterile packaging breached, visible biological contaminants present) | Modify | Modification of definition | Add "biological contaminants" and remove "contaminate". | None | Applicable for Fourte | REJECT | Proposal is sufficiently covered by current definition. Adding this example will exclude other types of contamination.
45 | 2025-08-22 | AbbVie | A | 2025 | A030205 | Particulates | Level 3 | Current - Substances that consist of separate particles that are introduced by the device during use. Proposed - Substances that consist of separate particles that are introduced by the device during or before use. | Modify | Modification of definition | Add "or before" for clarity. | None | Applicable for Dermal Fillers | ACCEPT | Minor editorial change |  | Substances that consist of separate particles that are introduced by the device during or before use.
46 | 2025-08-22 | AbbVie | A | 2025 | A0401 | Break | Level 2 | Current - Problem associated with undesired damage or breakage of those materials used in the device construction. Proposed - Problem associated with breakage of those materials used in the device construction. | Modify | Modification of definition | Remove "undesired damage or" for clarity. | None | Applicable for Fourte | REJECT | Removing "undesired" does not clarify the definition.
47 | 2025-08-22 | AbbVie | A | 2025 | A040102 | Loss or Failure to Bond | Level 3 | Current - Loss of or Failure to Bond Proposed - Loss or Failure to Bond | Modify | Modification of Term Name | Remove "of" for clarity. | None | For Breast implants | REJECT | Inclusion of 'loss of', is important and aligns with the definition.
48 | 2025-08-22 | AbbVie | A | 2025 | A040606 | Material Invagination | Level 3 | Current - Problem associated with an undesired material change in shape, characterized by the infolding of one part within another part of a structure. Proposed - Problem associated with an undesired material change in shape, characterized by the infolding of one part within another part of a structure, or creates an internal pocket within the device's structure. | Modify | Modification of definition | Add "or creates an internal pocket within the device's structure" for clarity. | None | For Breast implants | ACCEPT | Minor editorial change |  | New Definition: Problem associated with an undesired material change in shape, characterized by the infolding of one part within another part of a structure, or creates an internal pocket within the device's structure.
49 | 2025-08-22 | AbbVie | A | 2025 | A0407 | Material Discolored | Level 2 | Current - Problem associated with an undesired streak, pattern and/or a noticeable change in color from the rest of the materials used in the device construction. Proposed - Problem associated with an undesired streak, pattern and/or a noticeable change in original color. | Modify | Modification of definition | Add "original" and remove "from the rest of the materials used in the device construction.". | None | For Breast implants | ACCEPT | Minor editorial change |  | New Definition: Problem associated with an undesired streak, pattern and/or a noticeable change in original color.
50 | 2025-08-22 | AbbVie | A | 2025 | A0504 | Leak/Splash | Level 2 | Current - Problem associated with the escape of a liquid (including blood and bodily fluids), gas or radiation from the vessel or container in which it is housed. Proposed - Problem associated with the escape of a liquid (including blood and bodily fluids), gel, gas or radiation from the vessel or container in which it is housed. | Modify | Modification of definition | Add "gel" for clarity. | None | Applicable for Dermal Fillers | ACCEPT | Minor editorial change |  | Problem associated with the escape or ingress of a liquid (including blood and bodily fluids), gel, gas or radiation from or into a vessel or container.
51 | 2025-08-22 | AbbVie | E | 2025 | A05xx | Injector Slider Resistance | Level 2 | Slider resistance on injector | Add | New Term | No current IMDRF code Needed for XEN product No corresponding MedDRA PT | None | Need to differentiate between the injector and Xen device | REJECT | These are not clinical signs or symptoms.
52 | 2025-08-22 | AbbVie | A | 2025 | A1415 | Air/Gas in Device | Level 2 | Currently - Problem associated with air in the device (e.g. air in line). Proposed - Problem associated with air in the device (e.g. air in line, syringe, or bubbles in gel breast implant). | Modify | Modification of definition | Provides clarification and additional examples | None | Material integrity complaints of "bubbles" for gel breast implants. | MODIFY | Minor editorial change |  | New Definition: Problem associated with air or gas in the device (e.g. air in line).
53 | 2025-08-22 | AbbVie | E | 2025 | A1502xx | Unsuccessful Stent Placement | Level 3 | Unsuccessful final stent placement | Add | New Term | No current IMDRF code Needed for XEN product MedDRA does have LLT: Procedural failure but a more specific IMDRF code would be required as there is usually a back up stent that may be used. Having an unsuccessful placement is rare and noteworthy, | None | Related to XEN product. Device place could not be adjusted, used back up device. Looking to have specificity of the stent placement and not the "injector". There are times the injector may be in the appropriate position, but the stent still was not successfully placed upon being released from the injector. | REJECT | These are not clinical signs or symptoms.
54 | 2025-08-22 | AbbVie | A | 2025 | A16xx | Preventative Maintenance | Level 2 | Alert or notification of a routine health check for regular/preventative maintenance. | Add | New Term | No current IMDRF code for "health checks" or regular/preventative maintenance | None | Applicable for CoolSculpting | REJECT | This is not a malfunction so inapropriate to include in A codes.
55 | 2025-08-22 | AbbVie | A | 2025 | A1801 | Contamination | Level 2 | Current - Problem associated with the undesired introduction of impurities either chemical or microbiological in nature, or of foreign matter into or onto the device. Proposed - Problem associated with the undesired introduction of impurities either chemical or microbiological in nature, or of foreign matter onto the device. | Modify | Modification of definition | Remove "into or" for clarity. | None | Applicable for Dermal Fillers | REJECT | There may be cases in which contamination occurs in the device.
56 | 2025-08-22 | AbbVie | A | 2025 | A180101 | Biofilm coating in Device | Level 3 | Current - Biofilm coating in Device Proposed - Biofilm coating in/on Device | Modify | Modification of Term Name | Add "on device" to align with the IMDRF defintion. | None | For Breast implants | ACCEPT | Minor editorial change |  | New Term: Biofilm Coating in/on Device
57 | 2025-08-22 | AbbVie | A | 2025 | A1801xx | Transmission of Communicable Disease | Level 3 | Unintensional transmission of communicable diseases | Add | New Term | No current IMDRF code. The change is necessary to accommodate human tissue product. | None | Applicable to human tissue (e.g. process problem when present on donor tissue). "Revolve" product due to user cross contamination. | REJECT | This issue is better suited to Annex E.
58 | 2025-08-22 | AbbVie | A | 2025 | A210101 | Expiration Date Error | Level 3 | Current - Problem associated with errors in identification of expiration date. Proposed - Problem associated with the wrong expiration being displayed. | Modify | Modification of definition | Remove "with errors in identification of expiration date." and add "with the wrong expiration being displayed". | None | Applicable for Keller Funnel | MODIFY | Agree this is important to capture. Have decided to amend term and definition to include all expiration date problems. |  | New Term: Expiration Date Problem New Definition: Any problem involving the accuracy, legibility, presence, or interpretation of an expiration date on a medical device, its labeling, or associated documentation.
59 | 2025-08-22 | AbbVie | A | 2025 | A210103 | Inaccurate Information | Level 3 | Current - Problem associated with imprecise, inexact information. Proposed - Problem associated with imprecise, incorrect inexact information. | Modify | Modification of definition | Add "incorrect" to align with the IMDRF Term "Inaccurate". | None | For Breast implants | ACCEPT | Minor editorial change |  | New Definition: Problem associated with imprecise, incorrect inexact information.
60 | 2025-08-22 | AbbVie | A | 2025 | A210107 | Device label look-alike | Level 3 | Current - Problem associated with medical device label that looks similar to another product leading to confusion. The other product may be another device or a different product such as a medicine or cleaning agent. Use A020505 for Device packaging look-alike. Proposed - Problem associated with medical device labeling that looks similar to another product leading to confusion. The other product may be another device or a different product altogether. Use A020505 for Device packaging look-alike. | Modify | Modification of definition | Change "label" to "labeling" and remove "such as a medicine or cleaning agent"and add "altogether" for consistency. | None | Applicable for Keller Funnel | MODIFY | Agree to modify but retain examples in definition as these are important. |  | New Definition: Current - Problem associated with medical device label that looks similar to another product leading to confusion. The other product may be another device or a different product such as a medicine or cleaning agent, or other product altogether. Use A020505 for packaging look-alike issues.
61 | 2025-08-22 | AbbVie | A | 2025 | A210107 | Device label look-alike | Level 3 | Current - Device label look-alike Proposed - Device labeling look-alike | Modify | Modification of Term Name | Change "label" to "labeling for consistency. | None | Applicable for Keller Funnel | ACCEPT | Minor editorial change |  | New Term: Device Labeling Look-Alike
62 | 2025-08-22 | AbbVie | E | 2025 | Current - E1803 Proposed - E17xx | Nodule | Level 2 | Current - A small lump, swelling or collection of tissue. Proposed - A small lump, swelling, inflammatory, or collection of tissue. | Modify | Modification of Code hierarchy and definition | Either change E1803 to E17xx or add Level 2 E1803 code within the E17 hierachy (as this aligns w/ MedDRA). Provides clarification of an "inflammatory nodule". | None | For dermal fillers it's necessary to differentiate nodules from Imflammatory versus non-imflammatory. Subcutaneous Nodules are defined as non-imflammatory and Imflammatory nodules are coded to Nodule. | REJECT | Code with inflammation. This is not confined to skin issues
63 | 2025-08-22 | AbbVie | E | 2025 | E0119xx | Vasovagal Episode | Level 3 | A common type of fainting caused by a sudden drop in blood pressure and heart rate due to an overreaction of the nervous system. This overreaction can be triggered by various factors like medical procedures, emotional stress, or prolong standing. | Add | New Term | Differentiate from syncope and presyncope. Often related to a medical procedure. | None | Associated w/ Dermal Fillers, Coolsculpting, and breast implant procedures. | REJECT | Covered by E011903
64 | 2025-08-22 | AbbVie | E | 2025 | E01xx | On and off phenomenon | Level 2 | A syndrome of alternating periods of being symptom-free (the “on” period) and an abrupt reappearance of symptoms (the “off” period). | Add | New Term | The addition is necessary to to accommodate a combination product. Aligns with MedDRA 10030312 On and off phenomenon | None | Would assist with combination device for Parkinson's complaints. | REJECT | Suggested to code with decreased therapeutic response. This may not be device related.
65 | 2025-08-22 | AbbVie | E | 2025 | E0202xx | Paranoia | Level 3 | Excessive mistrust or suspicion of people -thinking incorrectly, for example, that someone is trying to hurt you, that people are doing things to deliberately annoy you or that there is a conspiracy against you. | Add | New Term | Aligns with MedDRA 10033864 Paranoia | None | Needed for harzard related to an AbbVie device. | MODIFY | Accept code but modify as level 2. | E0210 | New Term: Paranoia New Definition: Excessive mistrust or suspicion of people -thinking incorrectly, for example, that someone is trying to hurt you, that people are doing things to deliberately annoy you or that there is a conspiracy against you.
66 | 2025-08-22 | AbbVie | E | 2025 | E05xx | Raynaud's Phenomenon | Level 2 | A condition that causes the blood vessels in the extremities to narrow, restricting blood flow. | Add | New Term | No current IMDRF code. Aligns with MedDRA 10037912 Raynaud's phenomenon | None | Associated w/ breast implant compliants. | REJECT | To be coded under E0401
67 | 2025-08-22 | AbbVie | E | 2025 | E05xx | Telangiectasia | Level 2 | Telangiectasia, also known as Spider veins or thread veins, are prominent clusters of small, damaged blood vessels visible in the skin. They appear as thin red, blue, or purple lines. | Add | New Term | No current IMDRF code Aligns with MedDRA 10043189 Telangiectasia | None | For dermal fillers and coolsculpting this code is needed. It's currently getting coded to vasodilation. | ACCEPT | Agree to add | E0526 | New Term: Telangiectasia New Definition: Telangiectasia, also known as Spider veins or thread veins, are prominent clusters of small, damaged blood vessels visible in the skin. They appear as thin red, blue, or purple lines.
68 | 2025-08-22 | AbbVie | E | 2025 | E07XX | Airway Burns | Level 2 | Burn to upper respiratory tract from inhalation of heated gases. | Add | New Term | No current IMDRF code Aligns with LLT's under the MedDRA PT: Airway burns (10067204) | None | Applicable to Resonic, which is using compressed propane gas in a chiller module, that can leak out and ignite from a spark. | ACCEPT | NOTE: Look at moving primary from skin to injury for most burn items | E0762 | New Term: Airway Burns New Definition: Burn to upper respiratory tract from inhalation of heated gases.
69 | 2025-08-22 | AbbVie | E | 2025 | E0822xx | Absence of Subconjunctival Bleb | Level 3 | Absense of a small, blister-like fluid collection that forms on the surface of the eye (underneath the conjunctiva) after surgery (e.g. trabeculectomy), allowing fluid to drain out and reduce eye pressure. | Add | New Term | No current IMDRF code and needed to accommodate XEN device. No corresponding MedDRA PT | None | Needed for XEN product. IMDRF lacking coding to accurately capture this issue. Needed specifity for coding and trending. Code needed in scenerios when possible situations are not understood such as stent incorrect placement or stent drainage issue, etc. | MODIFY | In response to the codes for bleb related issues, we have added a new code for 'Occular Procedural Complications' under E2135 | E2135 | Ocular Procedural Complication Any complication arising from ocular surgical procedures involving the conjunctiva or adjacent tissues, including but not limited to subconjunctival bleb formation (e.g., trabeculectomy, glaucoma filtration procedures). This may encompass issues such as bleb absence, infection, leakage, perforation, structural thinning, scarring, or other adverse changes affecting the surgical site or surrounding ocular structures.
70 | 2025-08-22 | AbbVie | E | 2025 | E0822xx | Blebitis | Level 3 | Infection affecting the tissue forming the fluid-filled subconjunctival blister (bleb). | Add | New Term | No current IMDRF code Needed for XEN product Aligns with MedDRA 10069166 Blebitis | None | Needed for XEN product. IMDRF lacking coding to accurately capture this issue. Needed specifity for coding and trending. | MODIFY | In response to the codes for bleb related issues, we have added a new code for 'Occular Procedural Complications' under E2135 | E2135 | Ocular Procedural Complication Any complication arising from ocular surgical procedures involving the conjunctiva or adjacent tissues, including but not limited to subconjunctival bleb formation (e.g., trabeculectomy, glaucoma filtration procedures). This may encompass issues such as bleb absence, infection, leakage, perforation, structural thinning, scarring, or other adverse changes affecting the surgical site or surrounding ocular structures.
71 | 2025-08-22 | AbbVie | E | 2025 | E0822xx | Cyclodialysis | Level 3 | Is a separation of the longitudinal fibers of the ciliary muscle from the scleral spur, creating an abnormal pathway for aqueous humor drainage. | Add | New Term | No current IMDRF code Needed for XEN product Aligns with MedDRA 10050156 Glaucoma surgery | None | Needed for XEN product. IMDRF lacking coding to accurately capture this issue. Needed specifity for coding and trending. | REJECT | Cyclodialysis is a procedure.
72 | 2025-08-22 | AbbVie | E | 2025 | E0822xx | Glaucoma Progression | Level 3 | Worsening of the increased pressure in the eyeball due to obstruction of the outflow of aqueous humor. | Add | New Term | No current IMDRF code Needed for XEN product Aligns with MedDRA 10018304 Glaucoma | None | Needed for XEN product. IMDRF lacking coding to accurately capture this issue. Needed specifity for coding and trending. | REJECT | This scenario can be coded using Glaucoma and Disease Progression
73 | 2025-08-22 | AbbVie | E | 2025 | E0822xx | Malignant Glaucoma | Level 3 | Elevated IOP with shallow or flat anterior chamber | Add | New Term | No current IMDRF code Needed for XEN product Aligns with MedDRA 10079469 Malignant Glaucoma | None | Needed for XEN product. IMDRF lacking coding to accurately capture this issue. Needed specifity for coding and trending. | REJECT | This scenario can be coded by Glaucoma
74 | 2025-08-22 | AbbVie | E | 2025 | E0822xx | Subconjunctival Bleb Leakage | Level 3 | A leak of the tissue forming the bleb which may occur spontaneously without history of trauma or during surgery. | Add | New Term | No current IMDRF code MedDRA 10069165 Conjunctival filtering bleb leak | None | Needed for XEN when the bleb leakage occurs spontaneously without history of trauma or during surgery and would not be coded to a post-operative leak. | MODIFY | In response to the codes for bleb related issues, we have added a new code for 'Occular Procedural Complications' under E2135 | E2135 | Ocular Procedural Complication Any complication arising from ocular surgical procedures involving the conjunctiva or adjacent tissues, including but not limited to subconjunctival bleb formation (e.g., trabeculectomy, glaucoma filtration procedures). This may encompass issues such as bleb absence, infection, leakage, perforation, structural thinning, scarring, or other adverse changes affecting the surgical site or surrounding ocular structures.
75 | 2025-08-22 | AbbVie | E | 2025 | E0822xx | Subconjunctival Bleb Perforation | Level 3 | A perforation of the tissue forming the bleb which may occur spontaneously without history of trauma or during surgery. | Add | New Term | No current IMDRF code. Needed for XEN product. No corresponding MedDRA PT | None | Needed for XEN when the bleb perforation occurs spontaneously without history of trauma or during surgery and would not be coded to a post-operative leak. | MODIFY | In response to the codes for bleb related issues, we have added a new code for 'Occular Procedural Complications' under E2135 | E2135 | Ocular Procedural Complication Any complication arising from ocular surgical procedures involving the conjunctiva or adjacent tissues, including but not limited to subconjunctival bleb formation (e.g., trabeculectomy, glaucoma filtration procedures). This may encompass issues such as bleb absence, infection, leakage, perforation, structural thinning, scarring, or other adverse changes affecting the surgical site or surrounding ocular structures.
76 | 2025-08-22 | AbbVie | E | 2025 | E0822xx | Subconjunctival Bleb Thinning | Level 3 | Thinning of the tissue forming the bleb, making it more prone to leakage and infection. | Add | New Term | No current IMDRF code. Needed for XEN product. No corresponding MedDRA PT | None | Related to XEN product. Needed for coding specifity and trending. | MODIFY | In response to the codes for bleb related issues, we have added a new code for 'Occular Procedural Complications' under E2135 | E2135 | Ocular Procedural Complication Any complication arising from ocular surgical procedures involving the conjunctiva or adjacent tissues, including but not limited to subconjunctival bleb formation (e.g., trabeculectomy, glaucoma filtration procedures). This may encompass issues such as bleb absence, infection, leakage, perforation, structural thinning, scarring, or other adverse changes affecting the surgical site or surrounding ocular structures.
77 | 2025-08-22 | AbbVie | E | 2025 | E0826xx | Hypotony Maculopathy | Level 3 | Ocular hypotony complicated by papilledema and/or folding of the retina and choroid in the posterior pole. | Add | New Term | No current IMDRF code Needed for XEN product Aligns with MedDRA 10081899 Hypotony maculopathy | None | Needed for XEN product. IMDRF lacking coding to accurately capture this issue. Needed specifity for coding and trending. | MODIFY | Accepted by added as level 2 and changed to maculopathy. | E0862 | New Term: Maculopathy New Definition: any pathological condition of the macula of the eye.
78 | 2025-08-22 | AbbVie | E | 2025 | E08xx | Angle Recession | Level 2 | Separation between the longitudinal and circular muscle of the ciliary body (commonly caused by blunt injury to the eye) | Add | New Term | No current IMDRF code and needed to accommodate XEN device. No corresponding MedDRA PT | None | Needed for XEN product. IMDRF lacking coding to accurately capture this issue. Needed specifity for coding and trending. | REJECT | This can be coded under Eye Injury (E0819)
79 | 2025-08-22 | AbbVie | E | 2025 | E08xx | Conjunctival Perforation | Level 2 | A rupture in the conjunctival tissue due to traumatic or pathologic processes. | Add | New Term | No current IMDRF code Needed for XEN product No corresponding MedDRA PT | None | Needed for XEN product. IMDRF lacking coding to accurately capture this issue. Needed specifity for coding and trending. | REJECT | Code as eye injury or new ocular procedure code
80 | 2025-08-22 | AbbVie | E | 2025 | E08xx | Irido-corneal Touch | Level 2 | The Iris and cornea come into contact with each other. | Add | New Term | No current IMDRF code Needed for XEN product No corresponding MedDRA PT | None | Needed for XEN product. IMDRF lacking coding to accurately capture this issue. Needed specifity for coding and trending. | REJECT | This can be coded as Eye Injury (E0819) and Malposition of Device (A150202)
81 | 2025-08-22 | AbbVie | E | 2025 | E08xx | Iridodialysis | Level 2 | Condition where the iris tissue becomes detached from its root. | Add | New Term | No current IMDRF code Needed for XEN product Aligns with MedDRA 10022942 Iridodialysis | None | Needed for XEN product. IMDRF lacking coding to accurately capture this issue. Needed specifity for coding and trending. | REJECT | This can be coded under Eye Injury (E0819)
82 | 2025-08-22 | AbbVie | E | 2025 | E08xx | Iris Atrophy | Level 2 | Iris tissue degeneration/decrease in size and becomes thinner. | Add | New Term | No current IMDRF code Needed for XEN product Aligns with MedDRA 10022948 Iris atrophy | None | Needed for XEN product. IMDRF lacking coding to accurately capture this issue. Needed specifity for coding and trending. | REJECT | This can be coded under Eye Injury (E0819) and Atrophy (E2351)
83 | 2025-08-22 | AbbVie | E | 2025 | E08xx | Iris Rupture | Level 2 | Forcible tearing or disruption of the Iris tissue. | Add | New Term | No current IMDRF code Needed for XEN product No corresponding MedDRA PT | None | Needed for XEN product. IMDRF lacking coding to accurately capture this issue. Needed specifity for coding and trending. | REJECT | This can be coded as Eye Injury (E0819) and Rupture (E2013)
84 | 2025-08-22 | AbbVie | E | 2025 | E08xx | Scleritis | Level 2 | Inflammation of the sclera | Add | New Term | No current IMDRF code Needed for XEN product Aligns with MedDRA 10039705 Scleritis | None | Related to XEN product. Needed for coding specifity and trending. Scleritis or inflammation of sclera. | REJECT | This can be coded under Eye Injury (E0819) and Inflammation (E2326)
85 | 2025-08-22 | AbbVie | E | 2025 | E08xx | Shallow Anterior Chamber | Level 2 | Shallow Anterior chamber | Add | New Term | No current IMDRF code Needed for XEN product Aligns with MedDRA 10040532 Shallow Anterior Chamber | None | Related to XEN product. Needed for coding specifity and trending. | REJECT | This can be coded as Intraocular Pressure Decreased (E0826)
86 | 2025-08-22 | AbbVie | E | 2025 | E08xx | Subconjunctival Bleb Injury | Level 2 | Injury related to the tissue forming the fluid-filled subconjunctival blister (bleb). | Add | New Term | No current IMDRF code. Needed for XEN product and other Level 3 bleb related terms listed below. No corresponding MedDRA PT | None | Related to XEN product. Needed for coding specifity and trending. | MODIFY | In response to the codes for bleb related issues, we have added a new code for 'Occular Procedural Complications' under E2135 | E2135 | Ocular Procedural Complication Any complication arising from ocular surgical procedures involving the conjunctiva or adjacent tissues, including but not limited to subconjunctival bleb formation (e.g., trabeculectomy, glaucoma filtration procedures). This may encompass issues such as bleb absence, infection, leakage, perforation, structural thinning, scarring, or other adverse changes affecting the surgical site or surrounding ocular structures.
87 | 2025-08-22 | AbbVie | E | 2025 | E08xx or E05xx | Retinal Vein Occlusion | Level 2 | Retinal vein occlusion | Add | New Term | No current IMDRF code Needed for XEN product Aligns with MedDRA 10038907 Retinal Vein Occlusion | None | Related to XEN product. Needed for coding specifity and trending. | ACCEPT | Agree to add | E0861 | New Term: Retinal Vein Occlusion New Definition: Retinal vein occlusion
88 | 2025-08-22 | AbbVie | E | 2025 | E08xx or E21xx | Eye Surgical Wound Leak | Level 2 | Pertains to leakage from the surgical wound or paracentesis of the cornea. | Add | New Term | No current IMDRF code Needed for XEN product MedDRA does have many location specific discharge/wound secreation codes. IMDRF does not. This new term is a combination of wound secretion and location secondary to the XEN procedure. | None | Needed for XEN product. IMDRF lacking coding to accurately capture this issue. Needed specifity for coding and trending. The seidel test will be positive. | REJECT | This can be coded as Fluid Discharge or new Occular Procedural Complications'
89 | 2025-08-22 | AbbVie | E | 2025 | E12XX | Thyroiditis | Level 2 | Thyroiditis reflects inflammation of the thyroid gland. | Add | New Term | No current IMDRF code Aligns with MedDRA 10043778 Thyroiditis | None | Associated w/ breast implant compliants. | ACCEPT | Agree to add | E1217 | New Term: Thyroiditis New Definition: Thyroiditis reflects inflammation of the thyroid gland.
90 | 2025-08-22 | AbbVie | E | 2025 | E1402 | Currently - Breast Discomfort/Pain Proposed - Breast Pain | Level 2 | Currently - Having pain, tenderness or discomfort in the breast. Proposed - Having pain or tenderness in the breast. | Modify | Modification of term and definition | Remove "discomfort" from the term name and definition, to better align w/ MedDRA 10006298 Breast pain. | None | Breast Implant complaints of "discomfort - hard to explain" and not necessary "pain". Would like to have ability to utilize code of "discomfort" only and code clear complaints of "pain" separately. | REJECT | Important to keep 'discomfort' included. There is a separate code for discomfort in cases that involve only discomfort. However, when the distinction is unclear and borders between pain, tenderness, and discomfort, this code can be used.
91 | 2025-08-22 | AbbVie | E | 2025 | E14xx | Nipple Discharge | Level 2 | Leakage of fluid from the nipple. | Add | New Term | No current IMDRF code. Needed for breast implants. Aligns with MedDRA 10006223 Breast discharge | None | Pt reports discharge from nipple with breast implants. Usually from an MD. | REJECT | This can be coded as E2315 - Fluid Discharge
92 | 2025-08-22 | AbbVie | E | 2025 | E14xx | Symmastia | Level 2 | Symmastia is a condition where the breasts join together in the middle of the chest, resulting in a lack of cleavage. | Add | New Term | No current IMDRF code Aligns with MedDRA 10073930 Symmastia | None | Usually reported with breast implant shifting, malposition, migration of device. | ACCEPT | Agree to add | E1421 | New Term: Symmastia New Definition: Symmastia is a condition where the breasts join together in the middle of the chest, resulting in a lack of cleavage.
93 | 2025-08-22 | AbbVie | E | 2025 | E1604xx | Scleroderma | Level 3 | Scleroderma is an autoimmune disease that causes inflammation and fibrosis (thickening) in the skin and other areas of the body. | Add | New Term | No current IMDRF code Aligns with MedDRA 10039710 Scleroderma | None | Needed for Breast Implants and currently using "Autoimmune Disorder" for coding and trending. Would like availibility of the code for specificity. | REJECT | To be coded under E0401
94 | 2025-08-22 | AbbVie | E | 2025 | E1604xx | Sjogren's Syndrome | Level 3 | Sjögren’s syndrome, is a chronic (long-lasting) autoimmune disorder that happens when the immune system attacks the glands that make moisture in the eyes, mouth, and other parts of the body. | Add | New Term | No current IMDRF code Aligns with MedDRA 10040767 Sjogren's syndrome | None | Associated w/ breast implant compliants. | REJECT | To be coded under E0401
95 | 2025-08-22 | AbbVie | E | 2025 | E16XX | Chest Wall Deformation/Depression | Level 2 | Ribcage deformity caused by external pressure. | Add | New Term | No current IMDRF code. Needed for large sized breast implants. Partially aligns with MedDRA 10061619 Deformity | None | From larger breast implants, possibly associated with cap con and or sleeping on abdomen. | REJECT | To be coded under E2308
96 | 2025-08-22 | AbbVie | E | 2025 | E1720xx | Acne | Level 3 | Clogged (blackheads or whiteheads), inflammed, or infected sebaceous glands in the skin, characterized by painful pus-filled and/or red papules and pustules which may progress to cysts and nodules | Add | New Term | No current IMDRF E code for acne; (3) separate codes E1719 skin infection, E1720 inflammation/irritation, and E1710 skin lesion Aligns with MedDRA 10000496 Acne | none, more specific term to use in place of the current coding | Acne outbreak post dermal filler injection and code is needed for IMDRF for specificity and trending. | ACCEPT | Agree to add | E1730 | New Term: Acne New Definition: Clogged (blackheads or whiteheads), inflammed, or infected sebaceous glands in the skin, characterized by painful pus-filled and/or red papules and pustules which may progress to cysts and nodules
97 | 2025-08-22 | AbbVie | E | 2025 | E1725 | Current: Injection Site Pooling Proposed: Infusion/Injection Site Swelling | Level 2 | Current: Liquid or a pocket of air trapped under the skin, typically due to injection problems. Does not include inflammation. Proposed: Liquid or a pocket of air trapped under the skin (e.g. pooling), typically due to infusion or injection problems. Does not include inflammation. | Modify | Modify Code Name and Definition from "Injection Site Pooling" to "Infusion/Injection Site Swelling" to encompass both "injection site" and "infusion site" swelling. Modify the definition to clarify swelling may include "pooling". This modification will align closer with the MedDRA mapping. | Aligns with MedDRA 10053505 Infusion Site Swelling Aligns with MedDRA 10053425 Injection Site Swelling | Modification of the existing "Injection Site Pooling", aligning the MedDRA PT code to the IMDRF Code Name and clarification of the IMDRF Definition that this emcompasses "pooling". | Injection Site Pooling naming does not align with the assigned MedRA PT making it difficult to determine the appropriate code to select. This modification will provide clarification. | MODIFY | Combine terms as proposed but amend to align with MedDRA. |  | New Term: Administration Site Swelling Liquid or a pocket of air trapped under the skin (e.g. pooling), typically due to infusion or injection problems. Does not include inflammation.
98 | 2025-08-22 | AbbVie | E | 2025 | E17xx | Induration | Level 2 | Localized thickening or firmness or hardening of soft tissue of the body. | Add | New Term | Currenting using E040203 Local Reaction Code since no existing code for Induration. Aligns with MedDRA 10060708 Induration | None | Reported Symptoms of dermal fillers and coolsculpting. | REJECT | This can be coded by E040203. The term induration is included in definition
99 | 2025-08-22 | AbbVie | E | 2025 | E2107 | Failure of Implant | Level 2 | Failure of the Implant (e.g. degredation and normal aging of the implant). | Add | New Term | Was retired in 2025, however this code was used for breast and other implants. Please reinstate this code. | None | Scenerios of breast and other types of implants that fail as a result of degredation and normal aging of implant. | REJECT | Removed in 2025. Change request log 2025_02 when code was retired: "Agreed to retire this term because it is a device problem.
100 | 2025-08-22 | AbbVie | E | 2025 | E210X | Hepatitis Exposure | Level 3 | Exposure to Hepatitis from a patient who is infected or a carrier of Hepatitis, during surgery to place/remove medical device. | Add | New Term | No current IMDRF code Aligns with LLT's under the MedDRA PT: Exposure to Communicable Diseases (10049711) | None | HCP exposed to infectious hepatitis during procedure to place breast implants. | REJECT | This can be coded as F1007 'exposure to contaminated device' and E1102 'Hepatitis'
101 | 2025-08-22 | AbbVie | E | 2025 | E2111 | Current: Injection Site Reaction Proposed: Infusion/Injection Site Reaction | Level 2 | Current: An intense reaction (usually immunologic) developing at the site of injection Proposed: An intense reaction (usually immunologic) developing at the site of infusion or injection | Modify | Modify Code Name and Definition to encompass both "injection site" and "infusion site" reaction | A code for "Infusion Site Reaction" is needed. Either create a new code for "Infusion Site Action" or combine to the existing "Injection Site Reaction". Aligns with MedDRA 10054996 Infusion Site Reaction Aligns with MedDRA 10022095 Injection Site Reaction | None | Will help to differentiate systemic reactions/hypersensitives | MODIFY | Amending term to 'administration site reaction' to incorporate all |  | New Term: Administration Site Reaction An intense reaction (usually immunologic) developing at the site of infusion or injection
102 | 2025-08-22 | AbbVie | E | 2025 | E21xx | Post Surgical Site Edema | Level 2 | Edema at a surgical site is a common and generally normal post-operative response, caused by the body's inflammatory response to tissue trauma. | Add | New Term | No current IMDRF code Needed for breast implants and other products MedDRA has several options: Post procedural Edema Incision Site Edema | None | Needed for Breast Implants for coding and trending specifity. | REJECT | This can be coded as E2115 - post operative wound infection and general E2338 - Swelling/ Edema.
103 | 2025-08-22 | AbbVie | E | 2025 | E21xx | Subcutaneous Emphysema | Level 2 | Condition when air gets into the tissues under the skin and in the soft tissues. | Add | New Term | No current IMDRF code Needed for breast implants products. Possibly surgery complication or result of an infectious process. Aligns with MedDRA 10042344 Subcutaneous emphysema | None | Needed for Breast Implants and currently using "Skin Disorder" for coding and trending. Would like availibility of the code for specificity. | DUPLICATE | Duplicate of row 16
104 | 2025-08-22 | AbbVie | E | 2025 | E21xx | Ocular Device Exposure | Level 2 | Device exposure through the conjunctiva. | Add | New Term | No current IMDRF code Needed for XEN product No corresponding MedDRA PT | None | Related to XEN product. Needed for coding specifity and trending. | REJECT | Migration of the device would be a under device problem codes. This is not appropriate as an E code
105 | 2025-08-22 | AbbVie | E | 2025 | E21xx | Infusion/Injection site erythema | Level 2 | Erythema at the infusion or injection site. | Add | New Term | No current code Aligns with MedDRA 10048634 Infusion Site Erythema Aligns with MedDRA 10022061 Injection Site Erythema | None | Use for scenerio where the erythema is related to the infusion/injection site reaction. | REJECT | Use generic E171601 - Erythema.
106 | 2025-08-22 | AbbVie | E | 2025 | E21xx | Infusion/Injection Site Pain | Level 2 | Pain at the infusion or injection site. | Add | New Term | No current code Aligns with MedDRA 10053483 Infusion Site Pain Aligns with MedDRA 10022086 Injection Site Pain | None | Use for scenerio where the pain is related to the infusion/injection site reaction. | MODIFY | Amending term to 'administration site reaction' to incorporate all | E2134 | New Term: Administration Site Pain
107 | 2025-08-22 | AbbVie | E | 2025 | E21xx | Medical Device Site Reaction | Level 2 | An intense reaction (usually immunologic) developing at the site of the Medical Device | Add | New Term | Needed to differentiate a systemic reaction to a localized Medical Device Site reaction Aligns with MeddRA 10062680 Medical Device Site Reaction | None | Will help to differentiate systemic reactions/hypersensitives from a reaction that is localized. | REJECT | This can be coded as E040203 Local Reaction
108 | 2025-08-22 | AbbVie | E | 2025 | E2311 | Discomfort | Level 2 | Current: A feeling of mental or physical uneasiness, pain, or distress. Propose: A negative physical and/or emotional state causing unpleasant feelings or sensations. | Modify | Modification of definition | Pain has its own code, distress has its own code and uneasiness is not a harm. Proposed definition: https://pmc.ncbi.nlm.nih.gov/articles/PMC8944041/ | None | Breast Implant complaints of "discomfort - hard to explain" and not necessary "pain". Would like to have ability to utilize code of "discomfort" only. | REJECT | Important to keep 'discomfort' included. There is a separate code for discomfort in cases that involve only discomfort. However, when the distinction is unclear and borders between pain, tenderness, and discomfort, this code can be used.
109 | 2025-08-22 | AbbVie | E | 2025 | E23XX | Flu-Like Symptoms | Level 2 | A group of symptoms that are similar to those caused by the influenza (flu) virus. These include fever, chills, headache, muscle or body aches, cough, sore throat, runny nose, fatigue, nausea, vomiting, and diarrhea. | Add | New Term | No current IMDRF code Aligns with MedDRA 10022004 Influenza like illness | None | Procine and Revolve products are monitored for potential pyrogenic reactions. | REJECT | It is recommended to code the individual symptoms where necessary.
110 | 2025-08-28 | DITTA/GMTA | E | 2025 | E0522 | Coronary Spasm | Level 3 | Contraction of muscular walls of blood vessels | Add | Add new harm as a Level 3 term for "Coronary Spasm" under the Level 2 term "Vasospasm" | Provides a more specific term for coronary spasms | None | When applying pulsed field ablation energy near a coronary artery, such as during CTI line or mitral isthmus line, reports of coronary artery spasm have occurred. | REJECT | The definition provided is covered by the term Vasospasm
111 | 2025-08-28 | DITTA/GMTA | E | 2025 | E0107 | Cognitive Changes | Level 2 | Changes in perception, thinking, or remembering. | Modify | Modification of term and definition. Split cognitive changes into two codes cognitive changes (sensory) and cognitive changes (thinking and memory) | Split cognitive changes into two codes to support more granular post market trending. | None | Sensory: Changes in taste, smell Thinking and memory: Memory loss | MODIFY | Modified to add a new code for olfactory changes. All other issues provided in definition/example already have codes. | E0145 | NewTerm: Olfactory Changes New Definition: Conditions characterized by an alteration in the function or perception of smell, including diminished, distorted, or lost olfactory sensation.
112 | 2025-08-28 | DITTA/GMTA | E | 2025 | E0122 | Dystonia | Level 3 | Movement disorder characterized by involuntary muscle contractions. | Add | Add new Level 3 term of "Dystonia" under the Level 2 term "Movement Disorder" | Provides more general term for events in which patients with dystonia experience a return, exacerbation, or worsening of the disorder without detailing the specific symptoms (e.g., spasm, dyskinesia). | None | A patient being treated with DBS for dystonia experiences an exacerbation of underlying dystonia (e.g., dystonic crisis, dystonic storm) | ACCEPT | Agree to add | E012210 | New Term: Dystonia New Definition: Movement disorder characterized by involuntary muscle contractions.
113 | 2025-08-28 | DITTA/GMTA | E | 2025 | E0122 | Eyelid Apraxia | Level 3 | Difficulty initiating or maintaining eyelid elevation. | Add | Add new Level 3 term of "Eyelid Apraxia" under the Level 2 term "Movement Disorder" | Provides a more specific term for instances in which there is difficulty initiating or maintaining eyelid elevation due to miscoordination of the muscles responsible for opening the eyelid | None | Rare, but potential complication of Deep Brain Stimulation. | MODIFY | Agree to add Apraxia | E012211 | New Term: Apraxia New Definition: inability to perform particular purposive actions
114 | 2025-08-28 | DITTA/GMTA | F | 2025 | F0101 | Therapeutic Response Decreased | Level 2 | A reduction in the desirable and beneficial effects resulting from a medical treatment. | Modify | A reduction or loss in the desirable and beneficial effects resulting from a medical treatment. | Include loss of response in addition to reduction in therapeutic response | None | The reported therapeutic response or benefit to medical treatment was decreased or absent. | MODIFY | Agreed to modify term with slight editorial amendment |  | New definition: A reduction in or complete loss of the desirable and beneficial effects resulting from a medical treatment.
115 | 2025-08-28 | DITTA/GMTA | F | 2025 | TBD | Conversion from Robotic-Assisted Procedure to Laparoscopic Approach | N/A | A surgical procedure initially intended to be completed using robotic assistance that is converted intraoperatively to a conventional laparoscopic approach. This change may result from robotic system limitations, technical issues, or surgeon preference to complete the procedure safely and effectively. Clarifying Notes (optional): This term applies regardless of whether the conversion results in patient harm. Excludes conversions to open surgery, which are captured under existing terms. Intended to enable differentiation of robotic-specific workflow disruptions from other modified surgical procedures. | Add | Adding a code to capture robotic to laparoscopic procedural conversions | This type of health effect/health impact is currently captured under F1906 Modified Surgical Procedure; however, many other scenarios are also captured under this F code. Capturing robotic conversions to laparoscopic is extremely important for our robotic system, and having a designated code would allow for more precise trending, clearer reporting, and differentiation of this specific event from other surgical modifications. It would also support risk analysis, performance monitoring, and targeted mitigation strategies specific to robotic workflows. | Introducing a new code specific to robotic-to-laparoscopic conversion would not negate or replace F1906 but would help disaggregate data currently grouped under it. This refinement would enable more accurate classification of surgical modifications, reduce ambiguity, and improve consistency in coding practices. It would also minimize misclassification of robotic-specific events, ensuring existing terms retain their broader applicability for other types of procedural changes. | During a robotic-assisted laparoscopic cholecystectomy, the procedure was converted to a standard laparoscopic approach due to a technical issue with the robotic system (e.g., loss of instrument articulation). The conversion was necessary to complete the surgery safely and resulted in no additional patient harm. Under current coding, this event would fall under F1906 Modified Surgical Procedure; however, the specific nature of the conversion (robotic to laparoscopic) is critical for post-market trending and analysis of robotic system performance and would be more accurately captured with the proposed term. | MODIFY | Agreed to accept term as this is important to capture. Modified request to remove reference to laproscopic procedure and make more general to incorporate other procedures. | F190701 | New Term: Conversion from Robotic-Assisted Procedure to Other Surgical Approach New Definition: A surgical procedure initially intended to be completed using robotic assistance that is converted intraoperatively to a different surgical approach. This change may result from robotic system limitations, technical issues, or surgeon preference to complete the procedure safely and effectively.
116 | 2025-08-28 | DITTA/GMTA | A | 2025 | new | AI/ML Problem | Level 2 | Problem associated with an AI/ML software failure | Add | Adding new Level 2 category for AI/ML related failure modes. Recommend to add under A11. | Category at level 2 is needed for AI/ML related problems | N/A | Inaccurate or unintended AI/ML actions, bias, decisions, or guidance | MODIFY | Example not provided so this issue was difficult to assess. We have added "AI/ML" to the definition of A11 to assist with this issue. |  | New Definition (A11): Problem associated with written programs, codes, and/or software system (including AI/ML) that affects device performance or communication with another device
117 | 2025-08-28 | DITTA/GMTA | A | 2025 | A1105 | Computer System Security Problem | Level 2 | Problem associated with a software system (including an AI/ML system) resulting in a loss of confidentiality, integrity, or availability of written program code, application software, or data or entire device (e.g. unauthorized access to or modification, data leakage) | Modify | Modification of definition to include AI/ML specific terms | This change is necessary to include considerations for AI/ML | A1105 | Examples: Data leakage, adversarial attack, data poisoning | REJECT | Covered by A1105
118 | 2025-08-28 | DITTA/GMTA | D | 2025 | D0107 | Inadquate Cybersecurity | Level 2 | Problems traced to inadequate protection of computers, servers, mobile devices, electronic systems, networks, programs, and data from malicious attacks, damage, data leakage, or unauthorized access. | Modify | Modification of definition to include AI/ML specific terms | This change is necessary to include considerations for AI/ML | D0107 | Example: The AI inadvertently reproduces sensitive, proprietary, or private information seen during training. | ACCEPT |  |  | New Definition: Problems traced to inadequate protection of computers, servers, mobile devices, electronic systems, networks, programs, and data from malicious attacks, damage, data leakage, or unauthorized access.
119 | 2025-08-28 | DITTA/GMTA | D | 2025 | new | Inadequate Machine Learning/AI | Level 3 | Problems traced to inadequate AI/ML | Add | Add new Level 3 under D0108 (Inadequate Software Design) | The change is necessary to properly account for investigation conclusion that is specific to AI/ML software systems | NA | Examples: Overfitting, Overtrust, Domain Shift, Reward Hacking, Catastrophic Forgetting, Agentic AI Failure Propagation, Prompt Injection and Jailbreaking, Mode Collapse | MODIFY | Agreed this would add value. Instead of adding a new level 3 code, it was decided to add AI/ML to the definition of D0108 |  | Cause traced to inadequate or inappropriate design of the software (including AI/ML) component of a medical device, or Software as a Medical Device.
120 | 2025-08-28 | DITTA/GMTA | A | 2025 | A072204 | High Stimulation Impedance | Level 3 | Problem associated with higher than intended stimulation impedance levels | Add | add the term 'high stimulation impedance' | The current IMDRF coding system includes only a single, undifferentiated category for impedance-related issues, which lacks the specificity needed to reflect the diverse clinical and technical contexts in which impedance is monitored across medical devices. Impedance abnormalities, whether in stimulation, sensing, or therapeutic delivery, can indicate distinct underlying issues such as lead malfunction, tissue interface degradation, or system failure, each with different risk profiles and clinical implications. The Introduction of six new, specific impedance-related codes would enable more precise documentation and analysis of device-related events, thereby improving the overall quality of incident reporting. This enhancement would also align the IMDRF coding system with current clinical practice, where distinct impedance types are routinely separately ultimately enhancing patient safety and ensuring more appropriate regulatory evaluation. | Additional more specific term to high impedance A072201 | Pacemaker lead shows high out of range pacing impedance (>2500 ohms). OR Tachycardia lead shows increase of the pacing impedance. OR CRTD lead shows increased pacing impedance | REJECT | This information would become a level 4 code. The manufacturer can add level 4 codes to their own system if this level of granularity is required.
121 | 2025-08-28 | DITTA/GMTA | A | 2025 | A072205 | Low Stimulation Impedance | Level 3 | Problem associated with lower than intended stimulation impedance levels | Add | add the term 'low stimulation impedance' | The current IMDRF coding system includes only a single, undifferentiated category for impedance-related issues, which lacks the specificity needed to reflect the diverse clinical and technical contexts in which impedance is monitored across medical devices. Impedance abnormalities, whether in stimulation, sensing, or therapeutic delivery, can indicate distinct underlying issues such as lead malfunction, tissue interface degradation, or system failure, each with different risk profiles and clinical implications. The Introduction of six new, specific impedance-related codes would enable more precise documentation and analysis of device-related events, thereby improving the overall quality of incident reporting. This enhancement would also align the IMDRF coding system with current clinical practice, where distinct impedance types are routinely separately ultimately enhancing patient safety and ensuring more appropriate regulatory evaluation. | Additional more specific term to low impedance A072202 | Pacemaker lead shows low out of range pacing impedance (<200 ohms). OR Tachycardia lead shows decrease of the pacing impedance. OR CRTD lead shows decreased pacing impedance | REJECT | This information would become a level 4 code. The manufacturer can add level 4 codes to their own system if this level of granularity is required.
122 | 2025-08-28 | DITTA/GMTA | A | 2025 | A072206 | Unstable Stimulation Impedance | Level 3 | The stimulation impedance varies without being consistently too high or too low. | Add | add the term 'unstable stimulation impedance' | The current IMDRF coding system includes only a single, undifferentiated category for impedance-related issues, which lacks the specificity needed to reflect the diverse clinical and technical contexts in which impedance is monitored across medical devices. Impedance abnormalities, whether in stimulation, sensing, or therapeutic delivery, can indicate distinct underlying issues such as lead malfunction, tissue interface degradation, or system failure, each with different risk profiles and clinical implications. The Introduction of six new, specific impedance-related codes would enable more precise documentation and analysis of device-related events, thereby improving the overall quality of incident reporting. This enhancement would also align the IMDRF coding system with current clinical practice, where distinct impedance types are routinely separately ultimately enhancing patient safety and ensuring more appropriate regulatory evaluation. | Additional more specific term to unstable impedance A072203 | Pacemaker lead shows fluctuating pacing impedance. OR Tachycardia lead shows fluctuation of the pacing impedance. OR CRTD lead shows fluctuating pacing impedance | REJECT | This information would become a level 4 code. The manufacturer can add level 4 codes to their own system if this level of granularity is required.
123 | 2025-08-28 | DITTA/GMTA | A | 2025 | A072207 | High Shock Impedance | Level 3 | Problem associated with higher than intended shock impedance levels | Add | add the term 'high shock impedance' | The current IMDRF coding system includes only a single, undifferentiated category for impedance-related issues, which lacks the specificity needed to reflect the diverse clinical and technical contexts in which impedance is monitored across medical devices. Impedance abnormalities, whether in stimulation, sensing, or therapeutic delivery, can indicate distinct underlying issues such as lead malfunction, tissue interface degradation, or system failure, each with different risk profiles and clinical implications. The Introduction of six new, specific impedance-related codes would enable more precise documentation and analysis of device-related events, thereby improving the overall quality of incident reporting. This enhancement would also align the IMDRF coding system with current clinical practice, where distinct impedance types are routinely separately ultimately enhancing patient safety and ensuring more appropriate regulatory evaluation. | Additional more specific term to high impedance A072201 | ICD lead shows high out of range shock impedance (>150 ohms). | REJECT | This information would become a level 4 code. The manufacturer can add level 4 codes to their own system if this level of granularity is required.
124 | 2025-08-28 | DITTA/GMTA | A | 2025 | A072208 | Low Shock Impedance | Level 3 | Problem associated with lower than intended shock impedance levels | Add | add the term 'low shock impedance' | The current IMDRF coding system includes only a single, undifferentiated category for impedance-related issues, which lacks the specificity needed to reflect the diverse clinical and technical contexts in which impedance is monitored across medical devices. Impedance abnormalities, whether in stimulation, sensing, or therapeutic delivery, can indicate distinct underlying issues such as lead malfunction, tissue interface degradation, or system failure, each with different risk profiles and clinical implications. The Introduction of six new, specific impedance-related codes would enable more precise documentation and analysis of device-related events, thereby improving the overall quality of incident reporting. This enhancement would also align the IMDRF coding system with current clinical practice, where distinct impedance types are routinely separately ultimately enhancing patient safety and ensuring more appropriate regulatory evaluation. | Additional more specific term to low impedance A072202 | ICD lead shows low out of range shock impedance (<25 ohms). | REJECT | This information would become a level 4 code. The manufacturer can add level 4 codes to their own system if this level of granularity is required.
125 | 2025-08-28 | DITTA/GMTA | A | 2025 | A072209 | Unstable Shock Impedance | Level 3 | The shock impedance varies without being consistently too high or too low. | Add | add the term 'unstable shock impedance' | The current IMDRF coding system includes only a single, undifferentiated category for impedance-related issues, which lacks the specificity needed to reflect the diverse clinical and technical contexts in which impedance is monitored across medical devices. Impedance abnormalities, whether in stimulation, sensing, or therapeutic delivery, can indicate distinct underlying issues such as lead malfunction, tissue interface degradation, or system failure, each with different risk profiles and clinical implications. The Introduction of six new, specific impedance-related codes would enable more precise documentation and analysis of device-related events, thereby improving the overall quality of incident reporting. This enhancement would also align the IMDRF coding system with current clinical practice, where distinct impedance types are routinely separately ultimately enhancing patient safety and ensuring more appropriate regulatory evaluation. | Additional more specific term to unstable impedance A072203 | ICD lead shows fluctuating shock impedance. | REJECT | This information would become a level 4 code. The manufacturer can add level 4 codes to their own system if this level of granularity is required.
126 | 2025-08-28 | DITTA/GMTA | A | 2025 | A140506 | Intermittent Infusion | Level 3 | Problem associated with the infusion not being stead- characterised by intermittent stoppages to the flow. | Modify | modification of term to include "flow" : "Intermittent Infusion or flow", and to the definition to make it widder than infusion problems only "Problem associated with the flow not being stead- characterised by intermittent stoppages to the flow." | This code is under A1405 "Improper Flow or Infusion" but it is referencing infusion only. There is no code for intermittent flow that are not related to infusion (for example intermittent irrigation). | no impact | For an ultrasonic surgical aspirator system that has 3 functions: fragmentation of the tissue, irrigation and aspiration (suction of the irrigation fluid and fragmented tisssue), the user reported that the irrigation was working intermittently. | ACCEPT | Accepted and slight editorial amendment to change stead- to steady |  | New Term: Intermittent Infusion or Flow New Definition: Problem associated with the flow not being steady characterised by intermittent stoppages to the flow.
127 | 2025-08-28 | DITTA/GMTA | A | 2025 | A210102 | Illegible Information | Level 3 | Problem associated with information unable to be read or deciphered | Modify | Change term to: Illegible or Faded Information. Change definition to: Problem associated with information difficult or impossible to be read or deciphered. | The level 2 term, A2101 - Device Markings/Labelling Problem, includes in its definition a sentence related to markings/labelings 'worn out'. A2101 reads: "... Problems can include but are not limited to this material being unclear, missing, worn out". The level 3 term, A210102 – Illegible Information, and its definition, does not allow for difficulties to read the markings when it comes to the markings being worn out. In consequence, the proposal includes that this definition is broaden so it can also accommodate for markings difficult to be read. | None | The laser lines on the instruments are worn and difficult to see clearly. They appear faded to a gray color from the original black color. | MODIFY | Amend term to include or difficult to read. Amend per definition |  | New Term: Illegible or Difficult to Read Information
128 | 2025-08-28 | DITTA/GMTA | A | 2025 | New level 3 code under A0205 | Product Mix-Up Incorrect product (note: avoid terms that may not translate - mix-up for example) | 3rd level | Problem associated with a device whose identity does not match to what is written on the label. | Add | Add new level 3 sub-code under A0205. Consider updating language under A210106 - Wrong Label since the symptom would be the part did not match the label but the reason would not be known (label issue or product mispackaing) | There is no suitable code/term available: Product mix-ups can be caused by multiple factors: incorrect product was introduced in the correct package; incorrect label was affixed to the correct product; or incorrect manufacturing programming was selected, rendering an incorrect product. | None | Product A was ordered. The product received, showed it was product A; however, when opened, the product's identity was B. | REJECT | A210106 can be used.
129 | 2025-08-28 | DITTA/GMTA | A | 2025 | New level 3 code under A0406 | Thinning of Material | 3rd level | Device is thinned due to deformation. | Add | Add new level 3 sub-code under A0406, Term: 'Thinning of Material', Description 'Device is thinned due to deformation. | There is a code for Thickening of Material (A040612). (Alternatively, modify A040611 - Stretched to include thinning of Material). | None | Addition only for consistency. | MODIFY | Agree with alternative proposal, being - modify A040611 - Stretched to include thinning of Material. |  | New Definition: Problem associated with an increase, elongation, or thinning in a materials' dimension
130 | 2025-08-28 | DITTA/GMTA | B | 2025 | B | Risk Assessment Review | Level 1 | The investigation involved the review and analysis of risk files of the device to support the identification of possible causes for the adverse event and assess whether the occurrence and severity are within anticipated parameters. It should be noted that risk assessment review typically is not considered sufficient as a stand-alone method, but should be used in conjunction with other investigation methods for providing for instance complementary information. | Add | Addition, see left. | ISO 13485 integrates a risk-based approach throughout the entire lifecycle of medical devices, including post-market activities. This standard mandates a proactive process of identifying, evaluating, and controlling risks to ensure patient safety. The change is necessary to accommodate an investigation task that is constantly used to ensure the medical device is safe and effective. MIR form 7.3.1 section 4.2 – ‘Cause of investigation and conclusion’ asks for risk assessment review, we would like to propose coding all types investigations with Annex B instead of in other sections of the MIR form. | None | Risk management review is used in investigation of quality complaints and adverse events to identify whether the event is already included in the risk files. If it was not, it drives its inclusion; if it was, an assessment is conducted to determine whether the severity, occurrence rate and probability of harm are within the anticipated levels; if it is not, the manufacturer takes action. | REJECT | Risk assessment review should remain documented as part of CAPA or risk management processes, not as an Annex B investigation code.
131 | 2025-08-28 | DITTA/GMTA | B | 2025 | B | Medical/Clinical Review | Level 1 | The investigation involved the analysis of the medical documentation (radiological images, lab tests, medical records, surgical notes, pathology reports, etc.) by a healthcare professional to support the identification of possible causes for the adverse event. It should be noted that medical/clinical review typically is not considered sufficient as a stand-alone method, but should be used in conjunction with other investigation methods for providing for instance complementary information. | Add | Addition, see left. | This proposed term is similar to B15; however, we are proposing that there is a separate code for a clinical/medical investigation because it must be conducted by a trained healthcare professional. The medical investigation is necessary to use D1001 - Adverse Event Related to Patient Anatomy and/or Conditions. This assessment does not fit under B15, as we would like to specifically call out a review by a healthcare professional. Of note, US/AU regulations include clauses/exceptions for not reporting certain adverse events based on medical judgement from a healthcare professional. | None | The code will be used in the investigation of any adverse event where there is patient harm and the user provides medical/clinical information (images, reports, tests, analysis, notes) for the manufacturer to assess so a root cause can be established. | REJECT | This scenario is covered by B15.
132 | 2025-08-28 | DITTA/GMTA | C | 2025 | New child of C0706 | Stress Problem Identified | Level 2 | Problems caused by either excessive or inadequate physical force exerted on it by another object resulting in problems e.g. wear, bending, deformation, fracture, fatigue. | Add | Add a level 3 sub-code under C0706 for additional codes to address type of endoleak - Type IIIa | Addition of new terms allows for identification of specific types of endoleak which will better define location, cause and impact to the patient. | None | Endovascular aneurysm repair (EVAR) procedures frequently involve the use of stent grafts featuring multiple connections and branches to various arteries. Accurate and detailed classification of endoleak types is essential to enhance the monitoring, reporting, and management of adverse events/serious incidents associated with these devices. Type IIIa endoleak is a graft material or modular failure with component disconnection [e.g., modular segments] at the attachment of aortic-aortic or aortic-bifurcated component. Implementing a dedicated C-code structure that distinctly captures each endoleak subtype would substantially improve the precision of device-related event tracking. This enhancement would support more informed clinical decision-making, strengthen regulatory oversight, and ultimately contribute to improved patient outcomes and more robust post-market surveillance. | REJECT | The existing combination of A050408 – Device Endoleak and C0706 – Stress Problem Identified already covers this situation.
133 | 2025-08-28 | DITTA/GMTA | C | 2025 | New child of C0706 | Stress Problem Identified | Level 2 | Problems caused by either excessive or inadequate physical force exerted on it by another object resulting in problems e.g. wear, bending, deformation, fracture, fatigue. | Add | Add a level 3 sub-code under C0706 for additional codes to address type of endoleak - Type IIIb | Addition of new terms allows for identification of specific types of endoleak which will better define location, cause and impact to the patient. | None | In endovascular aneurysm repair (EVAR) procedures, the use of complex stent grafts with multiple branches and arterial connections is common. To improve the monitoring and reporting of adverse events/serious incidents, it is essential to further refine the classification of endoleak types. The Oderich Classification (Expanded/Modern System) provides a more granular categorization of endoleaks. In that classification, a Type IIIb endoleak is a graft material or modular failure with component disconnection [e.g., modular segments] at the bifurcated iliac limb or iliac limb-iliac limb component. Implementing a dedicated C-code structure that distinctly captures each endoleak subtype would substantially improve the precision of device-related event tracking. This enhancement would support more informed clinical decision-making, strengthen regulatory oversight, and ultimately contribute to improved patient outcomes and more robust post-market surveillance. | REJECT | The existing combination of A050408 – Device Endoleak and C0706 – Stress Problem Identified already covers this situation.
134 | 2025-08-28 | DITTA/GMTA | C | 2025 | New child of C0706 | Stress Problem Identified | Level 2 | Problems caused by either excessive or inadequate physical force exerted on it by another object resulting in problems e.g. wear, bending, deformation, fracture, fatigue. | Add | Add a level 3 sub-code under E2121 for additional codes to address types of endoleak - Type IIIc | Addition of new terms allows for identification of specific types of endoleak which will better define location, cause and impact to the patient. | None | Endovascular aneurysm repair (EVAR) procedures frequently involve the use of stent grafts featuring multiple connections and branches to various arteries. Accurate and detailed classification of endoleak types is essential to enhance the monitoring, reporting, and management of adverse events/serious incidents associated with these devices. Type IIIc endoleak is an graft material or modular failure with seperation or disconnection of a branch component [e.g., bridging stent graft] Implementing a dedicated C-code structure that distinctly captures each endoleak subtype would substantially improve the precision of device-related event tracking. This enhancement would support more informed clinical decision-making, strengthen regulatory oversight, and ultimately contribute to improved patient outcomes and more robust post-market surveillance. | REJECT | The existing combination of A050408 – Device Endoleak and C0706 – Stress Problem Identified already covers this situation.
135 | 2025-08-28 | DITTA/GMTA | C | 2025 | New child of C0706 | Stress Problem Identified | Level 2 | Problems caused by either excessive or inadequate physical force exerted on it by another object resulting in problems e.g. wear, bending, deformation, fracture, fatigue. | Add | Add a level 3 sub-code under C0706 for additional codes to address type of endoleak - Type IIId | Addition of new terms allows for identification of specific types of endoleak which will better define location, cause and impact to the patient. | None | In endovascular aneurysm repair (EVAR) procedures, the use of complex stent grafts with multiple branches and arterial connections is common. To improve the monitoring and reporting of adverse events/serious incidents, it is essential to further refine the classification of endoleak types. The Oderich Classification (Expanded/Modern System) provides a more granular categorization of endoleaks. In that classification, a Type IIId endoleak specifically involves structural failures of the stent graft [tear, perforation or fracture of the side branch] Implementing a dedicated C-code structure that distinctly captures each endoleak subtype would substantially improve the precision of device-related event tracking. This enhancement would support more informed clinical decision-making, strengthen regulatory oversight, and ultimately contribute to improved patient outcomes and more robust post-market surveillance. | REJECT | The existing combination of A050408 – Device Endoleak and C0706 – Stress Problem Identified already covers this situation.
136 | 2025-08-28 | DITTA/GMTA | C | 2025 | C0602 | Inappropriate Material | 2nd level | Problems that occur due to the presence of a material that should not be present or part of the device. | Modify | Modify definition of C0602 to: "Problems that occur due to undesired presence, composition or structure of a material." | Child terms of C0602 do mention material structure and composition and presence of undesired material. The proposed change to the parent term derscription just includes all those options. | None | N/A | ACCEPT |  |  | Problems that occur due to undesired presence, composition or structure of a material.
137 | 2025-08-28 | DITTA/GMTA | C | 2025 | C160503 | Packaging Contains Unintended Material | 3rd level | Problems that occurred because unintended material was packaged with the device. | Modify | Modify definition: Unintended material was packaged with the device. | Remove word "problem", as not all such investigation findings are considered a problem of the product (i.e. when it meets specifications) | None | Particulate in a sterile package is reported however this meets the specifications defined by the manufacturer. | REJECT | Annex C uses “problem” wording consistently because these codes reflect investigation findings that explain why the device issue occurred. If no problem is identified,
138 | 2025-08-28 | DITTA/GMTA | D | 2025 | New child of D14 | Product meets specifications | 2nd level | The reported event is substantiated but determined to not be a deficiency of the product. | Add | Add new level 2 sub-code under D14: 'Product Meets Specifications', definition: 'The reported event is substantiated but determined to not be a deficiency of the product.' | Include a new code to address situations when a reported event is observed, however it is not a related to a product deficiency. | None | Particulate in a sterile package is reported however this met the specifications defined by the manufacturer. | REJECT | Captured by new C1902
139 | 2025-08-28 | DITTA/GMTA | E | 2025 | New child of E2121 | Endoleak | Level 2 | The observed patient condition is persistent blood flow/pressure in the aneurysm sac following an endovascular aneurysm repair (EVAR) procedure. Where other patient signs and symptoms indicate the presence of an endoleak those signs and symptoms should be coded. | Add | Add a level 3 sub-code under E2121 for additional codes to address types of endoleak - Type Ia [MedDRA: 10063824: Endoleak] | Addition of new terms allows for identification of specific types of endoleak which will better define location, cause and impact to the patient. | None | Endovascular aneurysm repair (EVAR) procedures frequently involve the use of stent grafts featuring multiple connections and branches to various arteries. Accurate and detailed classification of endoleak types is essential to enhance the monitoring, reporting, and management of adverse events/serious incidents associated with these devices. A Type Ia endoleak is an inadequate seal at the endograft proximal [main graft or fenestration seal] attachment. Implementing a dedicated E-code structure that distinctly captures each endoleak subtype would substantially improve the precision of adverse event/serious incident tracking. This enhancement would support more informed clinical decision-making, strengthen regulatory oversight, and ultimately contribute to improved patient outcomes and more robust post-market surveillance. | REJECT | This can be coded as Endoleak. A manufacturer can add more detailed terms in their own system, as required.
140 | 2025-08-28 | DITTA/GMTA | E | 2025 | New child of E2121 | Endoleak | Level 2 | The observed patient condition is persistent blood flow/pressure in the aneurysm sac following an endovascular aneurysm repair (EVAR) procedure. Where other patient signs and symptoms indicate the presence of an endoleak those signs and symptoms should be coded. | Add | Add a level 3 sub-code under E2121 for additional codes to address types of endoleak -Type Ib [MedDRA: 10063824: Endoleak] | Addition of new terms allows for identification of specific types of endoleak which will better define location, cause and impact to the patient. | None | Endovascular aneurysm repair (EVAR) procedures frequently involve the use of stent grafts featuring multiple connections and branches to various arteries. Accurate and detailed classification of endoleak types is essential to enhance the monitoring, reporting, and management of adverse events/serious incidents associated with these devices. Type Ib endoleak is an inadequate seal at the endograft distal [e.g., iliac limb or branch] attachment. Implementing a dedicated E-code structure that distinctly captures each endoleak subtype would substantially improve the precision of adverse event/serious incident tracking. This enhancement would support more informed clinical decision-making, strengthen regulatory oversight, and ultimately contribute to improved patient outcomes and more robust post-market surveillance. | REJECT | This can be coded as Endoleak. A manufacturer can add more detailed terms in their own system, as required.
141 | 2025-08-28 | DITTA/GMTA | E | 2025 | New child of E2121 | Endoleak | Level 2 | The observed patient condition is persistent blood flow/pressure in the aneurysm sac following an endovascular aneurysm repair (EVAR) procedure. Where other patient signs and symptoms indicate the presence of an endoleak those signs and symptoms should be coded. | Add | Add a level 3 sub-code under E2121 for additional codes to address types of endoleak - Type Ic [MedDRA: 10063824: Endoleak] | Addition of new terms allows for identification of specific types of endoleak which will better define location, cause and impact to the patient. | None | Endovascular aneurysm repair (EVAR) procedures frequently involve the use of stent grafts featuring multiple connections and branches to various arteries. Accurate and detailed classification of endoleak types is essential to enhance the monitoring, reporting, and management of adverse events/serious incidents associated with these devices. Type Ic endoleak is an inadequate seal at the endograft branch vessel origin in branched devices [e.g., SMA, renal] Implementing a dedicated E-code structure that distinctly captures each endoleak subtype would substantially improve the precision of adverse event/serious incident tracking. This enhancement would support more informed clinical decision-making, strengthen regulatory oversight, and ultimately contribute to improved patient outcomes and more robust post-market surveillance. | REJECT | This can be coded as Endoleak. A manufacturer can add more detailed terms in their own system, as required.
142 | 2025-08-28 | DITTA/GMTA | E | 2025 | New child of E2121 | Endoleak | Level 2 | The observed patient condition is persistent blood flow/pressure in the aneurysm sac following an endovascular aneurysm repair (EVAR) procedure. Where other patient signs and symptoms indicate the presence of an endoleak those signs and symptoms should be coded. | Add | Add a level 3 sub-code under E2121 for additional codes to address types of endoleak - Type II [MedDRA: 10063824: Endoleak] | Addition of new terms allows for identification of specific types of endoleak which will better define location, cause and impact to the patient. | None | Endovascular aneurysm repair (EVAR) procedures frequently involve the use of stent grafts featuring multiple connections and branches to various arteries. Accurate and detailed classification of endoleak types is essential to enhance the monitoring, reporting, and management of adverse events/serious incidents associated with these devices. Type II endoleak is from retrograde flow from collateral vessels. Implementing a dedicated E-code structure that distinctly captures each endoleak subtype would substantially improve the precision of adverse event/serious incident tracking. This enhancement would support more informed clinical decision-making, strengthen regulatory oversight, and ultimately contribute to improved patient outcomes and more robust post-market surveillance. | REJECT | This can be coded as Endoleak. A manufacturer can add more detailed terms in their own system, as required.
143 | 2025-08-28 | DITTA/GMTA | E | 2025 | E07XX | Hypercapnia | Respiratory System | A condition of abnormally elevated carbon dioxide (CO2) levels in the blood (MEDDRA/10020591) | Add | term added | It was reported to us in conjunction with procedural problems during knee arthroplasty. The symptom could not be coded as it is not covered in Annex E. | None | It was reported to us in conjunction with procedural problems during knee arthroplasty. | REJECT | This can be coded as Pulmonary Disfunction
144 | 2025-08-28 | DITTA/GMTA | E | 2025 | E07XX | Tidal volume decreased | Respiratory System | Using a smaller than usual amount of air (tidal volume) with each breath delivered to a patient on a ventilator (MEDDRA/10078272) | Add | term added | It was reported to us in conjunction with procedural problems during knee arthroplasty. The symptom could not be coded as it is not covered in Annex E. | None | It was reported to us in conjunction with procedural problems during knee arthroplasty. | REJECT | Per the definition, this is representative of an intentional setting on a ventilator rather than the result of a pulmonary function investigation.
145 | 2025-08-28 | DITTA/GMTA | E | 2025 | E1649 | Crepitus | Level 2 | Popping, cracking, or clicking sound in a joint. | Modify | Modification of definition | Current definition "Grating sound or sensation produced by friction between e.g. bone, cartilage, fractured parts of a bone." is from the Oxford dictionary. In the MSD manual ist is described as "popping, cracking, or clicking sound in a joint". Use of "crepitus" in this article https://pmc.ncbi.nlm.nih.gov/articles/PMC5851845/. | None | A clicking noise was reported after knee arthroplasty. | MODIFY | Agreed this is important to amend. Definition amended to incorporate both scenarios |  | New Definition: A grating, popping, cracking, or clicking sound or sensation in a joint, typically produced by friction between structures such as bone, cartilage, or fractured parts of a bone.
146 | 2025-08-28 | DITTA/GMTA | F | 2025 | New level 2 term under F26 | Potential Serious Risk (No Harm Occurred) | Level 2 | No harm occurred in relation to the adverse event. However, the event involved a device malfunction or use condition that, if repeated, could reasonably be expected to result in a serious adverse health impact (e.g., death, serious deterioration of health, or serious public health threat). (For Exposure to Contaminated Device/Risk of Infection use F1007) | Add | Add a level 2 term under F26. | Current IMDRF codes do not adequately capture adverse events that meet the MDR definition of a serious incident due to their potential consequences, even if no harm occurred. This term would fill a gap in reporting serious risks that were averted but remain relevant for vigilance and preventive action. | None | In 2023, a serious incident was reported to us in which a blunt scissor led to a serious complication during labour. A year later, the same customer reported a similar incident with a blunt scissors in which no complications occurred. Nevertheless, the incident was reported as a serious incident because there was a risk when using the product. | REJECT | F26 is sufficient for describing the actual event. Reporting requirements are decided by each regulator for their jurisdiction.
147 | 2025-08-28 | DITTA/GMTA | E | 2025 | E0819xx | Conjunctival retraction | Level 3 | The abnormal pulling back or shrinking of the conjunctiva | Add |  | For ocular devices, this is an adverse event that we see in clinical trails, but the closest current E code is 'E0819-Eye Injury'. Unfortunately this does not describe the adverse event specifically enough for reporting or risk management purposes. It should also be noted that there is a corresponding Meddra code 'Conjunctival retraction-10079837''. | None | 1) Patient presented with conjunctival retraction. The implant is still covered so refill was done. If retraction worsens, will have to consider explantation. 2) Noticed conjunctival retraction upon slit lamp exam by investigator Most commonly related to implant and/or implantation procedure | REJECT | Code as eye injury or new ocular procedure code
148 | 2025-08-28 | DITTA/GMTA | E | 2025 | E08xxxx | Conjunctival Bleb | Level 3 | A conjunctival bleb is a blister-like elevation on the conjunctiva due to escape of fluid under the conjunctiva | Add |  | Currently there are no terms that describe Conjunctival bleb or terms that closely represent this events. Also Conjunctival bleb is frequently reported with use of an ocular implant. It should also be noted that there is a corresponding Meddra Code: Conjunctival bleb- 10052114 |  | 1) Principal investigator had refill needle seated, assumed to be in the implant, he felt resistance and noticed that about 0.01 ml of drug was pushed under the conjunctiva forming a bleb 2) Principal investigator noted a conjunctival bleb overlying the implant during examination. | DUPLICATE | In response to the codes for bleb related issues, we have added a new code for 'Occular Procedural Complications' under E2135
149 | 2025-08-28 | DITTA/GMTA | E | 2025 | E08xxxx | Conjunctival filtering bleb leak | Level 3 | Escape of fluid from a conjunctival bleb | Add |  | Currently there are no terms that describe Conjunctival filtering bleb leak or terms that closely represent this events. Also Conjunctival filtering bleb leak is frequently reported with use of an ocular implant. It should also be noted that there is a corresponding Meddra code: Conjunctival filtering bleb leak- 10069165 |  | 1) Physician who perfomed week 120 refill contacted the medical monitor when he noticed mild leak of conjunctiva overlying implant and was recommended close monitoring for filtering bleb leak. | DUPLICATE | In response to the codes for bleb related issues, we have added a new code for 'Occular Procedural Complications' under E2135
150 | 2025-08-28 | DITTA/GMTA | E | 2025 | E08xxxx | Conjunctival Bleb | Level 3 | A conjunctival bleb is a blister-like elevation on the conjunctiva due to escape of fluid under the conjunctiva | Add |  | For occular devices, this is an adverse event that we see in clinical trails, but the closest current E code is 'E0819-Eye Injury'. Unfortunately this does not describe the adverse event specifically enough for reporting or risk management purposes. It should also be noted that there is a corresponding Meddra code: Conjunctival bleb- 10052114; Conjunctival filtering bleb leak- 10069165 | None | 1) Principal investigator had refill needle seated, assumed to be in the implant, he felt resistance and noticed that about 0.01 ml of drug was pushed under the conjunctiva forming a bleb 2) Principal investigator noted a conjunctival bleb overlying the implant during examination. 3) Physician who perfomed week 120 refill contacted the medical monitor when he noticed mild leak of conjunctiva overlying implant and was recommended close monitoring for filtering bleb leak. | DUPLICATE | In response to the codes for bleb related issues, we have added a new code for 'Occular Procedural Complications' under E2135
151 | 2025-08-28 | DITTA/GMTA | E | 2025 | E08xx or E0819xx | Conjunctival Erosion | Level 2 or 3 | A localized area of damage, thinning, or breakdown of the conjunctival surface | Add |  | For occular devices, this is an adverse event that we see in clinical trails, but the closest current E code is either 'E0819-Eye Injury' or 'E2006-Erosion'. Unfortunately these do not describe the adverse event specifically enough for reporting or risk management purposes. It should also be noted that there is a corresponding Meddra code 'Conjunctival erosion-10010716'. | None | Subject presented for a second unscheduled visit after being diagnosed with a conjunctival bleb post implantation procedure. Sub-investigator determined that the subject had a resolved conjunctival bleb, and diagnosed her with a moderate focal conjunctival avascularity. Per query from system, site is clarifying the diagnosis to a focal conjunctival erosion | REJECT | This Can be coded as Eye Injury (E0819)
152 | 2025-08-28 | DITTA/GMTA | E | 2025 | E08xx | Tenon Capsule Thickening | Level 2 | An abnormal thickening of Tenon’s capsule | Add |  | Currently there are no terms that describe Tenon capsule thickening or terms that closely represent this event. Meddra Code Tenon's capsule thickening- 10087613 | None | The sub-investigator noted tenon thickening during refill attempt that was unsuccessful. Unable to insert refill needle into implant septum. Most commonly related to implant and/or implantation procedure | REJECT | This can be coded as Eye Injury (E0819) and Fibrosis (E2313)
153 | 2025-08-28 | DITTA/GMTA | E | 2025 | E08xx | Subconjunctival Thickening | Level 2 | An abnormal increase in the density or bulk of the tissue beneath the conjunctiva | Add |  | Currently there are no terms that describe subconjunctival thickening or terms that closely represent this event. Currently no specific Meddra Code. Closest is 10061446 Conjunctival disorder | None | Principal investigator pushed the syringe plunger with resistance and no reflux exchange was noticed due to subconjunctivval thickening over implant. | REJECT | This can be coded as Eye Injury (E0819)
154 | 2025-08-28 | DITTA/GMTA | E | 2025 | E0819xx | Scleral thinning | Level 3 | A condition where the sclera, the white, tough outer layer of the eye that provides structural support and protection, becomes abnormally thin | Add |  | For ocular devices, this is an adverse event that we see in clinical trials, but the closest current E code is 'E0819-Eye Injury'. Unfortunately this does not describe the adverse event specifically enough for reporting or risk management purposes. | None | Patient presented with complaints of ocular discomfort, photophobia and mild foreign body sensation in the right eye for 4 days. On examination localized scleral thinning was noted in the superior temporal quadrant of the right eye . | REJECT | This can be coded as Eye Injury (E0819)
155 | 2025-08-28 | DITTA/GMTA | E | 2025 | E2404 | Death (Cause Unknown) | Level 2 | The cessation of all biological functions that sustain life, where the exact reason or underlying condition leading to death has not been determined or identified | Add |  | Death (cause unknown): Death (cause unknown) can be considered an harm for the cases where non information about the causes of death are available Medrra 10011906 LLT- 10063378 | None | Subject's daughter called to report patient had died. Records requested. Unknown cause of death. | REJECT | Death can be coded. If the death can be determined not to be related to the device, this can be indicated with F29, if there is a possible relationship or contribution of the device to the death it should be coded with F02
156 | 2025-08-28 | DITTA/GMTA | E | 2025 | E240101 | Health Effect Not Yet Characterized | Level 3 | It is not clear whether any health effects occurred, or health effect appears to have occurred but there is not yet enough information available to classify the clinical signs, symptoms and conditions. | Add |  | The current IMDRF coding system does not adequately differentiate the reasons behind incomplete event information, limiting the clarity and usefulness of incident reporting. To address this, we propose three new codes: Health Effect Not Yet Characterized, Unsuccessful Data Collection, and Reporter Declined to Provide Information. These additions will enhance transparency, improve documentation accuracy, and better reflect real-world reporting scenarios—ultimately supporting more reliable analysis and regulatory decision-making. | None |  | REJECT | This can be coded by E2401
157 | 2025-08-28 | DITTA/GMTA | E | 2025 | E240102 | Unsuccessful Data collection | Level 3 | Despite repeated inquiries, no further information about the events was provided. | Add |  | The current IMDRF coding system does not adequately differentiate the reasons behind incomplete event information, limiting the clarity and usefulness of incident reporting. To address this, we propose three new codes: Health Effect Not Yet Characterized, Unsuccessful Data Collection, and Reporter Declined to Provide Information. These additions will enhance transparency, improve documentation accuracy, and better reflect real-world reporting scenarios—ultimately supporting more reliable analysis and regulatory decision-making. | None |  | REJECT | This can be coded by E2401
158 | 2025-08-28 | DITTA/GMTA | E | 2025 | E240103 | Reporter declined to provide information | Level 3 | Reporter has been contacted but explicitly declined to provide further information. | Add |  | The current IMDRF coding system does not adequately differentiate the reasons behind incomplete event information, limiting the clarity and usefulness of incident reporting. To address this, we propose three new codes: Health Effect Not Yet Characterized, Unsuccessful Data Collection, and Reporter Declined to Provide Information. These additions will enhance transparency, improve documentation accuracy, and better reflect real-world reporting scenarios—ultimately supporting more reliable analysis and regulatory decision-making. | None |  | REJECT | This can be coded by E2401
159 | 2025-08-28 | DITTA/GMTA | E | 2025 | E2111 | Injection site reaction | Level 2 | A reaction (usually immunologic) developing at the site of injection | Modify | Remove (Intense) from the definition | The word "intense" is limiting the use of this code. "intense" tends to be aligned with harms with higher severity. ''Injection site reactions'' leading minor harms are excluded from this term due to the word intens doesn't represent the severity of minor reactions. | None | A day after administration, physician noticed mild erythema around the injection site. | REJECT | Intense required in definition to capture intense/unusual reactions
160 | 2025-08-28 | DITTA/GMTA | E | 2025 | E1725 | Injection Site Pooling | Level 2 | Liquid or a pocket of air trapped under the skin, typically due to injection problems. Sometimes refers to Subcutaneous Emphysema.Does not include inflammation. | Modify | Change the Term Definition and add to "Subcutaneos Emphysema'' to it | The definition of injection site pooling is overlapping with the definition for Subcutaneous emphysema. Please, modify the Term name and definition of E1725; Include SubCutaneos Emphysema instead. | None | Patient received an intramuscular injection of drug X in the left deltoid region at a local clinic. Withing 4 hours the patient reported crackling sensation and mild swelling at the injection site. On examination at the emergency department, crepitus was plapable on the upper arm without signs of erythema or warmth. X-ray confirmed the presence of subcutaneous emphysema localised to the soft tissues at and around the injection site. | REJECT | Covered by row 16 and 197
161 | 2025-08-28 | DITTA/GMTA | E | 2025 | E21xx | Infusion site reaction | Level 2 | A localized response at the site where a medication or fluid is administered through an intravenous (IV), subcutaneous, or other infusion route | Add |  | Infusion site reaction is intented to cover the reactions during intravenous delivery using infusion devices. This term differs from the E2111 "injection site reaction" because the latter is supposed to cover the subcutaneous reactions. | None | Approximately an hour post-infusion, the patient reported a reaction at the infusion site | DUPLICATE
162 | 2025-08-28 | DITTA/GMTA | E | 2025 | new level 2 code under E10 | Abdominal compartment syndrome | Level 2 | Abdominal compartment syndrome (ACS) is a condition that occurs when bleeding or swelling leads to rises of pressure in the abdominal cavity that causes end-organ damage, dysfunction or failure. | Add | Add term and description as suggested on left. | No code/term available to classify a known complication of e.g. NPWT (negative pressure wound therapy). Meddra: 10058808 | E10 | Patient had abdominal compartment syndrome during treatment of open abdomen with negative pressure wound therapy (device) to achieve fascial closure. | ACCEPT | Agree to add | E1040 | New Term: Abdominal Compartment Syndrome New Definition: Abdominal compartment syndrome (ACS) is a condition that occurs when bleeding or swelling leads to rises of pressure in the abdominal cavity that causes end-organ damage, dysfunction or failure.
163 | 2025-08-28 | DITTA/GMTA | E | 2025 | New level 2 code under E21 | Crust or scab | 2nd level | Crusts are dried serum, pus, or blood, usually mixed with epithelial and sometimes bacterial debris. | Add | Addition, see left. | Received reports of significance and a suitable term/code was unavailable (Meddra: 10039509 - Scab, 10076524 - Nasal crusting). | None | Nasal crusting following treatment of turbinate hypertrophy (e.g. endoscopic or ablation device) can require re-intervention or may take months to resolve on its own. | MODIFY | Agreed to add Scab. Proposed to move to E23, not E21. | E2353 | New Term: Scab New Definition: a dry, rough protective crust that forms over a cut or wound during healing.
164 | 2025-08-28 | DITTA/GMTA | F | 2025 | F31 | Cancelled/Aborted/ Treatment/Therapy or No Treatment/Therapy performed | Level 1 term | Patient treatment aborted, or unable to be completed. This applies to cases where there is no information whether the treatment will be carried out at a later time, or where it is known that treatment will not go ahead. If it is known that the treatment has been, or will be, carried out later then use F05 - Delay to Treatment/ Therapy | Modify | Change termonoligy and perhaps add additional level 2 term to F31. | Change term to Cancelled/aborted Treatment/Therapy or no Treatment/Therapy performed. Perhaps Add additional level to specify between cancelled/aborted and not perfomed. The current F31 - Cancelled/Aborted Treatment/Therapy is not sufficient if it is reported that eiter the patient denied treatment or no treatment was performed. | no impact | 1. Cancelled/aborted treatment/Therapy: Patient treatment aborted, or unable to be completed. This applies to cases where there is no information whether the treatment will be carried out at a later time, or where it is known that treatment will not go ahead. 2. No treatment/Therapy performed: No intervention has been planned/performed to treat the reported adverse event. (eg. At the patient's request, it was decided not to perform additional treatments or a wait and see aproach have been desided by health professional) | REJECT | The choice to deny treatment isn't a consequence of the device performance. No treatement can be captured under F1001 if it is a consequence of device performance
165 | 2025-08-28 | DITTA/GMTA | G | 2025 | G07004 | No Defective Part/Component/Sub-assembly | Level 2 | The device and/or device data was analyzed and found to be functioning properly. There was no product malfunction. This term is used when a device-related component cannot be assigned, as the reported event is not attributable to a failure of the overall device or its subcomponents. | Add | add the term 'No Defective Part/Component/Sub-assembly' | The change is necessary to enable appropriate classification of complaint cases in which the device was analyzed and found to be operationg properly, with no device malfunction or component failure. Currently, no existing IMDRF Annex G term accuratly captures this scenario. While G07001 ("Part/Component/Sub-assembly not applicable") is inteded for cases involving the overall device, it implies that the device caused the event. The absence of a suitable code for "in Specification" outcomes can lead to inconsistent classification. | Additional more specific term to Annex G: G07 - Other Terms/Codes Related to Components | It was reported that an ICD activated the elective replacement indicator (ERI) prematurely. A thorough analysis of the device data and the returned ICD confirmed that the device was functioning properly. The activation of the ERI battery status was due to high-output programming, which led to increased energy consumption. No device malfunction or component failure was present. | REJECT | The outcome of investigations is captured with D codes.
166 | 2025-08-28 | DITTA/GMTA | G | 2025 | G07005 | Part/Component/Sub-assembly Not Determinable | Level 2 | The device and/or device data was analyzed and the reported event could be confirmed. However, no deviation or failure could be identified. The root cause remains unknwon despite all available information. This term is used when the event cannot be attributed to any specific component nor explained by external factors or influences. | Add | add the term 'Part/Component/Sub-assembly Not Determinable' | The change is necessary to support the classification of complaints for which a thorough investigation has been completed, but no failure, malfunction or identifiable root cause could be determined. The existing IMDRF code G07003 (“Insufficient Component Information”) is restricted to situations where further information is expected or needed. Likewise, G07001 assumes the device is affected overall, which is inappropriate when no impact or component association can be established. In such cases, the only available option is G07002, which should be reserved for cases requiring a new code proposal. G07005 addresses this specific gap by providing a term for concluded investigations with unknwon root cause. | Additional more specific term to Annex G: G07 - Other Terms/Codes Related to Components | It was reported that an ICD activated the elective replacement indicator (ERI) prematurly and unexpectedly. The data confirmed the premature battery depletion. However, a thorough analysis of the device data and the device itself, including the electronic module, the battery and other relevant device components, revealed no malfunction. | REJECT | The code G07001 - Part/Component/Sub-assembly Term not Applicable covers this under the definition as it has the phrase "it would not be appropriate to link the reported incident to a single part, component, or sub-assembly" which is clearly true if no such part can be idetnitifed.
167 | 2025-08-28 | DITTA/GMTA | G | 2025 | G07003 | Insufficient Component Information | Level 2 | There is not enough information available to classify the medical device component or there is not enough information available to determine whether a medical device component is affected. | Modify | modify the description of "Insufficient Component Information" | The change is necessary to remove the word “yet” which implies that additional information is expected, which is not consistently the case. In addition, to support the classification of complaints for which it is not possible to determine whether a device component is at all affected. | Broadening term description to Annex G07003 - Insufficient Component Information | It was reported that the elective replacement indicator (ERI) was activated. However, no device or data are available for analysis. Based solely on the complaint description it is not possible to determine whether the ERI activation is as expected or prematurely. Therefore, no relationship to a device component can be established. | ACCEPT |  |  | New Definition: There is not enough information available to classify the medical device component or there is not enough information available to determine whether a medical device component is affected.
168 | 2025-08-28 | DITTA/GMTA | G | 2025 | G04034 | Connector/Coupler - Male | Level 3 | A connector which has a male luer design. | Add | Addition | There is not a clear distinction in the connector/coupler coding to suggest the type of connector/coupler ie.Male or Female. | Additional more specific term to Annex G: G04034 | It was reported that an extension set was leaking. Sample analysis revealed the male luer connector outer diameter was out of specification. | MODIFY | Modify. Specifying luer is not appropriate in the defintion, but agree to add subcodes for male and female | G0403402 | New Term: Connector/Coupler - Male New Definition: A connector with a male design
169 | 2025-08-28 | DITTA/GMTA | G | 2025 | G04034 | Connector/Coupler - Female | Level 3 | A connector which has a female luer design. | Add | Addition | There is not a clear distinction in the connector/coupler coding to suggest the type of connector/coupler ie.Male or Female. | Additional more specific term to Annex G: G04034 | It was reported that an extension set was leaking. Sample analysis revealed the female luer connector inner diameter was out of specification. | MODIFY | Modify. Specifying luer is not appropriate in the defintion, but agree to add subcodes for male and female | G0403403 | New Term: Connector/Coupler - Female New Definition: A connector with a female design
170 | 2025-08-28 | DITTA/GMTA | G | 2025 | G04135 | Valve - Needle Free Connector | Level 3 | A valve which is designed to to be accessed with a flat tip male luer, and seals once the male luer is disconnected. | Add | Addition | Addition of a needle free connector | Additional more specific term to Annex G: G04135 | It was reported that the valve was difficult to access. Sample evaluation determined that the needle free connector valve had reknit. | REJECT | Reject. This appears to be a luer value which is already captured under G0413502
171 | 2025-08-28 | DITTA/GMTA | G | 2025 | G04 | Spike | Level 2 | A component used to access or pierce a container. | Add | Addition | Replace the use of Trocar, which does not fit the definition of a spike. | Additional more specific term to Annex G: G04 | It was reported that the spike of a IV admin set was dull and could not access the IV solution. | MODIFY | agree, but refine the definition to be “A non-surgical component designed to access or pierce a container (e.g., IV fluid bag or bottle) to allow fluid transfer through an administration set.” | G04143 | Spike A non-surgical component designed to access or pierce a container (e.g., IV fluid bag or bottle) to allow fluid transfer through an administration set
172 | 2025-08-28 | DITTA/GMTA | G | 2025 | G04052 | Clamp - Roller | Level 3 | A clamp which contains a roller wheel to allow full, or partial occlusion of system. | Add | Addition | Provide more granularity in the type of clamp used. | Additional more specific term to Annex G: G04052 | It was reported that the IV set was not accurate in delivery. The sample evaluation determined the roller clamp was engaged. | MODIFY | Create sub termfor control clamp and add definition with examples | G0413504 | New Term: Control Clamp New definition: A clamp mechanism designed to occlude flow within a medical device system. This includes roller clamps, slide clamps, and pinch clamps.
173 | 2025-08-28 | DITTA/GMTA | G | 2025 | G04052 | Clamp - Slide | Level 3 | A clamp which occludes flow by being slid into place. | Add | Addition | Provide more granularity in the type of clamp used. | Additional more specific term to Annex G: G04052 | It was reported that the tubing was cut. The sample evaluation revealed that the slide clamp contained a bur. | MODIFY | Create sub termfor control clamp and add definition with examples | G0413504 | New Term: Control Clamp New definition: A clamp mechanism designed to occlude flow within a medical device system. This includes roller clamps, slide clamps, and pinch clamps.
174 | 2025-08-28 | DITTA/GMTA | G | 2025 | G04052 | Clamp - Pinch | Level 3 | A clamp which occludes flow by being pressed shut. | Add | Addition | Provide more granularity in the type of clamp used. | Additional more specific term to Annex G: G04052 | It was reported that the tubing was kinked near the pinch clamp. Sample evaluation revelaed pinch clamp was within tolerance and kink was not confirmed. | MODIFY | Create sub termfor control clamp and add definition with examples | G0413504 | New Term: Control Clamp New definition: A clamp mechanism designed to occlude flow within a medical device system. This includes roller clamps, slide clamps, and pinch clamps.
175 |  | FDA | E | 2025 | E0512 | Vasculitis | Level 2 | Typo in definition | Modify | Inflammation of any one of the blood vessels, including the arteries, veins, and rest of the vasculature system in the body. | Editorial change for consistency |  |  | ACCEPT | Editorial |  | New Definition: Inflammation of any one of the blood vessels, including the arteries, veins, and rest of the vasculature system in the body.
176 | 2025-08-29 | FDA | A | 2025 | A090804 | False Positive Result | Level 3 | Problem associated with the device incorrectly reporting that something has been detected and may mislead the operatorr/user/HCP to take certain actions. | Modify | Problem associated with the device incorrectly reporting that something has been detected and may mislead the operator/user/HCP to take certain actions. | Typo in definition | None | N/A, editorial change | ACCEPT | Editorial |  | New Definition: Problem associated with the device incorrectly reporting that something has been detected and may mislead the operator/user/HCP to take certain actions.
177 | 2025-08-29 | IMDRF AET WG | A | 2025 | A03 | Chemical Problem | Level 1 | Problem associated with any from the documented specifications of the device that relate to any chemical characterization, i.e., element, compound, or mixture. | Modify | Add 'deviation' to definition. |  | nil |  | ACCEPT | Editorial |  | New Definition: Problem associated with any deviation from the documented specifications of the device that relate to any chemical characterization, i.e., element, compound, or mixture.
178 | 2025-08-29 | IMDRF AET WG | F | 2025 | F05 | Delay to Treatment/ Therapy | Level 1 | Patient treatment was delayed as a consequence of device performance. | Modify | Consider adding “interrupted” | Modify definition | NA - Editorial |  | ACCEPT | Editorial |  | Patient treatment was delayed or interrupted as a consequence of device performance.
179 |  | IMDRF AET WG | A | 2025 | A072102 | Electrical Shorting | Level 3 | Problem associated with an electric current travelling along an accidental path (unintended path) in a circuit. | Modify | Problem associated with an electric current traveling along an accidental path (unintended path) in a circuit. | Correct spelling inconsistencies. It is stated that “IMDRF Terms and Definitions use American spelling” but some do not. | NA - Editorial | NA - Editorial | ACCEPT | Editorial |  | New Definition: Problem associated with an electric current traveling along an accidental path (unintended path) in a circuit.
180 |  | IMDRF AET WG | A | 2025 | A21 | Labelling, Instructions for Use or Training Problem | Level 1 | Problem associated with device markings/labelling, instructions for use, training and maintenance documentation or guidelines. | Modify | Amend to 'Labeling' in both name and definition | Correct spelling inconsistencies. It is stated that “IMDRF Terms and Definitions use American spelling” but some do not. | NA - Editorial | NA - Editorial | ACCEPT | Editorial |  | New Term: Labeling, Instructions for Use or Training Problem New Definition: Problem associated with device markings/labeling, instructions for use, training and maintenance documentation or guidelines
181 |  | IMDRF AET WG | A | 2025 | A2101 | Device Markings/Labelling Problem | Level 3 | Problem associated with the written, printed or graphic material accompanying or affixed to the device or any of its packaging. This includes verbal instructions relating to identification, technical description, and usage provided by the device manufacturers. Problems can include but are not limited to this material being unclear, missing, worn out, incorrect or inaccurate. | Modify | Amend to 'Labeling' in name | Correct spelling inconsistencies. It is stated that “IMDRF Terms and Definitions use American spelling” but some do not. | NA - Editorial | NA - Editorial | ACCEPT | Editorial |  | New Term: Device Markings/Labeling Problem
182 | 2025-08-29 | IMDRF AET WG | B | 2025 | B31 | Labelling Review | Level 1 | A review of the reported information against the product labelling, including instructions for use was conducted to determine if the complaint resulted from inadequacies in the product labelling or a failure by the user to adhere to the labelling. Note: labelling review alone is not sufficient as an investigation for the incident. | Modify | Amend to 'Labeling' in both name and definition | Correct spelling inconsistencies. It is stated that “IMDRF Terms and Definitions use American spelling” but some do not. | C0107 - amended as per row 9 | NA - Editorial | ACCEPT | Editorial |  | Labeling Review A review of the reported information against the product labeling, including instructions for use was conducted to determine if the complaint resulted from inadequacies in the product labeling or a failure by the user to adhere to the labeling. Note: labeling review alone is not sufficient as an investigation for the incident.
183 | 29/08/2025 | IMDRF AET WG | C | 2025 | C0107 | Unintended Prescence of Allergans | Level 2 | Unintended or unexpected presence of allergens in the device. If the presence of the allergen is expected but not adequately labelled, then use "Labelling Problem". | Modify | Amend to 'Labeling' in definition | Correct spelling inconsistencies. It is stated that “IMDRF Terms and Definitions use American spelling” but some do not. | None | NA - Editorial | ACCEPT | See PMDA comment row 195. Amended to 'label' and added C05. |  | Unintended or unexpected presence of allergens in the device. If the presence of the allergen is expected but not adequately labeled, then use C05 or child terms.
184 | 29/08/2025 | IMDRF AET WG | C | 2025 | C0501 | Inadequate Labelling and/or Instructions for Use | Level 2 | Inadequate information on the labels or in the instructions for use e.g. steps that are difficult to follow or that are missing. | Modify | Amend to 'Labeling' in term | Correct spelling inconsistencies. It is stated that “IMDRF Terms and Definitions use American spelling” but some do not. | NA - Editorial | NA - Editorial | ACCEPT | Editorial |  | Inadequate Labeling and/or Instructions for Use
185 | 29/08/2025 | IMDRF AET WG | C | 2025 | C1403 | Problem Related to Variant/Mutant | Level 2 | Device performance affected by variant/mutant which is not covered in the labelling. | Modify | Amend to 'Labeling' in definition | Correct spelling inconsistencies. It is stated that “IMDRF Terms and Definitions use American spelling” but some do not. | NA - Editorial | NA - Editorial | ACCEPT | Editorial |  | Device performance affected by variant/mutant which is not covered in the labeling.
186 |  | IMDRF AET WG | E | 2025 | E0142 | Idiopathic Intracranial Hypertension | Level 2 | A condition that mimics the presence of a tumour (Pseudotumour Cerebri) or hydrocephalus, caused by chronic increase in intracranial pressure. | Modify | Amend Tumour to Tumor in definition | Correct spelling inconsistencies. It is stated that “IMDRF Terms and Definitions use American spelling” but some do not. | NA - Editorial | NA - Editorial | ACCEPT | Editorial |  | A condition that mimics the presence of a tumor (Pseudotumour Cerebri) or hydrocephalus, caused by chronic increase in intracranial pressure.
187 |  | IMDRF AET WG | E | 2025 | E02 | Mental, Emotional and Behavioural Disorders | Level 1 | Mental, Emotional and Behavioural Disorders | Modify | Amend Behavioural to Behavioral | Correct spelling inconsistencies. It is stated that “IMDRF Terms and Definitions use American spelling” but some do not. | NA - Editorial | NA - Editorial | ACCEPT | Editorial |  | Mental, Emotional and Behavioral Disorders
188 |  | IMDRF AET WG | E | 2025 | E0206 | Unspecified Mental, Emotional or Behavioural Problem | Level 2 | The report describes a non-specific Mental, Emotional or Behavioural problem. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific problem but the relevant term does not exist. | Modify | Amend Behavioural to Behavioral | Correct spelling inconsistencies. It is stated that “IMDRF Terms and Definitions use American spelling” but some do not. | NA - Editorial | NA - Editorial | ACCEPT | Editorial |  | The report describes a non-specific Mental, Emotional or Behavioral problem. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific problem but the relevant term does not exist.
189 |  | IMDRF AET WG | E | 2025 | E1401 | Abnormal Vaginal Discharge | Level 2 | Secretions from the vagina with a smell, colour, or texture which is different from normal. | Modify | Amend colour to color in definition | Correct spelling inconsistencies. It is stated that “IMDRF Terms and Definitions use American spelling” but some do not. | NA - Editorial | NA - Editorial | ACCEPT | Editorial |  | Secretions from the vagina with a smell, color, or texture which is different from normal.
190 |  | IMDRF AET WG | E | 2025 | E1804 | Solid Tumour | Level 2 | Abnormal mass of tissue that usually does not contain cysts or liquid areas. Solid tumours may be benign (not cancerous), or malignant (cancerous). Different types of solid tumours are named for the type of cells that form them. If tumours are associated with those under the term cancer, these terms should be used. | Modify | Amend Tumour to Tumor in term and definition | Correct spelling inconsistencies. It is stated that “IMDRF Terms and Definitions use American spelling” but some do not. | NA - Editorial | NA - Editorial | ACCEPT | Editorial |  | Abnormal mass of tissue that usually does not contain cysts or liquid areas. Solid tumors may be benign (not cancerous), or malignant (cancerous). Different types of solid tumors are named for the type of cells that form them. If tumors are associated with those under the term cancer, these terms should be used.
191 | 2025-08-29 | IMDRF AET WG | F | 2025 | F15 | Recognised Device or Procedural Complication | Level 1 | Patient experiencing a complication that is well documented in association with the device or procedure. | Modify | Amend recognised to recognize in term | Correct spelling inconsistencies. It is stated that “IMDRF Terms and Definitions use American spelling” but some do not. | NA - Editorial | NA - Editorial | ACCEPT | Editorial |  | New Term: Recognized Device or Procedural Complication
192 | 2025-08-29 | IMDRF AET WG | F | 2025 | F1910 | Cancelled/Aborted Surgical Procedure | Level 2 | Surgical procedure aborted, or unable to be completed. This applies to cases where there is no information whether the procedure will be carried out at a later time, or where it is known that surgery will not go ahead. If it is known that the procedure has been, or will be, carried out later then use F1909 - Surgical Procedure Delayed | Modify | Amend cancelled to canceled in term | Correct spelling inconsistencies. It is stated that “IMDRF Terms and Definitions use American spelling” but some do not. | NA - Editorial | NA - Editorial | ACCEPT | Editorial |  | New Term: Canceled/Aborted Surgical Procedure
193 | 29/08/2025 | IMDRF AET WG | F | 2025 | F31 | Cancelled/Aborted Treatment/Therapy | Level 1 | Patient treatment aborted, or unable to be completed. This applies to cases where there is no information whether the treatment will be carried out at a later time, or where it is known that treatment will not go ahead. If it is known that the treatment has been, or will be, carried out later then use F05 - Delay to Treatment/ Therapy | Modify | Amend cancelled to canceled | Correct spelling inconsistencies. It is stated that “IMDRF Terms and Definitions use American spelling” but some do not. | NA - Editorial | NA - Editorial | ACCEPT | Editorial |  | New Term: Canceled/Aborted Treatment/Therapy
194 | 2025-08-29 | IMDRF AET WG | A | 2025 | A24 | Adverse Event Without Identified Device or Use Problem | Level 1 | An adverse event (e.g. patient harm) appears to have occurred, but there does not appear to have been a problem with the device or the way it was used. | Modify | Change to: No apparent device problem | The existence of an adverse event is indicated by A, E, and F codes in combination so not appropriate to code it just using A. | Retire A25 | NA - Editorial | ACCEPT | Editorial |  | New Term: No Apparent Device Problem
195 | 29/08/2025 | PMDA, Japan | C | 2025 | C0107 | Unintended Presence of Allergens | Level 2 | Unintended or unexpected presence of allergens in the device. If the presence of the allergen is expected but not adequately labelled, then use "Labelling Problem". | Modify | It would be preferable if "Labelling Problem" in the definition is replaced with "C05 or child terms". | This reference is not perfectlly alighned with the expression of actual term in the terminology and there is no reference of codes (In annex C, we have C05 "Labeling and Instructions for Use/Maintenance" or relevant child terms. Even if selection of A terms is intended, relevant term A21 "Labelling, Instructions for Use or Training Problem" maches the refered term neither.), so it might be confusing. Also, term name can be changed in the future revision, so how about refering other term by using code only in principle. | The terms that refer in the definition the use of any other terms (PMDA have specified the terms of this type, but it needs considered if we immediately amend all of them). | No example because of the editorial change. | ACCEPT | Editorial |  | Unintended or unexpected presence of allergens in the device. If the presence of the allergen is expected but not adequately labeled, then use C05 or child terms.
196 | 29/08/2025 | PMDA, Japan | E | 2025 | E0139 | Unspecified Nervous System Problem | Level 2 | The report describes a non-specific problem with the nervous system. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific injury but the relevant term does not exist. | Modify | It would be preferable if "Appropriate Term / Code Not Available" in the definition is replaced with "E2402". | This reference is not perfectlly alighned with the expression of actual term because E2402 was modified to "Appropriate Clinical Signs, Symptoms and Conditions Term/Code Not Available" in the last revision. Also, term name can be changed in the future revision, so how about refering other term by using code only in principle. | The terms that refer in the definition the use of any other terms (PMDA have specified the terms of this type, but it needs considered if we immediately amend all of them). | No example because of the editorial change. | ACCEPT | Editorial |  | The report describes a non-specific problem with the nervous system. Note: Please use E2402 if the report describes a specific injury but the relevant term does not exist.
197 | 29/08/2025 | PMDA, Japan | E | 2025 | E0206 | Unspecified Mental, Emotional or Behavioural Problem | Level 2 | The report describes a non-specific Mental, Emotional or Behavioural problem. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific problem but the relevant term does not exist. | Modify | It would be preferable if "Appropriate Term / Code Not Available" in the definition is replaced with "E2402". | This reference is not perfectlly alighned with the expression of actual term because E2402 was modified to "Appropriate Clinical Signs, Symptoms and Conditions Term/Code Not Available" in the last revision. Also, term name can be changed in the future revision, so how about refering other term by using code only in principle. | The terms that refer in the definition the use of any other terms (PMDA have specified the terms of this type, but it needs considered if we immediately amend all of them). | No example because of the editorial change. | ACCEPT | Editorial |  | The report describes a non-specific Mental, Emotional or Behavioural problem. Note: Please use E2402 if the report describes a specific injury but the relevant term does not exist.
198 | 29/08/2025 | PMDA, Japan | E | 2025 | E0310 | Unspecified Blood or Lymphatic problem | Level 2 | The report describes a non-specific problem with the blood or lymphatic system. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific problem but the relevant term does not exist. | Modify | It would be preferable if "Appropriate Term / Code Not Available" in the definition is replaced with "E2402". | This reference is not perfectlly alighned with the expression of actual term because E2402 was modified to "Appropriate Clinical Signs, Symptoms and Conditions Term/Code Not Available" in the last revision. Also, term name can be changed in the future revision, so how about refering other term by using code only in principle. | The terms that refer in the definition the use of any other terms (PMDA have specified the terms of this type, but it needs considered if we immediately amend all of them). | No example because of the editorial change. | ACCEPT | Editorial |  | The report describes a non-specific problem with the blood or lymphatic system. Note: Please use E2402 if the report describes a specific injury but the relevant term does not exist.
199 | 29/08/2025 | PMDA, Japan | E | 2025 | E0405 | Unspecified Immune System Problem | Level 2 | The report describes a non-specific problem with the immune system. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific problem but the relevant term does not exist. | Modify | It would be preferable if "Appropriate Term / Code Not Available" in the definition is replaced with "E2402". | This reference is not perfectlly alighned with the expression of actual term because E2402 was modified to "Appropriate Clinical Signs, Symptoms and Conditions Term/Code Not Available" in the last revision. Also, term name can be changed in the future revision, so how about refering other term by using code only in principle. | The terms that refer in the definition the use of any other terms (PMDA have specified the terms of this type, but it needs considered if we immediately amend all of them). | No example because of the editorial change. | ACCEPT | Editorial |  | The report describes a non-specific problem with the immune system. Note: Note: Please use E2402 if the report describes a specific injury but the relevant term does not exist.
200 | 29/08/2025 | PMDA, Japan | E | 2025 | E0518 | Unspecified Vascular Problem | Level 2 | The report describes a non-specific problem with the vascular system. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific problem but the relevant term does not exist. | Modify | It would be preferable if "Appropriate Term / Code Not Available" in the definition is replaced with "E2402". | This reference is not perfectlly alighned with the expression of actual term because E2402 was modified to "Appropriate Clinical Signs, Symptoms and Conditions Term/Code Not Available" in the last revision. Also, term name can be changed in the future revision, so how about refering other term by using code only in principle. | The terms that refer in the definition the use of any other terms (PMDA have specified the terms of this type, but it needs considered if we immediately amend all of them). | No example because of the editorial change. | ACCEPT | Editorial |  | The report describes a non-specific problem with the vascular system. Note: Please use E2402 if the report describes a specific injury but the relevant term does not exist.
201 | 29/08/2025 | PMDA, Japan | E | 2025 | E0623 | Unspecified Heart Problem | Level 2 | The report describes a non-specific problem with the heart. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific problem but the relevant term does not exist. | Modify | It would be preferable if "Appropriate Term / Code Not Available" in the definition is replaced with "E2402". | This reference is not perfectlly alighned with the expression of actual term because E2402 was modified to "Appropriate Clinical Signs, Symptoms and Conditions Term/Code Not Available" in the last revision. Also, term name can be changed in the future revision, so how about refering other term by using code only in principle. | The terms that refer in the definition the use of any other terms (PMDA have specified the terms of this type, but it needs considered if we immediately amend all of them). | No example because of the editorial change. | ACCEPT | Editorial |  | The report describes a non-specific problem with the heart. Note: Please use E2402 if the report describes a specific injury but the relevant term does not exist.
202 | 29/08/2025 | PMDA, Japan | E | 2025 | E0752 | Unspecified Respiratory Problem | Level 2 | The report describes a non-specific problem with the respiratory system. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific problem but the relevant term does not exist. | Modify | It would be preferable if "Appropriate Term / Code Not Available" in the definition is replaced with "E2402". | This reference is not perfectlly alighned with the expression of actual term because E2402 was modified to "Appropriate Clinical Signs, Symptoms and Conditions Term/Code Not Available" in the last revision. Also, term name can be changed in the future revision, so how about refering other term by using code only in principle. | The terms that refer in the definition the use of any other terms (PMDA have specified the terms of this type, but it needs considered if we immediately amend all of them). | No example because of the editorial change. | ACCEPT | Editorial |  | The report describes a non-specific problem with the respiratory system. Note: Please use E2402 if the report describes a specific injury but the relevant term does not exist.
203 | 29/08/2025 | PMDA, Japan | E | 2025 | E0844 | Unspecified Eye / Vision Problem | Level 2 | The report describes a non-specific problem with the eye(s) or vision. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific problem but the relevant term does not exist. | Modify | It would be preferable if "Appropriate Term / Code Not Available" in the definition is replaced with "E2402". | This reference is not perfectlly alighned with the expression of actual term because E2402 was modified to "Appropriate Clinical Signs, Symptoms and Conditions Term/Code Not Available" in the last revision. Also, term name can be changed in the future revision, so how about refering other term by using code only in principle. | The terms that refer in the definition the use of any other terms (PMDA have specified the terms of this type, but it needs considered if we immediately amend all of them). | No example because of the editorial change. | ACCEPT | Editorial |  | The report describes a non-specific problem with the eye(s) or vision. Note: Please use E2402 if the report describes a specific injury but the relevant term does not exist.
204 | 29/08/2025 | PMDA, Japan | E | 2025 | E0906 | Unspecified Ear or Labyrinth Problem | Level 2 | The report describes a non-specific problem with the ear or labyrinth. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific problem but the relevant term does not exist. | Modify | It would be preferable if "Appropriate Term / Code Not Available" in the definition is replaced with "E2402". | This reference is not perfectlly alighned with the expression of actual term because E2402 was modified to "Appropriate Clinical Signs, Symptoms and Conditions Term/Code Not Available" in the last revision. Also, term name can be changed in the future revision, so how about refering other term by using code only in principle. | The terms that refer in the definition the use of any other terms (PMDA have specified the terms of this type, but it needs considered if we immediately amend all of them). | No example because of the editorial change. | ACCEPT | Editorial |  | The report describes a non-specific problem with the ear or labyrinth. Note: Please use E2402 if the report describes a specific injury but the relevant term does not exist.
205 | 29/08/2025 | PMDA, Japan | E | 2025 | E1033 | Unspecified Gastrointestinal Problem | Level 2 | The report describes a non-specific problem with the gastrointestinal system. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific problem but the relevant term does not exist. | Modify | It would be preferable if "Appropriate Term / Code Not Available" in the definition is replaced with "E2402". | This reference is not perfectlly alighned with the expression of actual term because E2402 was modified to "Appropriate Clinical Signs, Symptoms and Conditions Term/Code Not Available" in the last revision. Also, term name can be changed in the future revision, so how about refering other term by using code only in principle. | The terms that refer in the definition the use of any other terms (PMDA have specified the terms of this type, but it needs considered if we immediately amend all of them). | No example because of the editorial change. | ACCEPT | Editorial |  | The report describes a non-specific problem with the gastrointestinal system. Note: Please use E2402 if the report describes a specific injury but the relevant term does not exist.
206 | 29/08/2025 | PMDA, Japan | E | 2025 | E1107 | Unspecified Hepatic or Biliary Problem | Level 2 | The report describes a non-specific problem with the liver or biliary system. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific problem but the relevant term does not exist. | Modify | It would be preferable if "Appropriate Term / Code Not Available" in the definition is replaced with "E2402". | This reference is not perfectlly alighned with the expression of actual term because E2402 was modified to "Appropriate Clinical Signs, Symptoms and Conditions Term/Code Not Available" in the last revision. Also, term name can be changed in the future revision, so how about refering other term by using code only in principle. | The terms that refer in the definition the use of any other terms (PMDA have specified the terms of this type, but it needs considered if we immediately amend all of them). | No example because of the editorial change. | ACCEPT | Editorial |  | The report describes a non-specific problem with the liver or biliary system. Note: Please use E2402 if the report describes a specific injury but the relevant term does not exist.
207 | 29/08/2025 | PMDA, Japan | E | 2025 | E1311 | Unspecified Kidney or Urinary Problem | Level 2 | The report describes a non-specific problem with the kidneys or urinary system. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific problem but the relevant term does not exist. | Modify | It would be preferable if "Appropriate Term / Code Not Available" in the definition is replaced with "E2402". | This reference is not perfectlly alighned with the expression of actual term because E2402 was modified to "Appropriate Clinical Signs, Symptoms and Conditions Term/Code Not Available" in the last revision. Also, term name can be changed in the future revision, so how about refering other term by using code only in principle. | The terms that refer in the definition the use of any other terms (PMDA have specified the terms of this type, but it needs considered if we immediately amend all of them). | No example because of the editorial change. | ACCEPT | Editorial |  | The report describes a non-specific problem with the kidneys or urinary system. Note: Please use E2402 if the report describes a specific injury but the relevant term does not exist.
208 | 29/08/2025 | PMDA, Japan | E | 2025 | E1417 | Unspecified Reproductive System or Breast Problem | Level 2 | The report describes a non-specific problem with the breasts or reproductive system. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific problem but the relevant term does not exist. | Modify | It would be preferable if "Appropriate Term / Code Not Available" in the definition is replaced with "E2402". | This reference is not perfectlly alighned with the expression of actual term because E2402 was modified to "Appropriate Clinical Signs, Symptoms and Conditions Term/Code Not Available" in the last revision. Also, term name can be changed in the future revision, so how about refering other term by using code only in principle. | The terms that refer in the definition the use of any other terms (PMDA have specified the terms of this type, but it needs considered if we immediately amend all of them). | No example because of the editorial change. | ACCEPT | Editorial |  | The report describes a non-specific problem with the breasts or reproductive system. Note: Please use E2402 if the report describes a specific injury but the relevant term does not exist.
209 | 29/08/2025 | PMDA, Japan | E | 2025 | E1635 | Unspecified Musculoskeletal problem | Level 2 | The report describes a non-specific problem with the musculoskeletal system. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific problem but the relevant term does not exist. | Modify | It would be preferable if "Appropriate Term / Code Not Available" in the definition is replaced with "E2402". | This reference is not perfectlly alighned with the expression of actual term because E2402 was modified to "Appropriate Clinical Signs, Symptoms and Conditions Term/Code Not Available" in the last revision. Also, term name can be changed in the future revision, so how about refering other term by using code only in principle. | The terms that refer in the definition the use of any other terms (PMDA have specified the terms of this type, but it needs considered if we immediately amend all of them). | No example because of the editorial change. | ACCEPT | Editorial |  | The report describes a non-specific problem with the musculoskeletal system. Note: Please use E2402 if the report describes a specific injury but the relevant term does not exist.
210 | 29/08/2025 | PMDA, Japan | E | 2025 | E1729 | Unspecified Skin and Subcutaneous Tissue Problem | Level 2 | The report describes a non-specific problem with the skin or subcutaneous tissue. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific problem but the relevant term does not exist. | Modify | It would be preferable if "Appropriate Term / Code Not Available" in the definition is replaced with "E2402". | This reference is not perfectlly alighned with the expression of actual term because E2402 was modified to "Appropriate Clinical Signs, Symptoms and Conditions Term/Code Not Available" in the last revision. Also, term name can be changed in the future revision, so how about refering other term by using code only in principle. | The terms that refer in the definition the use of any other terms (PMDA have specified the terms of this type, but it needs considered if we immediately amend all of them). | No example because of the editorial change. | ACCEPT | Editorial |  | The report describes a non-specific problem with the skin or subcutaneous tissue. Note: Please use E2402 if the report describes a specific injury but the relevant term does not exist.
211 | 29/08/2025 | PMDA, Japan | E | 2025 | E2015 | Unspecified Tissue Injury | Level 2 | The report describes a non-specific injury to the skin or subcutaneous tissue. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific injury but the relevant term does not exist. | Modify | It would be preferable if "Appropriate Term / Code Not Available" in the definition is replaced with "E2402". | This reference is not perfectlly alighned with the expression of actual term because E2402 was modified to "Appropriate Clinical Signs, Symptoms and Conditions Term/Code Not Available" in the last revision. Also, term name can be changed in the future revision, so how about refering other term by using code only in principle. | The terms that refer in the definition the use of any other terms (PMDA have specified the terms of this type, but it needs considered if we immediately amend all of them). | No example because of the editorial change. | ACCEPT | Editorial |  | The report describes a non-specific injury to the skin or subcutaneous tissue. Note: Please use E2402 if the report describes a specific injury but the relevant term does not exist.
212 | 2025-08-29 | Health Canada | A | 2025 | A0103 | Loosening of Implant Not Related to Bone-Ingrowth | Level 2 | Problem associated with the loss of direct anchorage of an implanted device over time or due to an injury. | Modify | Loosening of Implant not Related to Bone-Ingrowth | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Loosening of Implant not Related to Bone-Ingrowth
213 | 2025-08-29 | Health Canada | A | 2025 | A0107 | Implant subsidence | Level 2 | Problem related to the settling or collapse of bone adjacent to an implanted device, in direction of the loading force. | Modify | Implant Subsidence | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Implant Subsidence
214 | 2025-08-29 | Health Canada | A | 2025 | A020505 | Device packaging look-alike | Level 3 | Problem associated with medical device packaging that looks similar to another product leading to confusion. The other product may be another device or a different product such as a medicine or cleaning agent. Use A210107 for Device label look-alike. | Modify | Device Packaging Look-Alike | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Device Packaging Look-Alike
215 | 2025-08-29 | Health Canada | A | 2025 | A050406 | Intravalvular regurgitation | Level 3 | Problem associated with the escape or backward flow of blood through or across the valve. Also known as transvalvular regurgitation | Modify | Intravalvular Regurgitation | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Intravalvular Regurgitation
216 | 2025-08-29 | Health Canada | A | 2025 | A071209 | Extracardiac stimulation | Level 3 | Problem associated with undesired pacemaker induced stimulation of extracardiac (muscle) tissue (e.g. diaphragm, pectoral or intercostal muscle) | Modify | Extracardiac Stimulation | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Extracardiac Stimulation
217 | 2025-08-29 | Health Canada | A | 2025 | A072201 | High impedance | Level 3 | Problem associated with higher than intended electrical impedance levels between device and patient connections. | Modify | High Impedance | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | High Impedance
218 | 2025-08-29 | Health Canada | A | 2025 | A072202 | Low impedance | Level 3 | Problem associated with lower than intended electrical impedance levels between device and patient connections. | Modify | Low Impedance | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Low Impedance
219 | 2025-08-29 | Health Canada | A | 2025 | A090209 | Visual Prompts will not Clear | Level 3 | Problem with visual messages which continue to be displayed on/by the device after the appropriate action has been taken. | Modify | Visual Prompts Will not Clear | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Visual Prompts Will not Clear
220 | 2025-08-29 | Health Canada | A | 2025 | A090404 | Output above Specifications | Level 3 | Device output is exceeding the documented specifications of the device. | Modify | Output Above Specifications | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Output Above Specifications
221 | 2025-08-29 | Health Canada | A | 2025 | A090405 | Output below Specifications | Level 3 | Device output is below the documented specifications of the device. | Modify | Output Below Specifications | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Output Below Specifications
222 | 2025-08-29 | Health Canada | A | 2025 | A140507 | Reflux within Device | Level 3 | Problem associated with partial backflow, compromising the device's flow output. | Modify | Reflux Within Device | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Reflux Within Device
223 | 2025-08-29 | Health Canada | A | 2025 | A140508 | Restricted Flow rate | Level 3 | Problem associated with flow rate. Flow volume delivered over time is not reaching intended flow rate. | Modify | Restricted Flow Rate | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Restricted Flow Rate
224 | 2025-08-29 | Health Canada | A | 2025 | A160101 | Alarm Not Visible | Level 3 | The device does not display an alarm message when required. | Modify | Alarm not Visible | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Alarm not Visible
225 | 2025-08-29 | Health Canada | A | 2025 | A160201 | Fail-Safe Did Not Operate | Level 3 | Problem associated with the device fail-safe mechanism, which did not function or function in a non effective way, compromising safe use of the device. | Modify | Fail-Safe Did not Operate | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Fail-Safe Did not Operate
226 | 2025-08-29 | Health Canada | A | 2025 | A160602 | Premature End-of-Life Indicator | Level 3 | Problem with the early or unexpected activation of the end-of-life indicator. | Modify | Premature End-Of-Life Indicator | Editorial change for consistency | None | N/A, editorial change | REJECT | Editorial |  | Premature End-Of-Life Indicator
227 | 2025-08-29 | Health Canada | A | 2025 | A180101 | Biofilm coating in Device | Level 3 | Problem associated with the undesired introduction of a biofilm coating into or onto the device. | Modify | Biofilm Coating in Device | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Biofilm Coating in Device
228 | 2025-08-29 | Health Canada | A | 2025 | A210107 | Device label look-alike | Level 3 | Problem associated with medical device label that looks similar to another product leading to confusion. The other product may be another device or a different product such as a medicine or cleaning agent. Use A020505 for Device packaging look-alike. | Modify | Device Label Look-Alike | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Device Label Look-Alike
229 | 2025-08-29 | Health Canada | A | 2025 | A2501 | Prophylactic removal due to corrective action | Level 2 | Prophylactic removal of a device due to corrective action without malfunction or problem of the specific device involved | Modify | Prophylactic Removal Due to Corrective Action | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Prophylactic Removal Due to Corrective Action
230 | 2025-08-29 | Health Canada | A | 2025 | A27 | Appropriate Device Problem Term/Code Not Available | Level 1 | The device problem is not adequately described by any other term. Note: this code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. | Modify | Appropriate Device Problem Term/Code not Available | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Appropriate Device Problem Term/Code not Available
231 | 2025-08-29 | Health Canada | B | 2025 | B05 | Testing of Device from Other Lot/Batch Returned From User | Level 1 | The investigation employed relevant empirical testing of the device of another lot or batch than that of the suspected device in the reported adverse event in order to support the identification of possible causes for the adverse event. This includes devices without a lot/batch designation. The device was returned from the user. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the technical file or in the latest relevant standards. | Modify | Testing of Device from Other Lot/Batch Returned from User | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Testing of Device from Other Lot/Batch Returned from User
232 | 2025-08-29 | Health Canada | B | 2025 | B16 | Device Not Manufactured by Reporting Manufacturer | Level 1 | Further information was obtained which established that the manufacturer of the device involved was not the one to which it was initially attributed. This includes mistaken identity or falsified device. Use in conjunction with C20 and D13 or D14. | Modify | Device not Manufactured by Reporting Manufacturer | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Device not Manufactured by Reporting Manufacturer
233 | 2025-08-29 | Health Canada | B | 2025 | B17 | Device Not Returned | Level 1 | The actual device involved in the adverse event was not returned for testing despite requests by manufacturer. | Modify | Device not Returned | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Device not Returned
234 | 2025-08-29 | Health Canada | B | 2025 | B20 | Device Not Accessible for Testing | Level 1 | The actual device involved in the adverse event is not readily accessible for testing (e.g. remains implanted in patient). | Modify | Device not Accessible for Testing | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Device not Accessible for Testing
235 | 2025-08-29 | Health Canada | B | 2025 | B21 | Type of Investigation Not Yet Determined | Level 1 | Details to determine the type of investigation are not yet available, but are being sought. Do not use this code if the investigation is complete. | Modify | Type of Investigation not Yet Determined | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Type of Investigation not Yet Determined
236 | 2025-08-29 | Health Canada | B | 2025 | B23 | Specimen Requested But Not Provided | Level 1 | A patient specimen was required to adequately investigate the issue and was requested but not provided. | Modify | Specimen Requested But not Provided | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Specimen Requested But not Provided
237 | 2025-08-29 | Health Canada | B | 2025 | B28 | Device lost in transit | Level 1 | The actual device involved in the adverse event has been returned by the user but not received for testing. | Modify | Device Lost in Transit | Editorial change for consistency | None | N/A, editorial change | REJECT | These changes are reflected in the 2025 release. |  | Device Lost in Transit
238 | 2025-08-29 | Health Canada | B | 2025 | B29 | Device data not available for required analysis | Level 1 | Device data such as software logs are not available for analysis by the manufacturer | Modify | Device Data not Available for Required Analysis | Editorial change for consistency | None | N/A, editorial change | REJECT | These changes are reflected in the 2025 release. |  | Device Data not Available for Required Analysis
239 | 2025-08-29 | Health Canada | C | 2025 | C010604 | Problem due to Thrombosis Activation | Level 3 | The device causes the formation of blood clots in or along blood vessels resulting in disturbed or disrupted blood flow. | Modify | Problem Due to Thrombosis Activation | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Problem Due to Thrombosis Activation
240 | 2025-08-29 | Health Canada | C | 2025 | C0210 | Corrupted memory | Level 2 | Malfunction that resulted from a device memory corruption, including bit flips or single event errors. | Modify | Corrupted Memory | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Corrupted Memory
241 | 2025-08-29 | Health Canada | C | 2025 | C0403 | Device Not Compatible With Another Device | Level 2 | A device that malfunctions due to being used in combination with, or in the presence of, another device. | Modify | Device not Compatible With Another Device | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Device not Compatible With Another Device
242 | 2025-08-29 | Health Canada | C | 2025 | C1201 | Fail-safe Problem Identified | Level 2 | A system intended to prevent unsafe operation of the device did not operate correctly. | Modify | Fail-Safe Problem Identified | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Fail-Safe Problem Identified
243 | 2025-08-29 | Health Canada | C | 2025 | C1404 | Pre-analytical Handling Problem | Level 2 | Incorrect pre-analytical handling of patient's sample by the user. | Modify | Pre-Analytical Handling Problem | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Pre-Analytical Handling Problem
244 | 2025-08-29 | Health Canada | C | 2025 | C1701 | Misadjustment/misalignment identified | Level 2 | Problems caused by incorrect adjustment of device components. | Modify | Misadjustment/Misalignment Identified | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Misadjustment/Misalignment Identified
245 | 2025-08-29 | Health Canada | C | 2025 | C22 | Appropriate Investigation Findings Term/Code Not Available | Level 1 | Problems is not adequately described by any other term. Note: This code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. | Modify | Appropriate Investigation Findings Term/Code not Available | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Appropriate Investigation Findings Term/Code not Available
246 | 2025-08-29 | Health Canada | D | 2025 | D1401 | Unable to exclude device problem | Level 2 | The investigation findings do not clearly confirm the presence or absence of the reported problem with the device. | Modify | Unable to Exclude Device Problem | Editorial change for consistency | None | N/A, editorial change | REJECT | Term retired in 2026 maintenance
247 | 2025-08-29 | Health Canada | D | 2025 | D1402 | Device problem excluded | Level 2 | The investigation findings clearly confirm that there was no problem with the device. | Modify | Device Problem Excluded | Editorial change for consistency | None | N/A, editorial change | REJECT | Term retired in 2026 maintenance
248 | 2025-08-29 | Health Canada | D | 2025 | D15 | Cause Not Established | Level 1 | The investigation findings do not lead to a clear conclusion about the cause of the reported adverse event. | Modify | Cause not Established | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Cause not Established
249 | 2025-08-29 | Health Canada | D | 2025 | D16 | Conclusion Not Yet Available | Level 1 | A conclusion has yet to be established as the investigation is incomplete. Do not use this code if the investigation is complete. | Modify | Conclusion not Yet Available | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Conclusion not Yet Available
250 | 2025-08-29 | Health Canada | D | 2025 | D17 | Appropriate Investigation Conclusion Term/Code Not Available | Level 1 | The concluded cause is not adequately described by any other term. Note: This code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. | Modify | Appropriate Investigation Conclusion Term/Code not Available | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Appropriate Investigation Conclusion Term/Code not Available
251 | 2025-08-29 | Health Canada | E | 2025 | E0110 | Hyporesponsive to stimuli | Level 2 | Decreased sensitivity affecting one or more senses | Modify | Hyporesponsive to Stimuli | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Hyporesponsive to Stimuli
252 | 2025-08-29 | Health Canada | E | 2025 | E0111 | Hyperresponsive to stimuli | Level 2 | Increased sensitivity affecting one or more senses | Modify | Hyperresponsive to Stimuli | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Hyperresponsive to Stimuli
253 | 2025-08-29 | Health Canada | E | 2025 | E0119 | Loss of consciousness | Level 2 | A level of awareness that can be described as consistently not responsive to stimuli. | Modify | Loss of Consciousness | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Loss of Consciousness
254 | 2025-08-29 | Health Canada | E | 2025 | E0310 | Unspecified Blood or Lymphatic problem | Level 2 | The report describes a non-specific problem with the blood or lymphatic system. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific problem but the relevant term does not exist. | Modify | Unspecified Blood or Lymphatic Problem | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Unspecified Blood or Lymphatic Problem
255 | 2025-08-29 | Health Canada | E | 2025 | E0402 | Hypersensitivity/Allergic reaction | Level 2 | A local or general reaction of an organism following contact with a specific allergen to which it has been previously exposed and to which it has become sensitized. | Modify | Hypersensitivity/Allergic Reaction | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Hypersensitivity/Allergic Reaction
256 | 2025-08-29 | Health Canada | E | 2025 | E0404 | Positive antinuclear antibodies (ANA) | Level 2 | Indicates that antibodies directed against nuclear antigens have been detected. Almost invariably found in systemic lupus erythematosus and are frequently found in rheumatoid arthritis, scleroderma, Sjogren's Syndrome and mixed connective tissue disease. | Modify | Positive Antinuclear Antibodies (ANA) | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Positive Antinuclear Antibodies (ANA)
257 | 2025-08-29 | Health Canada | E | 2025 | E050102 | Aneurysm expansion | Level 3 | Expansion of aneurysm. | Modify | Aneurysm Expansion | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Aneurysm Expansion
258 | 2025-08-29 | Health Canada | E | 2025 | E050305 | Fat embolism | Level 3 | The blockage of a blood vessel by fat particles in the circulation. | Modify | Fat Embolism | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Fat Embolism
259 | 2025-08-29 | Health Canada | E | 2025 | E051501 | Aortic dissection | Level 3 | Spontaneous or induced tear within the wall of the aorta | Modify | Aortic Dissection | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Aortic Dissection
260 | 2025-08-29 | Health Canada | E | 2025 | E0521 | Coronary obstruction/occlusion | Level 2 | Partial or total occlusion or obstruction of the coronary ostia or artery. | Modify | Coronary Obstruction/Occlusion | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Coronary Obstruction/Occlusion
261 | 2025-08-29 | Health Canada | E | 2025 | E0618 | Non specific EKG/ECG Changes | Level 2 | An electrocardiographic finding of a change in cardiac electrical activity. | Modify | Non Specific EKG/ECG Changes | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Non Specific EKG/ECG Changes
262 | 2025-08-29 | Health Canada | E | 2025 | E0713 | Decreased Peak Expiratory Flow rate | Level 2 | Decrease in the maximum flow rate generated during a forceful exhalation, starting from full lung inflation. | Modify | Decreased Peak Expiratory Flow Rate | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Decreased Peak Expiratory Flow Rate
263 | 2025-08-29 | Health Canada | E | 2025 | E0714 | Increased Peak Expiratory Flow rate | Level 2 | An increase of the greatest rate of flow that can be achieved during forced expiration beginning with the lungs fully inflated. | Modify | Increased Peak Expiratory Flow Rate | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Increased Peak Expiratory Flow Rate
264 | 2025-08-29 | Health Canada | E | 2025 | E0756 | Nasal congestion | Level 2 | The blockage of the nasal passage due to mucosal edema. It is usually the result of a viral infection or an allergic reaction. | Modify | Nasal Congestion | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Nasal Congestion
265 | 2025-08-29 | Health Canada | E | 2025 | E0757 | Breathing-related sleep disorder | Level 2 | Clinical spectrum with frequent episodes of either complete or partial obstruction of the upper airway during sleep. Includes sleep apnea. Also known as Sleep disordered breathing (SDB). | Modify | Breathing-Related Sleep Disorder | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Breathing-Related Sleep Disorder
266 | 2025-08-29 | Health Canada | E | 2025 | E0858 | Choroidal detachment | Level 2 | A detachment of the choroid from the underlying sclera often due to the accumulation of fluid (blood for hemorrhagic, other fluids for serous detachment) in the suprachoroidal space | Modify | Choroidal Detachment | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Choroidal Detachment
267 | 2025-08-29 | Health Canada | E | 2025 | E0859 | Choroidal effusion | Level 2 | Fluid accumulation in the suprachoroidal space, can lead to detachment | Modify | Choroidal Effusion | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Choroidal Effusion
268 | 2025-08-29 | Health Canada | E | 2025 | E0908 | Ear infection | Level 2 | Infection of the ear. | Modify | Ear Infection | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Ear Infection
269 | 2025-08-29 | Health Canada | E | 2025 | E0912 | Ear injury | Level 2 | Injury to the outer, middle or inner ear. | Modify | Ear Injury | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Ear Injury
270 | 2025-08-29 | Health Canada | E | 2025 | E0913 | Cerumen impaction | Level 2 | A buildup and/or blockage of earwax in the ear canal. | Modify | Cerumen Impaction | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Cerumen Impaction
271 | 2025-08-29 | Health Canada | E | 2025 | E1029 | Stenosis of the esophagus | Level 2 | Narrowing or tightening of the esophagus | Modify | Stenosis of the Esophagus | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Stenosis of the Esophagus
272 | 2025-08-29 | Health Canada | E | 2025 | E120501 | Elevated ketones/ Diabetic Ketoacidosis | Level 3 | Elevated ketones, including metabolic acidosis produced by accumulation of ketone bodies resulting from uncontrolled diabetes mellitus. | Modify | Elevated Ketones/ Diabetic Ketoacidosis | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Elevated Ketones/ Diabetic Ketoacidosis
273 | 2025-08-29 | Health Canada | E | 2025 | E1317 | Kidney injury | Level 2 | Traumatic damage to the kidney(s). | Modify | Kidney Injury | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Kidney Injury
274 | 2025-08-29 | Health Canada | E | 2025 | E151202 | Partial Neonatal Hearing loss | Level 3 | Partial loss of the ability to detect or understand sounds present in an infant within its first month after birth. | Modify | Partial Neonatal Hearing Loss | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Partial Neonatal Hearing Loss
275 | 2025-08-29 | Health Canada | E | 2025 | E161201 | Aseptic loosening | Level 3 | Failure of the fixation of a prosthetic component in the absence of infection. | Modify | Aseptic Loosening | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Aseptic Loosening
276 | 2025-08-29 | Health Canada | E | 2025 | E161202 | Septic loosening | Level 3 | Failure of the fixation of a prosthetic component in the presence of infection. | Modify | Septic Loosening | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Septic Loosening
277 | 2025-08-29 | Health Canada | E | 2025 | E161605 | Joint locking | Level 3 | Joint is blocked or becomes stuck. For Pseudo-locking use E1616 - Loss of Range of Motion. | Modify | Joint Locking | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Joint Locking
278 | 2025-08-29 | Health Canada | E | 2025 | E1625 | Nonunion/Delayed-union Bone Fracture | Level 2 | Broken bone that doesn't heal or is delayed in healing. | Modify | Nonunion/Delayed-Union Bone Fracture | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Nonunion/Delayed-Union Bone Fracture
279 | 2025-08-29 | Health Canada | E | 2025 | E1626 | Extraskeletal ossification | Level 2 | The formation of bone or of a bony substance outside the normal skeleton. Includes ectopic or heterotopic ossification, myositis ossificans. | Modify | Extraskeletal Ossification | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Extraskeletal Ossification
280 | 2025-08-29 | Health Canada | E | 2025 | E1635 | Unspecified Musculoskeletal problem | Level 2 | The report describes a non-specific problem with the musculoskeletal system. Note: Please use "Appropriate Term / Code Not Available" if the report describes a specific problem but the relevant term does not exist. | Modify | Unspecified Musculoskeletal Problem | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Unspecified Musculoskeletal Problem
281 | 2025-08-29 | Health Canada | E | 2025 | E1638 | Cartilage damage | Level 2 | Damage of cartilage. | Modify | Cartilage Damage | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Cartilage Damage
282 | 2025-08-29 | Health Canada | E | 2025 | E1642 | Spinal deformity | Level 2 | Abnormal curvature and/or alignment of the spine. | Modify | Spinal Deformity | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Spinal Deformity
283 | 2025-08-29 | Health Canada | E | 2025 | E1644 | Compartment syndrome | Level 2 | Condition in which a closed muscle compartment's pressure increases to such an extent that the circulation is diminished, e.g. due to decrease in a compartment volume or an increase in the contents of a compartment, or both. | Modify | Compartment Syndrome | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Compartment Syndrome
284 | 2025-08-29 | Health Canada | E | 2025 | E1648 | Acetabular protrusion | Level 2 | Acetabular defect with a shift of the femoral head through the pelvic bone. | Modify | Acetabular Protrusion | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Acetabular Protrusion
285 | 2025-08-29 | Health Canada | E | 2025 | E1651 | Back pain | Level 2 | Pain in any part of the back. | Modify | Back Pain | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Back Pain
286 | 2025-08-29 | Health Canada | E | 2025 | E1652 | Loose body in joint | Level 2 | Presence of fibrous, bony, cartilaginous or osseocartilaginous fragments in a synovial joint. | Modify | Loose Body in Joint | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Loose Body in Joint
287 | 2025-08-29 | Health Canada | E | 2025 | E170405 | Partial thickness (Second Degree) Burn | Level 3 | A burn that affects the epidermis and the dermis. | Modify | Partial Thickness (Second Degree) Burn | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Partial Thickness (Second Degree) Burn
288 | 2025-08-29 | Health Canada | E | 2025 | E170406 | Full thickness (Third Degree) Burn | Level 3 | A burn to the skin that extends into deeper layers and underlying tissues. | Modify | Full Thickness (Third Degree) Burn | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Full Thickness (Third Degree) Burn
289 | 2025-08-29 | Health Canada | E | 2025 | E171501 | Abnormal scarring | Level 3 | Excessive or unusual scarring related to impaired wound healing (e.g. Keloid or Hypertrophic scar) | Modify | Abnormal Scarring | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Abnormal Scarring
290 | 2025-08-29 | Health Canada | E | 2025 | E1728 | Nail disorder | Level 2 | Abnormality of the nail, such as problems with the color, shape, texture, or thickness of the fingernails or toenails. | Modify | Nail Disorder | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Nail Disorder
291 | 2025-08-29 | Health Canada | E | 2025 | E1805 | Non-malignant Neoplasm | Level 2 | A new and abnormal growth of tissue in a part of the body which is not indicative of malignant disorder. Can include benign tumors. | Modify | Non-Malignant Neoplasm | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Non-Malignant Neoplasm
292 | 2025-08-29 | Health Canada | E | 2025 | E2016 | Perineal injury | Level 2 | Traumatic damage and/or injury to the perineum. | Modify | Perineal Injury | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Perineal Injury
293 | 2025-08-29 | Health Canada | E | 2025 | E2102 | Awareness during Anaesthesia | Level 2 | A state of consciousness that permits the formation of explicit and implicit memories while under general anesthesia. | Modify | Awareness During Anesthesia | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Awareness During Anesthesia
294 | 2025-08-29 | Health Canada | E | 2025 | E2123 | Medical device site infection | Level 2 | Infection local to an implanted or invasive medical device. | Modify | Medical Device Site Infection | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Medical Device Site Infection
295 | 2025-08-29 | Health Canada | E | 2025 | E2125 | Complex regional pain syndrome | Level 2 | Condition where a person experiences persistent severe and debilitating pain, involving an extremity or other body region, hyperesthesia, and localized autonomic dysfunction following injury to soft tissue or nerve | Modify | Complex Regional Pain Syndrome | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Complex Regional Pain Syndrome
296 | 2025-08-29 | Health Canada | E | 2025 | E2126 | Loss of fracture reduction | Level 2 | Loss of anatomical alignment after fracture reduction. | Modify | Loss of Fracture Reduction | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Loss of Fracture Reduction
297 | 2025-08-29 | Health Canada | E | 2025 | E2127 | Periprosthetic fracture | Level 2 | Fractures around joint replacement prosthetics or implants. These can occur intraoperatively or postoperatively. | Modify | Periprosthetic Fracture | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Periprosthetic Fracture
298 | 2025-08-29 | Health Canada | E | 2025 | E2128 | Implant site sclerosis | Level 2 | Pathological process consisting of hardening/thickening of tissue or structures. | Modify | Implant Site Sclerosis | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Implant Site Sclerosis
299 | 2025-08-29 | Health Canada | E | 2025 | E2129 | Peritoneal effluent abnormal | Level 2 | A change in the appearance or other characteristics of peritoneal dialysis effluent (e.g. cloudy or discolored). | Modify | Peritoneal Effluent Abnormal | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Peritoneal Effluent Abnormal
300 | 2025-08-29 | Health Canada | E | 2025 | E2130 | Graft failure | Level 2 | Failure of a transplanted graft. | Modify | Graft Failure | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Graft Failure
301 | 2025-08-29 | Health Canada | E | 2025 | E2206 | High metal ion levels | Level 2 | High blood or serum levels of ions which are attributed to one or more medical devices used, such as but not limited to Co, Cr, Ti, Ni, Mo. | Modify | High Metal Ion Levels | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | High Metal Ion Levels
302 | 2025-08-29 | Health Canada | E | 2025 | E233802 | Joint effusion | Level 3 | An abnormal collection of fluid in hollow spaces or between tissues of the joints. | Modify | Joint Effusion | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Joint Effusion
303 | 2025-08-29 | Health Canada | E | 2025 | E2343 | Hemodynamic instability | Level 2 | Abnormal or unstable blood pressure | Modify | Hemodynamic Instability | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Hemodynamic Instability
304 | 2025-08-29 | Health Canada | E | 2025 | E2344 | Voice alteration | Level 2 | Changes in the pitch or quality of the voice, which may sound weak, scratchy, raspy, or otherwise out of the ordinary. | Modify | Voice Alteration | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Voice Alteration
305 | 2025-08-29 | Health Canada | E | 2025 | E2345 | Heat illness | Level 2 | A set of symptoms brought on by exposure to heat that can include confusion, hot dry skin, seizures, elevated body temperature, altered mental status, etc. Includes heat stroke and heat exhaustion. | Modify | Heat Illness | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Heat Illness
306 | 2025-08-29 | Health Canada | E | 2025 | E2347 | Breath odor | Level 2 | Unusual breath odor (e.g., halitosis, fruity/sweet breath) | Modify | Breath Odor | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Breath Odor
307 | 2025-08-29 | Health Canada | E | 2025 | E2348 | Abnormal mucus discharge | Level 2 | Any deviation from the normal quantity, consistency, and/or color of mucus produced by the mucous membranes. | Modify | Abnormal Mucus Discharge | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Abnormal Mucus Discharge
308 | 2025-08-29 | Health Canada | E | 2025 | E2352 | Exercise capacity decreased | Level 2 | Reduced ability of the body to perform strenuous activities | Modify | Exercise Capacity Decreased | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Exercise Capacity Decreased
309 | 2025-08-29 | Health Canada | E | 2025 | E2402 | Appropriate Clinical Signs, Symptoms and Conditions Term/Code Not Available | Level 2 | The clinical signs, symptoms and conditions are not adequately described by any other term. Note: this code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. | Modify | Appropriate Clinical Signs, Symptoms and Conditions Term/Code not Available | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Appropriate Clinical Signs, Symptoms and Conditions Term/Code not Available
310 | 2025-08-29 | Health Canada | F | 2025 | F2308 | Medical treatment under anaesthesia | Level 2 | Patient required non-surgical treatment under anaesthesia. | Modify | Medical Treatment Under Anaesthesia | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Medical Treatment Under Anaesthesia
311 | 2025-08-29 | Health Canada | F | 2025 | F2601 | Problem identified before clinical use/exposure | Level 2 | The adverse event occurred during or immediately prior to a clinical procedure, but the problem was identified and corrected before harm could be caused. E.g.1. an IVD gave a false reading but it was recognized as incorrect before treatment decisions were made and without any clinically significant delay to diagnosis; E.g.2. a surgical device malfunctioned during pre-use checks and the malfunction was corrected or an alternative device was available without any clinically significant delay to the procedure. | Modify | Problem Identified Before Clinical Use/Exposure | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Problem Identified Before Clinical Use/Exposure
312 | 2025-08-29 | Health Canada | F | 2025 | F27 | Problem identified during non-clinical procedure. | Level 1 | The adverse event occurred while the device was being used in a non-clinical setting and/or procedure and could not have any impact on a patient E.g.1. an IVD gave a false reading when testing an External Quality Assessment (EQA) sample; E.g.2. a surgical device malfunctioned during cleaning and/or maintenance. | Modify | Problem Identified During Non-Clinical Procedure. | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Problem Identified During Non-Clinical Procedure.
313 | 2025-08-29 | Health Canada | F | 2025 | F28 | Appropriate Health Impact Term/Code Not Available | Level 1 | The Health Impact is not adequately described by any other term. Note: this code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. | Modify | Appropriate Health Impact Term/Code not Available | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Appropriate Health Impact Term/Code not Available
314 | 2025-08-29 | Health Canada | F | 2025 | F29 | Death not related to reported adverse event | Level 1 | A death has occurred, but it was confirmed that the event did not contribute to the death. Other terms may also be selected if it is possible to link the adverse event to other identifiable health impacts. | Modify | Death not Related to Reported Adverse Event | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | Death not Related to Reported Adverse Event
315 | 2025-08-29 | Health Canada | G | 2025 | G07001 | Part/Component/Sub-assembly Term not Applicable | Level 2 | The device does not have distinct parts, components, or sub-assemblies, or it would not be appropriate to link the reported incident to a single part, component, or sub-assembly. Use this term if the problem involves or affects the overall device rather than a specific component. | Modify | Part/Component/Sub-Assembly Term not Applicable | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | New Term: Part/Component/Sub-assembly Term not Applicable
316 | 2025-08-29 | Health Canada | G | 2025 | G07002 | Appropriate Component Term/Code Not Available | Level 2 | The parts, components, or sub-assemblies are not adequately described by any other term. Note: this code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. | Modify | Appropriate Component Term/Code not Available | Editorial change for consistency | None | N/A, editorial change | ACCEPT | Editorial |  | New Term: Appropriate Component Term/Code Not Available
317 | 2025-08-29 | BfArM | A | 2025 | A070909 | over-sensing | Level 3 | Problem related to failure of the device to properly filter cardiac signals resulting in inappropriate device response. | Modify | change term to "oversensing" | it's never spelled with a hyphen in incident reports. Searching for Oversensing (if you type fast) gives you no results in the browser |  | n/a | ACCEPT | Editorial |  | New Term: Oversensing
318 | 2025-08-29 | BfArM | A | 2025 | A070910 | under-sensing | Level 3 | Problem related to failure of the device to properly detect intrinsic cardiac activity and respond appropriately. | Modify | change term to "undersensing" | it's never spelled with a hyphen in incident reports. Searching for undersensing (if you type fast) gives you no results in the browser |  | n/a | ACCEPT | Editorial |  | New Term: Undersensing
319 | 2025-08-22 | AbbVie | E | 2025 | E19xx | Biofilm | Level 2 | Biofilms are microbial communities that are attached to a surface, including living tissue, implants, and medical devices, might contribute to a chronic inflammatory response and subclinical infection. | Add | New Term | No current IMDRF code Aligns with MedDRA 10069175 Product contamination microbial | None | MD comments there is biofilm when removing breast implants and other device types. This is an Annex E code due to infection in patient from device. | REJECT | Biofilm is not a clinical sign or symptom.
320 | 2025-08-31 | MedTech Europe | E | 2025 | New child of E2121 | Endoleak | Level 2 | The observed patient condition is persistent blood flow/pressure in the aneurysm sac following an endovascular aneurysm repair (EVAR) procedure. Where other patient signs and symptoms indicate the presence of an endoleak those signs and symptoms should be coded. | Add | Add a level 3 sub-code under E2121 for additional codes to address types of endoleak - Type Ia [MedDRA: 10063824: Endoleak] | Addition of new terms allows for identification of specific types of endoleak which will better define location, cause and impact to the patient. | None | Endovascular aneurysm repair (EVAR) procedures frequently involve the use of stent grafts featuring multiple connections and branches to various arteries. Accurate and detailed classification of endoleak types is essential to enhance the monitoring, reporting, and management of adverse events/serious incidents associated with these devices. A Type Ia endoleak is an inadequate seal at the endograft proximal [main graft or fenestration seal] attachment. Implementing a dedicated E-code structure that distinctly captures each endoleak subtype would substantially improve the precision of adverse event/serious incident tracking. This enhancement would support more informed clinical decision-making, strengthen regulatory oversight, and ultimately contribute to improved patient outcomes and more robust post-market surveillance. | DUPLICATE | Duplicate of row 139
321 | 2025-08-31 | MedTech Europe | E | 2025 | New child of E2121 | Endoleak | Level 2 | The observed patient condition is persistent blood flow/pressure in the aneurysm sac following an endovascular aneurysm repair (EVAR) procedure. Where other patient signs and symptoms indicate the presence of an endoleak those signs and symptoms should be coded. | Add | Add a level 3 sub-code under E2121 for additional codes to address types of endoleak -Type Ib [MedDRA: 10063824: Endoleak] | Addition of new terms allows for identification of specific types of endoleak which will better define location, cause and impact to the patient. | None | Endovascular aneurysm repair (EVAR) procedures frequently involve the use of stent grafts featuring multiple connections and branches to various arteries. Accurate and detailed classification of endoleak types is essential to enhance the monitoring, reporting, and management of adverse events/serious incidents associated with these devices. Type Ib endoleak is an inadequate seal at the endograft distal [e.g., iliac limb or branch] attachment. Implementing a dedicated E-code structure that distinctly captures each endoleak subtype would substantially improve the precision of adverse event/serious incident tracking. This enhancement would support more informed clinical decision-making, strengthen regulatory oversight, and ultimately contribute to improved patient outcomes and more robust post-market surveillance. | DUPLICATE | Duplicate of row 140
322 | 2025-08-31 | MedTech Europe | E | 2025 | New child of E2121 | Endoleak | Level 2 | The observed patient condition is persistent blood flow/pressure in the aneurysm sac following an endovascular aneurysm repair (EVAR) procedure. Where other patient signs and symptoms indicate the presence of an endoleak those signs and symptoms should be coded. | Add | Add a level 3 sub-code under E2121 for additional codes to address types of endoleak - Type Ic [MedDRA: 10063824: Endoleak] | Addition of new terms allows for identification of specific types of endoleak which will better define location, cause and impact to the patient. | None | Endovascular aneurysm repair (EVAR) procedures frequently involve the use of stent grafts featuring multiple connections and branches to various arteries. Accurate and detailed classification of endoleak types is essential to enhance the monitoring, reporting, and management of adverse events/serious incidents associated with these devices. Type Ic endoleak is an inadequate seal at the endograft branch vessel origin in branched devices [e.g., SMA, renal] Implementing a dedicated E-code structure that distinctly captures each endoleak subtype would substantially improve the precision of adverse event/serious incident tracking. This enhancement would support more informed clinical decision-making, strengthen regulatory oversight, and ultimately contribute to improved patient outcomes and more robust post-market surveillance. | DUPLICATE | Duplicate of row 141
323 | 2025-08-31 | MedTech Europe | E | 2025 | New child of E2121 | Endoleak | Level 2 | The observed patient condition is persistent blood flow/pressure in the aneurysm sac following an endovascular aneurysm repair (EVAR) procedure. Where other patient signs and symptoms indicate the presence of an endoleak those signs and symptoms should be coded. | Add | Add a level 3 sub-code under E2121 for additional codes to address types of endoleak - Type II [MedDRA: 10063824: Endoleak] | Addition of new terms allows for identification of specific types of endoleak which will better define location, cause and impact to the patient. | None | Endovascular aneurysm repair (EVAR) procedures frequently involve the use of stent grafts featuring multiple connections and branches to various arteries. Accurate and detailed classification of endoleak types is essential to enhance the monitoring, reporting, and management of adverse events/serious incidents associated with these devices. Type II endoleak is from retrograde flow from collateral vessels. Implementing a dedicated E-code structure that distinctly captures each endoleak subtype would substantially improve the precision of adverse event/serious incident tracking. This enhancement would support more informed clinical decision-making, strengthen regulatory oversight, and ultimately contribute to improved patient outcomes and more robust post-market surveillance. | DUPLICATE | Duplicate of row 142
324 | 2025-08-31 | MedTech Europe | C | 2025 | New child of C0706 | Stress Problem Identified | Level 2 | Problems caused by either excessive or inadequate physical force exerted on it by another object resulting in problems e.g. wear, bending, deformation, fracture, fatigue. | Add | Add a level 3 sub-code under C0706 for additional codes to address type of endoleak - Type IIIa | Addition of new terms allows for identification of specific types of endoleak which will better define location, cause and impact to the patient. | None | Endovascular aneurysm repair (EVAR) procedures frequently involve the use of stent grafts featuring multiple connections and branches to various arteries. Accurate and detailed classification of endoleak types is essential to enhance the monitoring, reporting, and management of adverse events/serious incidents associated with these devices. Type IIIa endoleak is a graft material or modular failure with component disconnection [e.g., modular segments] at the attachment of aortic-aortic or aortic-bifurcated component. Implementing a dedicated C-code structure that distinctly captures each endoleak subtype would substantially improve the precision of device-related event tracking. This enhancement would support more informed clinical decision-making, strengthen regulatory oversight, and ultimately contribute to improved patient outcomes and more robust post-market surveillance. | DUPLICATE | Duplicate of row 132
325 | 2025-08-31 | MedTech Europe | C | 2025 | New child of C0706 | Stress Problem Identified | Level 2 | Problems caused by either excessive or inadequate physical force exerted on it by another object resulting in problems e.g. wear, bending, deformation, fracture, fatigue. | Add | Add a level 3 sub-code under C0706 for additional codes to address type of endoleak - Type IIIb | Addition of new terms allows for identification of specific types of endoleak which will better define location, cause and impact to the patient. | None | In endovascular aneurysm repair (EVAR) procedures, the use of complex stent grafts with multiple branches and arterial connections is common. To improve the monitoring and reporting of adverse events/serious incidents, it is essential to further refine the classification of endoleak types. The Oderich Classification (Expanded/Modern System) provides a more granular categorization of endoleaks. In that classification, a Type IIIb endoleak is a graft material or modular failure with component disconnection [e.g., modular segments] at the bifurcated iliac limb or iliac limb-iliac limb component. Implementing a dedicated C-code structure that distinctly captures each endoleak subtype would substantially improve the precision of device-related event tracking. This enhancement would support more informed clinical decision-making, strengthen regulatory oversight, and ultimately contribute to improved patient outcomes and more robust post-market surveillance. | DUPLICATE | Duplicate of row 133
326 | 2025-08-31 | MedTech Europe | C | 2025 | New child of C0706 | Stress Problem Identified | Level 2 | Problems caused by either excessive or inadequate physical force exerted on it by another object resulting in problems e.g. wear, bending, deformation, fracture, fatigue. | Add | Add a level 3 sub-code under E2121 for additional codes to address types of endoleak - Type IIIc | Addition of new terms allows for identification of specific types of endoleak which will better define location, cause and impact to the patient. | None | Endovascular aneurysm repair (EVAR) procedures frequently involve the use of stent grafts featuring multiple connections and branches to various arteries. Accurate and detailed classification of endoleak types is essential to enhance the monitoring, reporting, and management of adverse events/serious incidents associated with these devices. Type IIIc endoleak is an graft material or modular failure with seperation or disconnection of a branch component [e.g., bridging stent graft] Implementing a dedicated C-code structure that distinctly captures each endoleak subtype would substantially improve the precision of device-related event tracking. This enhancement would support more informed clinical decision-making, strengthen regulatory oversight, and ultimately contribute to improved patient outcomes and more robust post-market surveillance. | DUPLICATE | Duplicate of row 134
327 | 2025-08-31 | MedTech Europe | C | 2025 | New child of C0706 | Stress Problem Identified | Level 2 | Problems caused by either excessive or inadequate physical force exerted on it by another object resulting in problems e.g. wear, bending, deformation, fracture, fatigue. | Add | Add a level 3 sub-code under C0706 for additional codes to address type of endoleak - Type IIId | Addition of new terms allows for identification of specific types of endoleak which will better define location, cause and impact to the patient. | None | In endovascular aneurysm repair (EVAR) procedures, the use of complex stent grafts with multiple branches and arterial connections is common. To improve the monitoring and reporting of adverse events/serious incidents, it is essential to further refine the classification of endoleak types. The Oderich Classification (Expanded/Modern System) provides a more granular categorization of endoleaks. In that classification, a Type IIId endoleak specifically involves structural failures of the stent graft [tear, perforation or fracture of the side branch] Implementing a dedicated C-code structure that distinctly captures each endoleak subtype would substantially improve the precision of device-related event tracking. This enhancement would support more informed clinical decision-making, strengthen regulatory oversight, and ultimately contribute to improved patient outcomes and more robust post-market surveillance. | DUPLICATE | Duplicate of row 135
328 | 2025-08-31 | MedTech Europe | A | 2025 | A072204 | High Stimulation Impedance | Level 3 | Problem associated with higher than intended stimulation impedance levels | Add | add the term 'high stimulation impedance' | The current IMDRF coding system includes only a single, undifferentiated category for impedance-related issues, which lacks the specificity needed to reflect the diverse clinical and technical contexts in which impedance is monitored across medical devices. Impedance abnormalities, whether in stimulation, sensing, or therapeutic delivery, can indicate distinct underlying issues such as lead malfunction, tissue interface degradation, or system failure, each with different risk profiles and clinical implications. The Introduction of six new, specific impedance-related codes would enable more precise documentation and analysis of device-related events, thereby improving the overall quality of incident reporting. This enhancement would also align the IMDRF coding system with current clinical practice, where distinct impedance types are routinely separately ultimately enhancing patient safety and ensuring more appropriate regulatory evaluation. | Additional more specific term to high impedance A072201 | Pacemaker lead shows high out of range pacing impedance (>2500 ohms). OR Tachycardia lead shows increase of the pacing impedance. OR CRTD lead shows increased pacing impedance | DUPLICATE | Duplicate of row 120
329 | 2025-08-31 | MedTech Europe | A | 2025 | A072205 | Low Stimulation Impedance | Level 3 | Problem associated with lower than intended stimulation impedance levels | Add | add the term 'low stimulation impedance' | The current IMDRF coding system includes only a single, undifferentiated category for impedance-related issues, which lacks the specificity needed to reflect the diverse clinical and technical contexts in which impedance is monitored across medical devices. Impedance abnormalities, whether in stimulation, sensing, or therapeutic delivery, can indicate distinct underlying issues such as lead malfunction, tissue interface degradation, or system failure, each with different risk profiles and clinical implications. The Introduction of six new, specific impedance-related codes would enable more precise documentation and analysis of device-related events, thereby improving the overall quality of incident reporting. This enhancement would also align the IMDRF coding system with current clinical practice, where distinct impedance types are routinely separately ultimately enhancing patient safety and ensuring more appropriate regulatory evaluation. | Additional more specific term to low impedance A072202 | Pacemaker lead shows low out of range pacing impedance (<200 ohms). OR Tachycardia lead shows decrease of the pacing impedance. OR CRTD lead shows decreased pacing impedance | DUPLICATE | Duplicate of row 121
330 | 2025-08-31 | MedTech Europe | A | 2025 | A072206 | Unstable Stimulation Impedance | Level 3 | The stimulation impedance varies without being consistently too high or too low. | Add | add the term 'unstable stimulation impedance' | The current IMDRF coding system includes only a single, undifferentiated category for impedance-related issues, which lacks the specificity needed to reflect the diverse clinical and technical contexts in which impedance is monitored across medical devices. Impedance abnormalities, whether in stimulation, sensing, or therapeutic delivery, can indicate distinct underlying issues such as lead malfunction, tissue interface degradation, or system failure, each with different risk profiles and clinical implications. The Introduction of six new, specific impedance-related codes would enable more precise documentation and analysis of device-related events, thereby improving the overall quality of incident reporting. This enhancement would also align the IMDRF coding system with current clinical practice, where distinct impedance types are routinely separately ultimately enhancing patient safety and ensuring more appropriate regulatory evaluation. | Additional more specific term to unstable impedance A072203 | Pacemaker lead shows fluctuating pacing impedance. OR Tachycardia lead shows fluctuation of the pacing impedance. OR CRTD lead shows fluctuating pacing impedance | DUPLICATE | Duplicate of row 122
331 | 2025-08-31 | MedTech Europe | A | 2025 | A072207 | High Shock Impedance | Level 3 | Problem associated with higher than intended shock impedance levels | Add | add the term 'high shock impedance' | The current IMDRF coding system includes only a single, undifferentiated category for impedance-related issues, which lacks the specificity needed to reflect the diverse clinical and technical contexts in which impedance is monitored across medical devices. Impedance abnormalities, whether in stimulation, sensing, or therapeutic delivery, can indicate distinct underlying issues such as lead malfunction, tissue interface degradation, or system failure, each with different risk profiles and clinical implications. The Introduction of six new, specific impedance-related codes would enable more precise documentation and analysis of device-related events, thereby improving the overall quality of incident reporting. This enhancement would also align the IMDRF coding system with current clinical practice, where distinct impedance types are routinely separately ultimately enhancing patient safety and ensuring more appropriate regulatory evaluation. | Additional more specific term to high impedance A072201 | ICD lead shows high out of range shock impedance (>150 ohms). | DUPLICATE | Duplicate of row 123
332 | 2025-08-31 | MedTech Europe | A | 2025 | A072208 | Low Shock Impedance | Level 3 | Problem associated with lower than intended shock impedance levels | Add | add the term 'low shock impedance' | The current IMDRF coding system includes only a single, undifferentiated category for impedance-related issues, which lacks the specificity needed to reflect the diverse clinical and technical contexts in which impedance is monitored across medical devices. Impedance abnormalities, whether in stimulation, sensing, or therapeutic delivery, can indicate distinct underlying issues such as lead malfunction, tissue interface degradation, or system failure, each with different risk profiles and clinical implications. The Introduction of six new, specific impedance-related codes would enable more precise documentation and analysis of device-related events, thereby improving the overall quality of incident reporting. This enhancement would also align the IMDRF coding system with current clinical practice, where distinct impedance types are routinely separately ultimately enhancing patient safety and ensuring more appropriate regulatory evaluation. | Additional more specific term to low impedance A072202 | ICD lead shows low out of range shock impedance (<25 ohms). | DUPLICATE | Duplicate of row 124
333 | 2025-08-31 | MedTech Europe | A | 2025 | A072209 | Unstable Shock Impedance | Level 3 | The shock impedance varies without being consistently too high or too low. | Add | add the term 'unstable shock impedance' | The current IMDRF coding system includes only a single, undifferentiated category for impedance-related issues, which lacks the specificity needed to reflect the diverse clinical and technical contexts in which impedance is monitored across medical devices. Impedance abnormalities, whether in stimulation, sensing, or therapeutic delivery, can indicate distinct underlying issues such as lead malfunction, tissue interface degradation, or system failure, each with different risk profiles and clinical implications. The Introduction of six new, specific impedance-related codes would enable more precise documentation and analysis of device-related events, thereby improving the overall quality of incident reporting. This enhancement would also align the IMDRF coding system with current clinical practice, where distinct impedance types are routinely separately ultimately enhancing patient safety and ensuring more appropriate regulatory evaluation. | Additional more specific term to unstable impedance A072203 | ICD lead shows fluctuating shock impedance. | DUPLICATE | Duplicate of row 125
334 | 2025-08-22 | AbbVie | E | 2025 | E23 | Patient Subjective Dissatification of Cosmetic Result | Level 2 | Perception issue by patient. Cosmetic result does not match expectations. Unable to objectively confirm. No device or procedure problems. | Add | New Term | No current IMDRF code. Needed for breast implant and dermal products. And this does not align w/ "Lack of Effect" since this is a patient perception issue. Aligns with MedDRA 10076571 Patient dissatisfaction with treatment | None | Patients dissatisfied with outcome of breast implant and other devices. | REJECT | This is not a clinical sign or symptom
335 | 2025-08-22 | AbbVie | E | 2025 | E21xx | Breast Implant Failure | Level 2 | Degredation of silicone shell due to normal aging of implant. | Add | New Term | Needed since the E2107¦ Failure of Implant code was retired 2025. | None | Scenerios when implant is replaced/removed as a result of implant rupture due normal aging of the silicone shell. Rupture is a device issue and also a harm. | REJECT | This is not a clinical sign or symptom
336 | 2025-08-22 | AbbVie | E | 2025 | A0206xx | Missing Gel Stent | Level 3 | Missing gel stent | Add | New Term | No current IMDRF code Needed for XEN product | None | Need to differentiate between the injector and XEN device | REJECT | This is not a clinical sign or symptom. A020602 would already cover this as 'Component missing'
337 | 2025-08-22 | AbbVie | A | 2025 | Axx | Customer Dissatisfaction | Level 1 | Customer dissatified with outcome of procedure without a device failure. | Add | New Term | No current IMDRF code | None | Apply to complaints that are not Lack of Effect. | REJECT | This is not a manifestation of a failure. If there was no malfunction reported, then this could be coded using A25. If a failure manifested at the device level, then the issue with the device should be coded using one of the existing A terms.
338 | 2025-08-22 | AbbVie | A | 2025 | Axx | Counterfeit | Level 1 | Suspicion and/or problem associated with a device identified as counterfeit, meaning it is deliberately and fraudulently misrepresented with respect to its identity, origin, or source, including labeling, packaging, or product content. This includes devices that are fake, unauthorized copies, or those produced or distributed without the required approvals or oversight. | Add | New Term | No current IMDRF code for counterfeit. | None | Applicable for Dermal Fillers | REJECT | This is not a manifestation of a failure. If there was no malfunction reported, then this could be coded using A25. If a failure manifested at the device level, then the issue with the device should be coded using one of the existing A terms.
339 | 2025-08-22 | AbbVie | A | 2025 | Axx | Diversion | Level 1 | Suspicion and/or problem associated with the distribution, sale, or use of a device in an unauthorized or unintended location, market, or channel. This includes any situation where a product is sold or made available outside of approved distribution pathways, resulting in the device being used in a setting not authorized by the manufacturer or regulatory authority. | Add | New Term | No current IMDRF code for diversion. | None | Applicable for Dermal Fillers | REJECT | This is not a manifestation of a failure. If there was no malfunction reported, then this could be coded using A25. If a failure manifested at the device level, then the issue with the device should be coded using one of the existing A terms.
340 | 2025-08-22 | AbbVie | A | 2025 | A0105xx | Lack of Effect Duration | Level 3 | Problem associated with a device or product exhibiting efficacy loss earlier than expected, specifically where the duration of the intended effect is shorter than specified or anticipated according to product labeling. | Add | New Term | Change needed to capture complaints that had effect, however the duration may not have lasted as long as expected | None | Applicable for Dermal Fillers | REJECT | Too granular, can be covered by existing A codes with their current definitions. Combination of these A codes with the F codes will cover the situations described in the rationale column.
341 | 2025-08-22 | AbbVie | A | 2025 | A0201xx | Label Quality Problem | Level 3 | Problem associated with the label that does not meet the specifications or requirements for which it was manufactured (e.g. materials, parts, manufacturing process). | Add | New Term | Change needed to capture labels that are not placed correctly within machine during manufacturing | None | Applicable for Dermal Fillers | REJECT | Too granular, can be covered by existing A codes with their current definitions. Combination of these A codes with the F codes will cover the situations described in the rationale column.
342 | 2025-08-22 | AbbVie | A | 2025 | A0201xx | Expired Device or Device Components Shipped | Level 3 | Problem associated with the shipment of expired device or device components from the manufacturer. | Add | New Term | No current IMDRF code for expired device shipped | None | Applicable for all devices | REJECT | Too granular, can be covered by existing A codes with their current definitions. Combination of these A codes with the F codes will cover the situations described in the rationale column.
343 | 2025-08-22 | AbbVie | A | 2025 | A0205xx | Delamination of packaging | Level 3 | Delamination of packaging material by separation into layers within a single piece of material (e.g., delamination of Tyvek lid). | Add | New Term | No current IMDRF code and speficity needed for opening packaging with delamination of the tyvek lid. | None | Applicable to Tyvek lid; have received complaints of material delamination. | REJECT | Too granular, can be covered by existing A codes with their current definitions. Combination of these A codes with the F codes will cover the situations described in the rationale column.
344 | 2025-08-22 | AbbVie | A | 2025 | A0205xx | Packaging Discolored | Level 3 | Problem associated with an undesired streak, pattern and/or a noticeable change in color from the rest of the materials used in the packaging construction. | Add | New Term | No current IMDRF code and speficity needed for opening packaging with discoloration of the tyvek lid. | None | Applicable to Tyvek lid; have received complaints of material discoloration. | REJECT | Too granular, can be covered by existing A codes with their current definitions. Combination of these A codes with the F codes will cover the situations described in the rationale column.
345 | 2025-08-22 | AbbVie | E | 2025 | E10XX | Gastroenteritis | Level 2 | An inflammation of the lining of the stomach and intestines. | Add | New Term | No current IMDRF code Aligns with MedDRA 10017888 Gastroenteritis | None | Had a case for dermal filler. | REJECT | Based on the definition, this could be coded as E1012 or E1036
346 | 2025-08-31 | MedTech Europe | A | 2025 | A140506 | Intermittent Infusion | A1405 | Problem associated with the infusion not being stead- characterised by intermittent stoppages to the flow. | Modify | modification of term to include "flow" : "Intermittent Infusion or flow", and to the definition to make it widder than infusion problems only "Problem associated with the flow not being stead- characterised by intermittent stoppages to the flow." | This code is under A1405 "Improper Flow or Infusion" but it is referencing infusion only. There is no code for intermittent flow that are not related to infusion (for example intermittent irrigation). | no impact | For an ultrasonic surgical aspirator system that has 3 functions: fragmentation of the tissue, irrigation and aspiration (suction of the irrigation fluid and fragmented tisssue), the user reported that the irrigation was working intermittently. | DUPLICATE | Duplicate of row 126
347 | 2025-08-31 | MedTech Europe | A | 2025 | A210102 | Illegible Information | Level 3 | Problem associated with information unable to be read or deciphered | Modify | Change term to: Illegible or Faded Information. Change definition to: Problem associated with information difficult or impossible to be read or deciphered. | The level 2 term, A2101 - Device Markings/Labelling Problem, includes in its definition a sentence related to markings/labelings 'worn out'. A2101 reads: "... Problems can include but are not limited to this material being unclear, missing, worn out". The level 3 term, A210102 – Illegible Information, and its definition, does not allow for difficulties to read the markings when it comes to the markings being worn out. In consequence, the proposal includes that this definition is broaden so it can also accommodate for markings difficult to be read. | None | The laser lines on the instruments are worn and difficult to see clearly. They appear faded to a gray color from the original black color. | DUPLICATE | Duplicate of row 127
348 | 2025-08-31 | MedTech Europe | A | 2025 | New level 3 code under A0205 | Incorrect product | 3rd level | Problem associated with a device whose identity does not match to what is written on the label. | Add | Add new level 3 sub-code under A0205. | There is no suitable code/term available: Product mix-ups can be caused by multiple factors: incorrect product was introduced in the correct package; incorrect label was affixed to the correct product; or incorrect manufacturing programming was selected, rendering an incorrect product. | None | Product A was ordered. The product received, showed it was product A; however, when opened, the product's identity was B. | REJECT | A210106 can be used.
349 | 2025-08-31 | MedTech Europe | A | 2025 | New level 3 code under A0406 | Thinning of Material | 3rd level | Device is thinned due to deformation. | Add | Add new level 3 sub-code under A0406, Term: 'Thinning of Material', Description 'Device is thinned due to deformation. | There is a code for Thickening of Material (A040612). (Alternatively, modify A040611 - Stretched to include thinning of Material). | None | Addition only for consistency. | DUPLICATE | Duplicate of row 129
350 | 2025-08-31 | MedTech Europe | B | 2025 | New level 1 B code | Risk Assessment Review | Level 1 | The investigation involved the review and analysis of risk files of the device to support the identification of possible causes for the adverse event and assess whether the occurrence and severity are within anticipated parameters. It should be noted that risk assessment review typically is not considered sufficient as a stand-alone method, but should be used in conjunction with other investigation methods for providing for instance complementary information. | Add | Addition, see left. | ISO 13485 integrates a risk-based approach throughout the entire lifecycle of medical devices, including post-market activities. This standard mandates a proactive process of identifying, evaluating, and controlling risks to ensure patient safety. The change is necessary to accommodate an investigation task that is constantly used to ensure the medical device is safe and effective. MIR form 7.3.1 section 4.2 – ‘Cause of investigation and conclusion’ asks for risk assessment review, we would like to propose coding all types investigations with Annex B instead of in other sections of the MIR form. | None | Risk management review is used in investigation of quality complaints and adverse events to identify whether the event is already included in the risk files. If it was not, it drives its inclusion; if it was, an assessment is conducted to determine whether the severity, occurrence rate and probability of harm are within the anticipated levels; if it is not, the manufacturer takes action. | DUPLICATE | Duplicate of proposal 130
351 | 2025-08-31 | MedTech Europe | B | 2025 | New level 1 B code | Medical/Clinical Review | Level 1 | The investigation involved the analysis of the medical documentation (radiological images, lab tests, medical records, surgical notes, pathology reports, etc.) by a healthcare professional to support the identification of possible causes for the adverse event. It should be noted that medical/clinical review typically is not considered sufficient as a stand-alone method, but should be used in conjunction with other investigation methods for providing for instance complementary information. | Add | Addition, see left. | This proposed term is similar to B15; however, we are proposing that there is a separate code for a clinical/medical investigation because it must be conducted by a trained healthcare professional. The medical investigation is necessary to use D1001 - Adverse Event Related to Patient Anatomy and/or Conditions. This assessment does not fit under B15, as we would like to specifically call out a review by a healthcare professional. Of note, US/AU regulations include clauses/exceptions for not reporting certain adverse events based on medical judgement from a healthcare professional. | None | The code will be used in the investigation of any adverse event where there is patient harm and the user provides medical/clinical information (images, reports, tests, analysis, notes) for the manufacturer to assess so a root cause can be established. | DUPLICATE | Duplicate of proposal 131
352 | 2025-08-31 | MedTech Europe | C | 2025 | C0602 | Inappropriate Material | 2nd level | Problems that occur due to the presence of a material that should not be present or part of the device. | Modify | Modify definition of C0602 to: "Problems that occur due to undesired presence, composition or structure of a material." | Child terms of C0602 do mention material structure and composition and presence of undesired material. The proposed change to the parent term derscription just includes all those options. | None | N/A | DUPLICATE | Duplicate of row 136
353 | 2025-08-31 | MedTech Europe | C | 2025 | C160503 | Packaging Contains Unintended Material | 3rd level | Problems that occurred because unintended material was packaged with the device. | Modify | Modify definition: Unintended material was packaged with the device. | Remove word "problem", as not all such investigation findings are considered a problem of the product (i.e. when it meets specifications) | None | Particulate in a sterile package is reported however this meets the specifications defined by the manufacturer. | DUPLICATE | Duplicate of row 137
354 | 2025-08-31 | MedTech Europe | D | 2025 | New child of D14 | Product meets specifications | 2nd level | The reported event is substantiated but determined to not be a deficiency of the product. | Add | Add new level 2 sub-code under D14: 'Product Meets Specifications', definition: 'The reported event is substantiated but determined to not be a deficiency of the product.' | Include a new code to address situations when a reported event is observed, however it is not a related to a product deficiency. | None | Particulate in a sterile package is reported however this met the specifications defined by the manufacturer. | DUPLICATE | Duplicate of row 138
355 | 2025-08-31 | MedTech Europe | E | 2025 | E07XX | Hypercapnia | Respiratory System | A condition of abnormally elevated carbon dioxide (CO2) levels in the blood (MEDDRA/10020591) | Add | term added | It was reported to us in conjunction with procedural problems during knee arthroplasty. The symptom could not be coded as it is not covered in Annex E. | none | It was reported to us in conjunction with procedural problems during knee arthroplasty. | DUPLICATE | Duplicate of row 143
356 | 2025-08-31 | MedTech Europe | E | 2025 | E07XX | Tidal volume decreased | Respiratory System | Using a smaller than usual amount of air (tidal volume) with each breath delivered to a patient on a ventilator (MEDDRA/10078272) | Add | term added | It was reported to us in conjunction with procedural problems during knee arthroplasty. The symptom could not be coded as it is not covered in Annex E. | none | It was reported to us in conjunction with procedural problems during knee arthroplasty. | DUPLICATE | Duplicate of row 144
357 | 2025-08-31 | MedTech Europe | E | 2025 | E1649 | Crepitus | Musculoskeletal System | Popping, cracking, or clicking sound in a joint. | Modify | modification of definition | Current definition "Grating sound or sensation produced by friction between e.g. bone, cartilage, fractured parts of a bone." is from the Oxford dictionary. In the MSD manual ist is described as "popping, cracking, or clicking sound in a joint". Use of "crepitus" in this article https://pmc.ncbi.nlm.nih.gov/articles/PMC5851845/. | none | A clicking noise was reported after knee arthroplasty. | DUPLICATE | Duplicate of row 145
358 | 2025-08-31 | MedTech Europe | F | 2025 | New level 2 term under F26 | Potential Serious Risk (No Harm Occurred) | Level 2 | No harm occurred in relation to the adverse event. However, the event involved a device malfunction or use condition that, if repeated, could reasonably be expected to result in a serious adverse health impact (e.g., death, serious deterioration of health, or serious public health threat). (For Exposure to Contaminated Device/Risk of Infection use F1007) | Add | Add a level 2 term under F26. | Current IMDRF codes do not adequately capture adverse events that meet the MDR definition of a serious incident due to their potential consequences, even if no harm occurred. This term would fill a gap in reporting serious risks that were averted but remain relevant for vigilance and preventive action. | None | In 2023, a serious incident was reported to us in which a blunt scissor led to a serious complication during labour. A year later, the same customer reported a similar incident with a blunt scissors in which no complications occurred. Nevertheless, the incident was reported as a serious incident because there was a risk when using the product. | DUPLICATE | Duplicate. See row 146
359 | 2025-08-31 | MedTech Europe | E | 2025 | E0819xx | Conjunctival retraction | Level 3 | The abnormal pulling back or shrinking of the conjunctiva | Add |  | For ocular devices, this is an adverse event that we see in clinical trails, but the closest current E code is 'E0819-Eye Injury'. Unfortunately this does not describe the adverse event specifically enough for reporting or risk management purposes. It should also be noted that there is a corresponding Meddra code 'Conjunctival retraction-10079837''. | None | 1) Patient presented with conjunctival retraction. The implant is still covered so refill was done. If retraction worsens, will have to consider explantation. 2) Noticed conjunctival retraction upon slit lamp exam by investigator Most commonly related to implant and/or implantation procedure | DUPLICATE | Duplicate. See row 147
360 | 2025-08-31 | MedTech Europe | E | 2025 | E08xxxx | Conjunctival Bleb | Level 3 | A conjunctival bleb is a blister-like elevation on the conjunctiva due to escape of fluid under the conjunctiva | Add |  | Currently there are no terms that describe Conjunctival bleb or terms that closely represent this events. Also Conjunctival bleb is frequently reported with use of an ocular implant. It should also be noted that there is a corresponding Meddra Code: Conjunctival bleb- 10052114 |  | 1) Principal investigator had refill needle seated, assumed to be in the implant, he felt resistance and noticed that about 0.01 ml of drug was pushed under the conjunctiva forming a bleb 2) Principal investigator noted a conjunctival bleb overlying the implant during examination. | DUPLICATE | Duplicate. See row 148
361 | 2025-08-31 | MedTech Europe | E | 2025 | E08xxxx | Conjunctival filtering bleb leak | Level 3 | Escape of fluid from a conjunctival bleb | Add |  | Currently there are no terms that describe Conjunctival filtering bleb leak or terms that closely represent this events. Also Conjunctival filtering bleb leak is frequently reported with use of an ocular implant. It should also be noted that there is a corresponding Meddra code: Conjunctival filtering bleb leak- 10069165 |  | 1) Physician who perfomed week 120 refill contacted the medical monitor when he noticed mild leak of conjunctiva overlying implant and was recommended close monitoring for filtering bleb leak. | DUPLICATE | Duplicate. See row 149
362 | 2025-08-31 | MedTech Europe | E | 2025 | E08xxxx | Conjunctival bleb | Level 3 | A conjunctival bleb is a blister-like elevation on the conjunctiva due to escape of fluid under the conjunctiva | Add |  | For occular devices, this is an adverse event that we see in clinical trails, but the closest current E code is 'E0819-Eye Injury'. Unfortunately this does not describe the adverse event specifically enough for reporting or risk management purposes. It should also be noted that there is a corresponding Meddra code: Conjunctival bleb- 10052114; Conjunctival filtering bleb leak- 10069165 | None | 1) Principal investigator had refill needle seated, assumed to be in the implant, he felt resistance and noticed that about 0.01 ml of drug was pushed under the conjunctiva forming a bleb 2) Principal investigator noted a conjunctival bleb overlying the implant during examination. 3) Physician who perfomed week 120 refill contacted the medical monitor when he noticed mild leak of conjunctiva overlying implant and was recommended close monitoring for filtering bleb leak. | DUPLICATE | Duplicate. See row 150
363 | 2025-08-31 | MedTech Europe | E | 2025 | E08xx or E0819xx | Conjunctival Erosion | Level 2 or 3 | A localized area of damage, thinning, or breakdown of the conjunctival surface | Add |  | For occular devices, this is an adverse event that we see in clinical trails, but the closest current E code is either 'E0819-Eye Injury' or 'E2006-Erosion'. Unfortunately these do not describe the adverse event specifically enough for reporting or risk management purposes. It should also be noted that there is a corresponding Meddra code 'Conjunctival erosion-10010716'. | None | Subject presented for a second unscheduled visit after being diagnosed with a conjunctival bleb post implantation procedure. Sub-investigator determined that the subject had a resolved conjunctival bleb, and diagnosed her with a moderate focal conjunctival avascularity. Per query from system, site is clarifying the diagnosis to a focal conjunctival erosion | DUPLICATE | Duplicate. See row 151
364 | 2025-08-31 | MedTech Europe | E | 2025 | E08xx | Tenon Capsule Thickening | Level 2 | An abnormal thickening of Tenon’s capsule | Add |  | Currently there are no terms that describe Tenon capsule thickening or terms that closely represent this event. Meddra Code Tenon's capsule thickening- 10087613 | None | The sub-investigator noted tenon thickening during refill attempt that was unsuccessful. Unable to insert refill needle into implant septum. Most commonly related to implant and/or implantation procedure | DUPLICATE | Duplicate. See row 152
365 | 2025-08-31 | MedTech Europe | E | 2025 | E08xx | Subconjunctival Thickening | Level 2 | An abnormal increase in the density or bulk of the tissue beneath the conjunctiva | Add |  | Currently there are no terms that describe subconjunctival thickening or terms that closely represent this event. Currently no specific Meddra Code. Closest is 10061446 Conjunctival disorder | None | Principal investigator pushed the syringe plunger with resistance and no reflux exchange was noticed due to subconjunctivval thickening over implant. | DUPLICATE | Duplicate. See row 153
366 | 2025-08-31 | MedTech Europe | E | 2025 | E0819xx | Scleral thinning | Level 3 | A condition where the sclera, the white, tough outer layer of the eye that provides structural support and protection, becomes abnormally thin | Add |  | For ocular devices, this is an adverse event that we see in clinical trials, but the closest current E code is 'E0819-Eye Injury'. Unfortunately this does not describe the adverse event specifically enough for reporting or risk management purposes. | None | Patient presented with complaints of ocular discomfort, photophobia and mild foreign body sensation in the right eye for 4 days. On examination localized scleral thinning was noted in the superior temporal quadrant of the right eye . | DUPLICATE | Duplicate. See row 154
367 | 2025-08-31 | MedTech Europe | E | 2025 | E2404 | Death (Cause Unknown) | Level 2 | The cessation of all biological functions that sustain life, where the exact reason or underlying condition leading to death has not been determined or identified | Add |  | Death (cause unknown): Death (cause unknown) can be considered an harm for the cases where non information about the causes of death are available Medrra 10011906 LLT- 10063378 | None | Subject's daughter called to report patient had died. Records requested. Unknown cause of death. | DUPLICATE | Duplicate. See row 155
368 | 2025-08-31 | MedTech Europe | E | 2025 | E240101 | Health Effect Not Yet Characterized | Level 3 | It is not clear whether any health effects occurred, or health effect appears to have occurred but there is not yet enough information available to classify the clinical signs, symptoms and conditions. | Add |  | The current IMDRF coding system does not adequately differentiate the reasons behind incomplete event information, limiting the clarity and usefulness of incident reporting. To address this, we propose three new codes: Health Effect Not Yet Characterized, Unsuccessful Data Collection, and Reporter Declined to Provide Information. These additions will enhance transparency, improve documentation accuracy, and better reflect real-world reporting scenarios—ultimately supporting more reliable analysis and regulatory decision-making. | None |  | DUPLICATE | Duplicate. See row 156
369 | 2025-08-31 | MedTech Europe | E | 2025 | E240102 | Unsuccessful Data collection | Level 3 | Despite repeated inquiries, no further information about the events was provided. | Add |  | The current IMDRF coding system does not adequately differentiate the reasons behind incomplete event information, limiting the clarity and usefulness of incident reporting. To address this, we propose three new codes: Health Effect Not Yet Characterized, Unsuccessful Data Collection, and Reporter Declined to Provide Information. These additions will enhance transparency, improve documentation accuracy, and better reflect real-world reporting scenarios—ultimately supporting more reliable analysis and regulatory decision-making. | None |  | DUPLICATE | Duplicate. See row 157
370 | 2025-08-31 | MedTech Europe | E | 2025 | E240103 | Reporter declined to provide information | Level 3 | Reporter has been contacted but explicitly declined to provide further information. | Add |  | The current IMDRF coding system does not adequately differentiate the reasons behind incomplete event information, limiting the clarity and usefulness of incident reporting. To address this, we propose three new codes: Health Effect Not Yet Characterized, Unsuccessful Data Collection, and Reporter Declined to Provide Information. These additions will enhance transparency, improve documentation accuracy, and better reflect real-world reporting scenarios—ultimately supporting more reliable analysis and regulatory decision-making. | None |  | DUPLICATE | Duplicate. See row 158
371 | 2025-08-31 | MedTech Europe | E | 2025 | E2111 | Injection site reaction | Level 2 | A reaction (usually immunologic) developing at the site of injection | Modify | Remove (Intense) from the definition | The word "intense" is limiting the use of this code. "intense" tends to be aligned with harms with higher severity. ''Injection site reactions'' leading minor harms are excluded from this term due to the word intens doesn't represent the severity of minor reactions. | None | A day after administration, physician noticed mild erythema around the injection site. | DUPLICATE | Duplicate. See row 159
372 | 2025-08-31 | MedTech Europe | E | 2025 | E1725 | Injection Site Pooling | Level 2 | Liquid or a pocket of air trapped under the skin, typically due to injection problems. Sometimes refers to Subcutaneous Emphysema. Does not include inflammation. | Modify | Change the Term Definition and add "Subcutaneos Emphysema'' to it | The definition of injection site pooling is overlapping with the definition for Subcutaneous emphysema. Please, modify the Term definition of E1725; Include SubCutaneos Emphysema instead. | None | Patient received an intramuscular injection of drug X in the left deltoid region at a local clinic. Withing 4 hours the patient reported crackling sensation and mild swelling at the injection site. On examination at the emergency department, crepitus was plapable on the upper arm without signs of erythema or warmth. X-ray confirmed the presence of subcutaneous emphysema localised to the soft tissues at and around the injection site. | DUPLICATE | Duplicate. See row 160
373 | 2025-08-31 | MedTech Europe | E | 2025 | E21xx | Infusion site reaction | Level 2 | A localized response at the site where a medication or fluid is administered through an intravenous (IV), subcutaneous, or other infusion route | Add |  | Infusion site reaction is intented to cover the reactions during intravenous delivery using infusion devices. This term differs from the E2111 "injection site reaction" because the latter is supposed to cover the subcutaneous reactions. | None | Approximately an hour post-infusion, the patient reported a reaction at the infusion site | DUPLICATE | Duplicate of row 161
374 | 2025-08-31 | MedTech Europe | E | 2025 | new level 2 code under E10 | Abdominal compartment syndrome | Level 2 | Abdominal compartment syndrome (ACS) is a condition that occurs when bleeding or swelling leads to rises of pressure in the abdominal cavity that causes end-organ damage, dysfunction or failure. | Add | Add term and description as suggested on left. | No code/term available to classify a known complication of e.g. NPWT (negative pressure wound therapy). Meddra: 10058808 | E10 | Patient had abdominal compartment syndrome during treatment of open abdomen with negative pressure wound therapy (device) to achieve fascial closure. | DUPLICATE | Duplicate of row 162
375 | 2025-08-31 | MedTech Europe | E | 2025 | New level 2 code under E21 | Crust or scab | 2nd level | Crusts are dried serum, pus, or blood, usually mixed with epithelial and sometimes bacterial debris. | Add | Addition, see left. | Received reports of significance and a suitable term/code was unavailable (Meddra: 10039509 - Scab, 10076524 - Nasal crusting). | None | Nasal crusting following treatment of turbinate hypertrophy (e.g. endoscopic or ablation device) can require re-intervention or may take months to resolve on its own. | DUPLICATE | Duplicate of row 163
376 | 2025-08-31 | MedTech Europe | F | 2025 | F31 | Cancelled/Aborted/ Treatment/Therapy or No Treatment/Therapy performed | Level 1 term | Patient treatment aborted, or unable to be completed. This applies to cases where there is no information whether the treatment will be carried out at a later time, or where it is known that treatment will not go ahead. If it is known that the treatment has been, or will be, carried out later then use F05 - Delay to Treatment/ Therapy | Modify | Change termonoligy and perhaps add additional level 2 term to F31. | Change term to Cancelled/aborted Treatment/Therapy or no Treatment/Therapy performed. Perhaps Add additional level to specify between cancelled/aborted and not perfomed. The current F31 - Cancelled/Aborted Treatment/Therapy is not sufficient if it is reported that eiter the patient denied treatment or no treatment was performed. | No impact | 1. Cancelled/aborted treatment/Therapy: Patient treatment aborted, or unable to be completed. This applies to cases where there is no information whether the treatment will be carried out at a later time, or where it is known that treatment will not go ahead. 2. No treatment/Therapy performed: No intervention has been planned/performed to treat the reported adverse event. (eg. At the patient's request, it was decided not to perform additional treatments or a wait and see aproach have been desided by health professional) | DUPLICATE | Duplicate. See row 164
377 | 2025-08-31 | MedTech Europe | G | 2025 | G07004 | No Defective Part/Component/Sub-assembly | Level 2 | The device and/or device data was analyzed and found to be functioning properly. There was no product malfunction. This term is used when a device-related component cannot be assigned, as the reported event is not attributable to a failure of the overall device or its subcomponents. | Add | add the term 'No Defective Part/Component/Sub-assembly' | The change is necessary to enable appropriate classification of complaint cases in which the device was analyzed and found to be operationg properly, with no device malfunction or component failure. Currently, no existing IMDRF Annex G term accuratly captures this scenario. While G07001 ("Part/Component/Sub-assembly not applicable") is inteded for cases involving the overall device, it implies that the device caused the event. The absence of a suitable code for "in Specification" outcomes can lead to inconsistent classification. | Additional more specific term to Annex G: G07 - Other Terms/Codes Related to Components | It was reported that an ICD activated the elective replacement indicator (ERI) prematurely. A thorough analysis of the device data and the returned ICD confirmed that the device was functioning properly. The activation of the ERI battery status was due to high-output programming, which led to increased energy consumption. No device malfunction or component failure was present. | DUPLICATE | Duplicate of row 165
378 | 2025-08-31 | MedTech Europe | G | 2025 | G07005 | Part/Component/Sub-assembly Not Determinable | Level 2 | The device and/or device data was analyzed and the reported event could be confirmed. However, no deviation or failure could be identified. The root cause remains unknwon despite all available information. This term is used when the event cannot be attributed to any specific component nor explained by external factors or influences. | Add | add the term 'Part/Component/Sub-assembly Not Determinable' | The change is necessary to support the classification of complaints for which a thorough investigation has been completed, but no failure, malfunction or identifiable root cause could be determined. The existing IMDRF code G07003 (“Insufficient Component Information”) is restricted to situations where further information is expected or needed. Likewise, G07001 assumes the device is affected overall, which is inappropriate when no impact or component association can be established. In such cases, the only available option is G07002, which should be reserved for cases requiring a new code proposal. G07005 addresses this specific gap by providing a term for concluded investigations with unknwon root cause. | Additional more specific term to Annex G: G07 - Other Terms/Codes Related to Components | It was reported that an ICD activated the elective replacement indicator (ERI) prematurly and unexpectedly. The data confirmed the premature battery depletion. However, a thorough analysis of the device data and the device itself, including the electronic module, the battery and other relevant device components, revealed no malfunction. | DUPLICATE | Duplicate of row 166
379 | 2025-08-31 | MedTech Europe | G | 2025 | G07003 | Insufficient Component Information | Level 2 | There is not enough information available to classify the medical device component or there is not enough information available to determine whether a medical device component is affected. | Modify | modify the description of "Insufficient Component Information" | The change is necessary to remove the word “yet” which implies that additional information is expected, which is not consistently the case. In addition, to support the classification of complaints for which it is not possible to determine whether a device component is at all affected. | Broadening term description to Annex G07003 - Insufficient Component Information | It was reported that the elective replacement indicator (ERI) was activated. However, no device or data are available for analysis. Based solely on the complaint description it is not possible to determine whether the ERI activation is as expected or prematurely. Therefore, no relationship to a device component can be established. | DUPLICATE | Duplicate of row 167
380 | 2025-10-15 | IMDRF AET WG | G | 2025 | G04068 | Holder | Level 2 | A component designed to hold another object | Modify | Add 'bracket' to definition to clarify difference between retainer and holder. |  |  |  | ACCEPT | Add 'bracket' to definition to clarify difference between retainer and holder. |  | New Definition: A component, such as a bracket, designed to hold another object
381 | 2025-10-15 | IMDRF AET WG | E | 2025 | E21XXX | Ocular Procedural Complication | Level 2 | Any complication arising from ocular surgical procedures involving the conjunctiva or adjacent tissues, including but not limited to subconjunctival bleb formation (e.g., trabeculectomy, glaucoma filtration procedures). This may encompass issues such as bleb absence, infection, leakage, perforation, structural thinning, scarring, or other adverse changes affecting the surgical site or surrounding ocular structures. | Add | [6 empty] | E2135 | Ocular procedural complication:10080964
382 | 2025-10-01 | IMDRF AET WG | A | 2025 | A1003 | Insufficient Cooling | Level 2 | Problem associated with the device insufficiently cooled in device active (working) or/and non-active (nonworking) state. | Modify | Modify definition to be more aligned with other Level 2 terms of similar nature, e.g. A1001, A1002 | Modify definition to be more aligned with other Level 2 terms of similar nature, e.g. A1001, A1002 | No impact | No impact | ACCEPT | Minor editorial change |  | Problem associated with the device, which has a cooling function, producing temperatures that are not as low as specified.
383 | 2025-10-24 | PMDA | G | 2025 | G02014 | Display |  | A component designed to present information visually. In the case of Touchscreen, use G0204004. | Modify |  |  | Editorial | Editorial | REJECT | No change required. |  | A component designed to present information visually. In the case of Touchscreen, use G0204004.
384 | 2025-10-24 | PMDA | G | 2025 | G04097 | Plate |  | A thin, flat sheet or strip used to join, strengthen or to form parts of another structure. For a flat, shallow container with a raised rim, use G03016 - Tray. | Modify |  |  | Editorial | Editorial | ACCEPT |  |  | A thin, flat sheet or strip used to join, strengthen or to form parts of another structure. For a flat, shallow container with a raised rim, use G03016
385 | 2025-10-24 | PMDA | C | 2025 | C0107 | Unintended Presence of Allergens | Level 2 | Unintended or unexpected presence of allergens in the device. If the presence of the allergen is expected but not adequately labelled, then use "Labelling Problem". | Modify | Unintended or unexpected presence of allergens in the device. If the presence of the allergen is expected but not adequately labelled, then use "Labelling Problem". |  | Editorial | Editorial | REJECT | Duplicate of 195
Note on P16: BfArM:
Note on P17: BfArM:

## Lists (hidden)
Annex | Change | Outcome
A | Add | ACCEPTED
B | Delete | REJECTED
C | Modify
D
E
F
G