## A
Annex Name: Annex A
Annex Title: Medical Device Problem
Release Number: 2026
Annex Description: Terms/codes for describing problems (malfunction, deterioration of function, failure) of medical devices that have occurred in pre- or post-market contexts (e.g. clinical studies, clinical evaluation or post-market surveillance). The MedDRA® trademark is registered by ICH. This information/material is provided "as is" without warranty of any kind. In no event shall the ICH or the authors of the original information/material be liable for any claim, damages or other liability arising from or in connection with the use of this information/material.
Annex Instructions: These terms allow capturing of the problems encountered at device(s) level through observational language without yet describing possible reasons or causes for the problems or failures observed. Annex A provides a comprehensive list of medical device problem terms and codes. It is recognized that not all jurisdictions may want to code to such detailed levels. The hierarchical structure will allow jurisdictions to choose the level of coding to use. IMDRF Terms and Definitions use American spelling. Wherever appropriate the word 'device' should be interpreted as meaning all or any of the following: medical device, in-vitro diagnostic medical device, accessory to a medical device, medical device parts and/or components. Wherever appropriate the word 'patient' should be interpreted as meaning all or any of the following: user, operator, any other person affected by the incident.
Annex Approval Date: January 2026
NOTE: Maximum field length Term: 100 characters Definition: 700 Status description: 700
Level 1 Term | Level 2 Term | Level 3 Term | Code | Definition | Non-IMDRF Code | Status | Status Description | CodeHierarchy
Patient Device Interaction Problem |  |  | A01 | Problem related to the interaction between the patient and the device. | MedDRA:10092649:Patient-device interaction issue |  |  | A01
| Patient-Device Incompatibility |  | A0101 | Problem associated with the interaction between the patient's physiology or anatomy and the device that affects the patient and/or the device. | MedDRA:10092649:Patient-device interaction issue |  |  | A01¦A0101
|  | Biocompatibility | A010101 | Problem associated with undesirable local or systemic effects due to exposure to medical device materials or leachates from those materials by a patient who has an implant or is receiving treatment with a device made from them. |  |  |  | A01¦A0101¦A010101
|  | Device Appears to Trigger Rejection | A010102 | The device appears to elicit undesired response in the patient to the presence of an implanted or invasive device, without inherent device failure, e.g. fibrous encapsulation, or inflammation of the tissue around the device, or extrusion of the device. |  |  |  | A01¦A0101¦A010102
|  | Inadequacy of Device Shape and/or Size | A010103 | The physical size and/or shape of the device was inadequate with regard to the patient's anatomy. |  |  |  | A01¦A0101¦A010103
| Osseointegration Problem |  | A0102 | Problem associated with interconnection between the bone tissue and the implanted device. | MedDRA:10092654:Device osseointegration issue |  |  | A01¦A0102
|  | Failure to Osseointegrate | A010201 | Problem associated with the failure to see direct anchorage of an implant by the formation of bony tissue around the implant without the growth of fibrous tissue at the bone-implant interface. |  |  |  | A01¦A0102¦A010201
|  | Loss of Osseointegration | A010202 | Problem associated with weakened integration of the device at the bone-implant interface due to loss of fibrous and/or bony tissue and leading to compromised anchorage of the device. i.e. Loosening/Lysis |  |  |  | A01¦A0102¦A010202
| Loosening of Implant not Related to Bone-Ingrowth |  | A0103 | Problem associated with the loss of direct anchorage of an implanted device over time or due to an injury. | MedDRA:10076364:Device loosening | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 212 of Change Log. | A01¦A0103
| Migration or Expulsion of Device |  | A0104 | Problem with an implanted or invasive device moving within the body, or being completely expelled from the body. | MedDRA:10012588:Device migration |  |  | A01¦A0104
|  | Expulsion | A010401 | Problem with all or part of an implanted or invasive device being completely expelled from its intended location within the body. |  |  |  | A01¦A0104¦A010401
|  | Migration | A010402 | Problem with all or part of an implanted or invasive device moving from its intended location within the body. |  |  |  | A01¦A0104¦A010402
| Lack of Effect |  | A0105 | Problems associated with a product having less efficacy than expected. No other malfunctions could be identified. | MedDRA:10059875:Device ineffective |  | Term was added on 27 January 2022. For details, see comment No. 23 of the Change Log (Release Number 2022). | A01¦A0105
| Device Stenosis |  | A0106 | Problem associated with the narrowing or obstruction of the device (e.g. prosthetic heart valves, stents, etc.). | MedDRA:10092650:Device stenosis |  | Term was added on 27 January 2022. For details, see comment No. 148, 176 of the Change Log (Release Number 2022). | A01¦A0106
| Implant Subsidence |  | A0107 | Problem related to the settling or collapse of bone adjacent to an implanted device, in direction of the loading force. | MedDRA:10074802:Implant subsidence | Modified (editorial) | Term added for Release Number 2024. See no. 44 of Change Log. Editorial change to term for Release Number 2026. See no. 213 of Change Log. | A01¦A0107
Manufacturing, Packaging or Shipping Problem |  |  | A02 | Problem associated with any deviations from the intended performance of the device that appear to relate to the manufacturing, packaging or shipping processes. | MedDRA:10083763:Device issue prior to use |  | Definition modified for Release Number 2025. See no. 42 of Change Request Log | A02
| Product Quality Problem |  | A0201 | Problem associated with an inherent device characteristic that is not satisfactory as specified or delivered. | MedDRA:10069327:Product quality issue |  |  | A02¦A0201
|  | Dull, Blunt | A020101 | Problem associated with a device not being as sharp as intended or expected. |  |  |  | A02¦A0201¦A020101
|  | Nonstandard Device | A020102 | Problem associated with the device that does not meet the specifications or requirements for which it was manufactured (e.g. materials, parts, manufacturing process). |  |  |  | A02¦A0201¦A020102
| Defective Component |  | A0202 | Problem associated with a device component having flaws of dimensional deviations greater than acceptable for the intended use. |  | Retired (2025) | Term retired for Release Number 2025. Use A020102 - Nonstandard Device or other codes which describe the specific defect with the component. See no. 1 of Change Request Log | A02¦A0202
| Defective Device |  | A0203 | Problem associated with having flaws or dimensional deviations greater than acceptable for the intended use of the device. |  | Retired (2025) | Term retired for Release Number 2025. Use A020102 - Nonstandard Device or other codes which describe the specific defect with the device. See no. 1 of Change Request Log | A02¦A0203
| Device Damaged Prior to Use |  | A0204 | Problem associated with packaging or shipping damage prior to the use of the device. |  | Retired (2025) | Term retired for Release Number 2025. Use A0207 - Device Damaged Prior to Receipt by User. See no. 106 of Change Request Log | A02¦A0204
| Packaging Problem |  | A0205 | Problem associated with the materials used to construct the cover or outer wrapping of the device. | MedDRA:10069405:Product packaging issue |  |  | A02¦A0205
|  | Difficult to Open or Remove Packaging Material | A020501 | Problem associated with difficulty for users to operate the device, specifically as it relates to the opening or removal of the packaging materials |  | Modified (technical) | Definition modified for Release Number 2026. See no. 41 of Change Request Log | A02¦A0205¦A020501
|  | Incomplete or Missing Packaging | A020502 | Problem associated with incomplete or missing packaging that may compromise the device operation as intended. |  | Modified (technical) | Definition modified for Release Number 2026. See no. 42 of Change Request Log | A02¦A0205¦A020502
|  | Unsealed Device Packaging | A020503 | Problem associated with the loss of packaging seal. |  |  |  | A02¦A0205¦A020503
|  | Tear, Rip or Hole in Device Packaging | A020504 | Problem associated with packaging damage (tear, rip or hole) prior to the use of the device. |  |  |  | A02¦A0205¦A020504
|  | Device Packaging Look-Alike | A020505 | Problem associated with medical device packaging that looks similar to another product leading to confusion. The other product may be another device or a different product such as a medicine or cleaning agent. Use A210107 if related to the device, not packaging. |  | Modified (editorial) | Term added for Release Number 2024. See no. 116 of Change Log. Editorial change to term for Release Number 2026. See no. 214 of Change Log. | A02¦A0205¦A020505
| Device Misassembled During Manufacturing /Shipping |  | A0206 | A device found incorrectly assembled when delivered to the user facility. | MedDRA:10092651:Device misassembly during manufacturing |  |  | A02¦A0206
|  | Component Misassembled | A020601 | A device found to have one or more components incorrectly assembled when delivered to the user facility. This includes cases in which components are interchanged. |  |  | Definition was modified on 27 January 2022 (Old definition: Component Misassembled: A device found to have one or more components incorrectly assembled when delivered to the user facility. For details, see comment No. 1 of the Change Log (Release Number 2022). | A02¦A0206¦A020601
|  | Component Missing | A020602 | A device component(s) found to be missing when delivered to the user facility. |  |  |  | A02¦A0206¦A020602
|  | Extra Components | A020603 | More than the usual number of components present/additional components from this or other devices. |  |  | Term was added on 27 January 2022. For details, see comment No. 222 of the Change Log (Release Number 2022). | A02¦A0206¦A020603
| Device Damaged Prior to Receipt by User |  | A0207 | Problem associated with packaging or shipping damage prior to receipt by the user. | MedDRA:10083748:Device damaged prior to use |  | Term and Definition modified for Release Number 2025. Previous term was "Shipping Damage or Problem". See no. 106 of Change Request Log. | A02¦A0207
|  | Delivered as Unsterile Product | A020701 | Problem associated with a device being received in such a manner to indicate that its sterility has been compromised (e.g. sterile packaging breached, visible contaminate present) |  |  |  | A02¦A0207¦A020701
Chemical Problem |  |  | A03 | Problem associated with any deviation from the documented specifications of the device that relate to any chemical characterization, i.e., element, compound, or mixture. | MedDRA:10069872:Device chemical property issue | Modified (technical) | Definition modified for Release Number 2026. See no. 177 of Change Request Log | A03
| Device Emits Odor |  | A0301 | Problem associated with an unexpected or inappropriate smell released by the device. | MedDRA:10083783:Device emits odor |  |  | A03¦A0301
| Device Ingredient or Reagent Problem |  | A0302 | Problem associated with any deviations from the documented specifications of the device that relate to any ingredient or reagent characterization. | MedDRA:10071303:Product ingredient issue |  |  | A03¦A0302
|  | Clumping in Device or Device Ingredient | A030201 | Problem associated with the aggregation of particles into irregular masses. |  |  |  | A03¦A0302¦A030201
|  | Coagulation in Device or Device Ingredient | A030202 | Problem associated with the undesired characterization of congealing, solidifying, thickening, curdling. |  |  |  | A03¦A0302¦A030202
|  | Precipitate in Device or Device Ingredient | A030203 | Problem associated with the separation of solid particles from a liquid as the result of a chemical or physical change. |  |  |  | A03¦A0302¦A030203
|  | Cross Reactivity | A030204 | Problem associated with the degree to which an antibody or antigen participates in cross reactions. |  |  |  | A03¦A0302¦A030204
|  | Particulates | A030205 | Substances that consist of separate particles that are introduced by the device during or before use. |  | Modified (editorial) | Definition modified for Release Number 2026. See no. 45 of Change Log. | A03¦A0302¦A030205
|  | High pH | A030206 | pH higher than expected and/or anticipated. |  |  |  | A03¦A0302¦A030206
|  | Low pH | A030207 | pH lower than expected and/or anticipated. |  |  |  | A03¦A0302¦A030207
|  | Unexpected Color | A030208 | Color of product is different from that expected. |  |  | Term was added on 18 March 2020 as an additional level 3 term under "Device Ingredient or Reagent Problem". | A03¦A0302¦A030208
| Improper Chemical Reaction |  | A0303 | Problem associated with an unexpected or incomplete chemical reaction or effect. | MedDRA:10092652:Device improper chemical reaction |  |  | A03¦A0303
Material Integrity Problem |  |  | A04 | Problem associated with any deviations from the documented specifications of the device that relate to the limited durability of all material used to construct device. | MedDRA:10083732:Device material integrity issue |  |  | A04
| Break |  | A0401 | Problem associated with undesired damage or breakage of those materials used in the device construction. | MedDRA:10012575:Device breakage |  |  | A04¦A0401
|  | Fracture | A040101 | Problem associated with a partial or full-thickness crack in the device materials. |  |  |  | A04¦A0401¦A040101
|  | Loss of or Failure to Bond | A040102 | Problem associated with lack or loss of adherence between materials intended to be joined together by any bonding such as an adhesive, heat welding, etc. |  |  | Definition modified for Release Number 2024. See no. 26 of Change Log | A04¦A0401¦A040102
|  | Material Fragmentation | A040103 | Problem associated with small pieces of the device breaking off unexpectedly. |  |  |  | A04¦A0401¦A040103
|  | Solder Joint Fracture | A040104 | Problem associated with undesired damage or breakage in a solder joint of materials used in the device construction. |  |  |  | A04¦A0401¦A040104
| Burst Container |  | A0402 | Problem associated with the pressure inside a vessel or container (tubing, bottle or lines) rising to such a degree that the container ruptures | MedDRA:10069302:Product container damaged | Modified (technical) | Definition and term modified for Release Number 2026. See no. 9 of Change Log. | A04¦A0402
| Explosion |  | A0403 | Problem associated with the violent bursting due to the sudden expansion of air, gas or fluid. | MedDRA:10083315:Product explosion |  |  | A04¦A0403
| Crack |  | A0404 | Problem associated with an undesired partial separation and/or a visible opening along the length or width in the materials that are used in the device construction. | MedDRA:10074154:Device cracked |  |  | A04¦A0404
| Degraded |  | A0405 | Problem associated with a undesired change in the chemical structure, physical properties, or appearance in the materials that are used in the device construction. | MedDRA:10074155:Device material degradation |  |  | A04¦A0405
|  | Calcified | A040501 | Problem associated with buildup of calcium salts on the device. |  |  |  | A04¦A0405¦A040501
|  | Corroded | A040502 | Problem associated with the chemical or electrochemical reaction between materials, usually a metal and its environment that produces a deterioration of the metal and its properties. |  |  |  | A04¦A0405¦A040502
|  | Material Erosion | A040503 | Problem associated with a progressive loss of a material from a solid surface. |  |  |  | A04¦A0405¦A040503
|  | Pitted | A040504 | Problem associated with the corrosion of a material's surface, confined to a point or small area that takes the form of cavities. |  |  |  | A04¦A0405¦A040504
|  | Flaked | A040505 | Problem associated with the detachment of small pieces of the coating film of a material. |  |  |  | A04¦A0405¦A040505
|  | Peeled/Delaminated | A040506 | Peeling or delamination of composite materials, including coatings, that occurs when layers are separated as a result of stress or impact and resulting in loss of mechanical toughness. |  |  |  | A04¦A0405¦A040506
|  | Naturally Worn | A040507 | Problem associated with material damage to a surface, usually involving progressive loss or displacement of material, due to relative motion between that surface and a contacting substance or substances. |  |  |  | A04¦A0405¦A040507
|  | Unraveled Material | A040508 | Problem due to the undesired unravelling of material (e.g. disentangled, unwound etc.). |  |  |  | A04¦A0405¦A040508
| Material Deformation |  | A0406 | Problem associated with an undesired material change in shape or property caused by external forces. | MedDRA:10083676:Device material deformation |  |  | A04¦A0406
|  | Deformation Due to Compressive Stress | A040601 | Problem associated with an undesired bulge, bend, bow, kink, or wavy condition observed in the device material resulting from compressive stresses. |  |  |  | A04¦A0406¦A040601
|  | Dent in Material | A040602 | Problem associated with a undesired change in shape, characterized by the presence of a slight hollow (dent) in the device surface. |  |  |  | A04¦A0406¦A040602
|  | Failure to Fold | A040603 | Problem associated with an undesired material change in physical property, characterized by failure to fold. |  |  |  | A04¦A0406¦A040603
|  | Failure to Unfold or Unwrap | A040604 | Problem associated with the comprising materials' deformation in that device fails to open its wrapping or open/extend in a certain manner i.e. balloon or lens. |  |  |  | A04¦A0406¦A040604
|  | Material Frayed | A040605 | Problem associated with the comprising materials having damaged edges. |  |  |  | A04¦A0406¦A040605
|  | Material Invagination | A040606 | Problem associated with an undesired material change in shape, characterized by the infolding of one part within another part of a structure, or creates an internal pocket within the device's structure. |  | Modified (technical) | Definition modified for Release Number 2026. See no. 48 of Change Request Log | A04¦A0406¦A040606
|  | Material Too Rigid or Stiff | A040607 | Problem associated with undesired rigidity of device material (resisting deformation in response to an applied force). |  |  | Definition modified for Release Number 2025. See no. 44 of Change Request Log | A04¦A0406¦A040607
|  | Material Too Soft/Flexible | A040608 | Problem associated with any device material that results in the material's inability to maintain the desired shape or support function. |  |  |  | A04¦A0406¦A040608
|  | Material Twisted/Bent | A040609 | Problem associated with deformations that lead to twisting or bending of the device. |  |  |  | A04¦A0406¦A040609
|  | Melted | A040610 | Problem associated with a solid device being transformed into a molten or liquid state. |  |  |  | A04¦A0406¦A040610
|  | Stretched | A040611 | Problem associated with an increase or elongation, or thinning in a materials' dimension. |  | Modified (technical) | Definition modified for Release Number 2026. See no. 129 of Change Request Log | A04¦A0406¦A040611
|  | Thickening of Material | A040612 | Device is thickened due to deformation. Note: Not to be used for thickening of liquid. |  |  | Term was added on 18 March 2020 as an additional level 3 term under "Material Deformation". | A04¦A0406¦A040612
| Material Discolored |  | A0407 | Problem associated with an undesired streak, pattern and/or a noticeable change in original color. | MedDRA:10069876:Device discoloration | Modified (technical) | Definition modified for Release Number 2026. See no. 49 of Change Request Log | A04¦A0407
| Material Disintegration |  | A0408 | Problem associated with material breaking into small particles. | MedDRA:10083766:Device material fragmentation |  |  | A04¦A0408
| Material Opacification |  | A0409 | Problem associated with an undesirable opaqueness or cloudiness. | MedDRA:10069879:Device material opacification |  |  | A04¦A0409
| Material Perforation |  | A0410 | Material constituting device is perforated possibly compromising the device's intended purpose. | MedDRA:10083727:Device material punctured |  |  | A04¦A0410
|  | Material Puncture/Hole | A041001 | Device material(s) punctured leading to undesired holes/openings. |  |  |  | A04¦A0410¦A041001
| Material Protrusion/Extrusion |  | A0411 | Problem associated with undesired physical appearance of device material, specifically when material extends beyond or above device surface. | MedDRA:10083726:Device material protrusion |  |  | A04¦A0411
| Material Rupture |  | A0412 | Problem associated with perforations that lead to bursting of the device. | MedDRA:10083719:Device material rupture |  |  | A04¦A0412
| Material Separation |  | A0413 | Problem associated with an undesired disassociation or breaking apart of the device. | MedDRA:10083678:Device material separation |  |  | A04¦A0413
| Material Split, Cut or Torn |  | A0414 | Problem associated with materials consisting the device are split, cut or torn due to external forces (e.g. wrenching or laceration) or internal forces (e.g. exceeding the tensile stress limits belonging to the materials used in the device construction). | MedDRA:10083740:Device material split |  |  | A04¦A0414
| Scratched Material |  | A0415 | Problem associated with an undesirable shallow cut or narrow groove in the surface of the device materials. | MedDRA:10083743:Device material scratched |  |  | A04¦A0415
| Failure to Degrade |  | A0416 | Failure of a material to degrade, absorb, dissolve, etc. in an appropriate manner, when the material is expected to do so. | MedDRA:10093154:Device material degradation failure |  | Term added for Release Number 2025. See no. 87 of Change Request Log | A04¦A0416
Mechanical Problem |  |  | A05 | Problems associated with mechanical actions or defects, including moving parts or subassemblies, etc. | MedDRA:10083638:Device mechanical issue |  |  | A05
| Detachment of Device or Device Component |  | A0501 | Problem associated with the separation of the device from its physical construct, integrity, or chassis. | MedDRA:10083644:Detachment of device |  |  | A05¦A0501
| Device Damaged by Another Device |  | A0502 | Problem associated with one device causing harm to another device. | MedDRA:10069870:Device damage |  |  | A05¦A0502
| Ejection Problem |  | A0503 | Problems associated with the inability or unexpected ejection or discharge of the device from its physical location. | MedDRA:10012578:Device expulsion |  | Editorial change to the definition on 18 March 2020. | A05¦A0503
|  | Failure to Eject | A050301 | Problem associated with the inability of the device to be ejected or discharged from its physical location. |  |  | Editorial change to the definition on 18 March 2020. | A05¦A0503¦A050301
|  | Unintended Ejection | A050302 | Problem associated with unexpected ejection of the device from its physical location. This includes but is not limited to devices such as clip appliers, film cartridge, and staples. |  |  | Editorial change to the definition on 18 March 2020. | A05¦A0503¦A050302
| Leak/Splash |  | A0504 | Problem associated with the escape or ingress of a liquid (including blood and bodily fluids), gel, gas or radiation from or into a vessel or container. | MedDRA:10012587:Device leakage | Modified (technical) | Definition was modified on 27 January 2022 (Old definition: Leak/Splash: Problem associated with the escape of a liquid, gas or radiation from the vessel or container in which it is housed). For details, see comment No. 14 of the Change Log (Release Number 2022). Definition modified for Release Number 2025. See no. 2 of Change Request Log. Definition modified for Release Number 2026. See no. 50 of Change Request Log. | A05¦A0504
|  | Fluid/Blood Leak | A050401 | Escape (Release, Discharge) of fluid (including blood or bodily fluid) through an unintended location; often accompanied by a loss of pressure and/or output. This includes both external and internal fluid leak. This does not include fluidized powder leak. |  |  | Term/Definition was modified on 27 January 2022 (Old term/definition: Fluid Leak: Escape (Release, Discharge) of fluid through an unintended location often accompanied by a loss of pressure and/or output). For details, see comment No. 15, 16, 17, 18 of the Change Log (Release Number 2022). | A05¦A0504¦A050401
|  | Gas/Air Leak | A050402 | Problem associated with the unintended escape of a gas/air from the container in which it is housed. This also includes both external and internal air/gas leak. |  |  | Term/Definition was modified on 27 January 2022 (Old term/definition: Gas Leak: Problem associated with the unintended escape of a gas from the container in which it is housed). For details, see comment No. 19, 20 of the Change Log (Release Number 2022). | A05¦A0504¦A050402
|  | Gel Leak | A050403 | Escape (Release, Discharge) of gel through an unintended location - as in leakage of ultrasound gel. Escape or release of gel from containment structures - as in gel filled implant leak. |  |  |  | A05¦A0504¦A050403
|  | Radiation Leak | A050404 | Escape of radiation (energy in the form of waves or subatomic particles, especially those that cause ionization) through containment structures, leading to unintended exposure. |  |  |  | A05¦A0504¦A050404
|  | Perivalvular Leak | A050405 | Problem associated with the escape of blood around a heart valve, particularly around its leaflets. Also known as paravalvular leak. |  |  | Definition was modified on 31 January 2023 (Old definition: Problem associated with the escape of blood around a heart valve, particularly around its leaflets.). For details, see comment No. 7 of the Change Log (Release Number 2023). | A05¦A0504¦A050405
|  | Intravalvular Regurgitation | A050406 | Problem associated with the escape or backward flow of blood through or across the valve. Also known as transvalvular regurgitation |  | Modified (editorial) | Term was added on 27 January 2022. For details, see comment No. 147, 175 of the Change Log (Release Number 2022). Term and Definition modified for Release Number 2024. Previous term was Central Regurgitation See no. 47 of Change Log. Editorial change for Release Number 2026. See no. 215 of Change Log. | A05¦A0504¦A050406
|  | Powder Leak | A050407 | Escape (release, discharge) of powder (including particulates) through an unintended location, often accompanied by a loss of pressure and/or output. |  |  | Term was added on 27 January 2022. For details, see comment No. 21 of the Change Log (Release Number 2022). | A05¦A0504¦A050407
|  | Device Endoleak | A050408 | The observed device problem is persistent blood flow/pressure in the aneurysm sac following an endovascular aneurysm repair (EVAR) procedure. Where other device problems lead to an endoleak the observed problems should be coded and E2121 selected as the E term. |  |  | Term was added on 31 January 2023. For details, see comment No. 33 of the Change Log (Release Number 2023). Term and Definition modified for Release Number 2025. Previous term was "Endoleaks". See no. 116 of Change Request Log. | A05¦A0504¦A050408
|  | Unspecified Valve Regurgitation | A050409 | Problem associated with the escape or backward flow of blood relative to a valve where it is not specified or not confirmed whether the leak is transvalvular/intravalvular or paravalvular/perivalvular. |  |  | Term added for Release Number 2024. See no. 46 of Change Log | A05¦A0504¦A050409
| Firing Problem |  | A0505 | Problem associated with the device not discharging as intended. | MedDRA:10070961:Mechanical device firing issue |  |  | A05¦A0505
|  | Failure to Fire | A050501 | Problem associated with failure of the device to discharge its load (e.g. surgical stapler failed to partially or completely deploy its staples). |  |  |  | A05¦A0505¦A050501
|  | Misfire | A050502 | Problem associated with a therapy or algorithm not being delivered or executed at the expected time. |  |  |  | A05¦A0505¦A050502
| Mechanical Jam |  | A0506 | The motion of the device is prevented or restricted. | MedDRA:10083776:Device mechanical jam |  |  | A05¦A0506
| Mechanics Altered |  | A0507 | Problem associated with a device mechanical functioning of machinery, moving parts or tools of device being changed or modified. | MedDRA:10083638:Device mechanical issue |  |  | A05¦A0507
|  | Failure to Align | A050701 | Problem associated with a circuit, equipment, or system whereby its functions fail to be properly synchronized or its relative positions properly oriented. |  |  |  | A05¦A0507¦A050701
|  | Failure to Cut | A050702 | Inability of the device to make an incision, pierce or open as intended. |  |  |  | A05¦A0507¦A050702
|  | Failure to Cycle | A050703 | Problem associated with the device failing to complete a series of processes or events. |  |  |  | A05¦A0507¦A050703
|  | Failure to Form Staple | A050704 | Problem associated with the device failing to connect tissue with a stapling device due to the staples not forming correctly. |  |  |  | A05¦A0507¦A050704
|  | Failure to Seal | A050705 | Problem associated with the device failing to seal the intended vessel with the device intended to provide a seal (e.g. electrosurgical hemostasis). |  |  | Term was added on 27 January 2022. For details, see comment No. 198 of the Change Log (Release Number 2022). Editorial change to reinstate missing definition for release number 2024.1 | A05¦A0507¦A050705
| Noise, Audible |  | A0508 | Problem associated with any unintended sound which emanates from the device (for example, squeaking from two parts rubbing together or buzzing sounds from electrical components). | MedDRA:10083747:Noise from device |  |  | A05¦A0508
| Physical Resistance/Sticking |  | A0509 | Problem associated with the lack of movement in the device due parts sticking or seizing. | MedDRA:10083674:Resistance to movement in device |  |  | A05¦A0509
|  | Adhesive Too Strong | A050901 | Strength of an adhesive bond is too great; typically used when the adhesive is meant as a temporary hold but is acting in a more permanent manner. |  |  | Term was added on 27 January 2022. For details, see comment No. 215 of the Change Log (Release Number 2022). | A05¦A0509¦A050901
| Retraction Problem |  | A0510 | Problem associated with drawing back the device or part of the device to an intended location. | MedDRA:10083718:Device retraction issue | Modified (technical) | Technical change to definition for Release Number 2026. See no. 34 of Change Log. | A05¦A0510
| Structural Problem |  | A0511 | Problem associated with the basic physical construction or physical make up of the device. | MedDRA:10069880:Device physical property issue |  |  | A05¦A0511
|  | Collapse | A051101 | Problem associated with the buckling or crushing of material from external forces. |  |  |  | A05¦A0511¦A051101
|  | Sharp Edges | A051102 | The device has undesirable sharp edges which can cause harm or damage. |  |  |  | A05¦A0511¦A051102
|  | Difficult to Fold, Unfold or Collapse | A051103 | Problem associated with the use of the device in terms of user experiencing difficulty to close or to spread out/extend length of the device, even if the operation is being performed according to labeled instructions for use. |  |  | Technical change to the term on 18 March 2020 to include difficulty in collapsing. | A05¦A0511¦A051103
|  | Difficult to Open or Close | A051104 | Problem associated with the use of the device in terms of user experiencing difficulty opening and closing the device, even if the operation is being performed according to labeled instructions for use. |  |  |  | A05¦A0511¦A051104
|  | Incomplete Coaptation | A051105 | Problem associated with the heart valve leaflet not closing properly. |  |  |  | A05¦A0511¦A051105
| Unintended Movement |  | A0512 | Problem associated with an undesired movement of the device, which may be related to the device malfunction, misdiagnosis, or mistreatment. | MedDRA:10028038:Movement of device from original site |  |  | A05¦A0512
|  | Device Dislodged or Dislocated | A051201 | Problems associated with the device not remaining in an expected location. |  |  |  | A05¦A0512¦A051201
|  | Device Tipped Over | A051202 | Problem associated with the inability of the device to stay in an upright position. |  |  |  | A05¦A0512¦A051202
|  | Device Fell | A051203 | Problem associated with the device or a component unexpectedly being dropped or moving down from an intended place. |  |  |  | A05¦A0512¦A051203
|  | Device Slipped | A051204 | Problem associated with the device moving or sliding from the intended position. |  |  |  | A05¦A0512¦A051204
|  | Unintended Collision | A051205 | Problem associated with the device impacting with another object. |  |  |  | A05¦A0512¦A051205
|  | Unintended System Motion | A051206 | Problem associated with any motion of the system or components that was not initiated by the user. |  |  |  | A05¦A0512¦A051206
|  | Unstable | A051207 | Problem associated with the mechanical stability of the device. |  |  |  | A05¦A0512¦A051207
|  | Vibration | A051208 | Problem associated with the undesirable mechanical oscillation. |  |  |  | A05¦A0512¦A051208
| Inadequate Lubrication |  | A0513 | The device shows inadequate lubrication. | MedDRA:10093149:Device mechanical parts lubrication inadequate |  | Term was added on 18 March 2020 as an additional level 2 term under "Mechanical Problem." | A05¦A0513
| Damaged Thread |  | A0514 | Problems associated with threads that are damaged either prior to use due to bad threading or repeated re-use (possibly with other damaged components) or threads that became stripped or damaged while being used. | MedDRA:10093155:Device mechanical thread damage |  | Term was added on 31 January 2023. For details, see comment No. 5 of the Change Log (Release Number 2023). | A05¦A0514
Optical Problem |  |  | A06 | Problem associated with transmission of visible light affecting the quality of the image transmitted or otherwise affecting the intended application of the visible light path. | MedDRA:10069863:Device optical issue |  |  | A06
| Misfocusing |  | A0601 | The problem relates to the poor focusing of the object or the focus is on the wrong object or in the wrong area. | MedDRA:10093144:Device optical misfocusing |  |  | A06¦A0601
| Optical Decentration |  | A0602 | Problem associated with being off-center of optical lenses. | MedDRA:10093150:Device optical decentration |  |  | A06¦A0602
| Optical Discoloration |  | A0603 | Problem associated with an undesired change of color. | MedDRA:10093229:Device optical discoloration |  |  | A06¦A0603
| Optical Distortion |  | A0604 | Problem associated with an optical defect in an image-forming system whereby the image is not the shape of an ideal image of the object. | MedDRA:10093156:Device optical distortion |  |  | A06¦A0604
| Optical Obstruction |  | A0605 | Problem associated with the blocking of optical devices, e.g. visual pathways. | MedDRA:10093145:Device light transmission obstruction |  |  | A06¦A0605
Electrical /Electronic Property Problem |  |  | A07 | Problem associated with the function of the electrical circuitry of the device. | MedDRA:10083688:Device electrical/electronic failure |  | Definition was modified on 31 January 2023 to replace "a failure" with "the function". For details, see comment No. 8 of the Change Log (Release Number 2023). | A07
| Capturing Problem |  | A0701 | Problem associated with the inability of the device to achieve successful depolarization and contraction of a cardiac chamber caused by a pacemaker output pulse. | MedDRA:10069867:Device capturing issue |  |  | A07¦A0701
|  | Failure to Capture | A070101 | Problem associated with the failure to achieve effective and consistent depolarization of the heart resulting from the electrical stimulus of the pacemaker. |  |  |  | A07¦A0701¦A070101
|  | High Capture Threshold | A070102 | Problem with the amount of output energy needed to cause cardiac depolarization being higher than expected/desired. |  |  |  | A07¦A0701¦A070102
|  | Intermittent Capture | A070103 | Problem associated with the ineffective and inconsistent depolarization of the heart. |  |  |  | A07¦A0701¦A070103
|  | Unstable Capture Threshold | A070104 | Problem with the amount of output energy needed to cause cardiac depolarization being unstable. |  |  |  | A07¦A0701¦A070104
| Continuous Firing |  | A0702 | Problem associated with the excessive production of electrical impulses over a period. | MedDRA:10070962:Electrical device firing issue |  |  | A07¦A0702
| Arcing |  | A0703 | Problem associated with electrical current flowing through a gap between two conductive surfaces, typically resulting in a visible flash of light. | MedDRA:10065963:Device electrical finding |  |  | A07¦A0703
|  | Arcing at Paddles | A070301 | Problem associated with electrical current flowing through a gap between paddles (conductive surfaces), typically resulting in a visible flash of light. |  |  |  | A07¦A0703¦A070301
|  | Arcing of Electrodes | A070302 | Problem associated with electrical current flowing through a gap between electrodes (conductive surfaces), typically resulting in a visible flash of light. |  |  |  | A07¦A0703¦A070302
| Sparking |  | A0704 | Problem associated with a flash of light related to an electrical discharge into a normally non conductive medium, such as air. Not associated with a discharge between two conductive surfaces. | MedDRA:10065963:Device electrical finding |  |  | A07¦A0704
| Battery Problem |  | A0705 | Problem associated with the internal power of the device (e.g. battery, transformer, fuel cell or other power sources). | MedDRA:10071172:Device battery issue |  |  | A07¦A0705
|  | Battery Problem: High Impedance | A070501 | Problem related to increased battery internal impedance. |  |  |  | A07¦A0705¦A070501
|  | Battery Problem: Low Impedance | A070502 | Problem related to decreased battery internal impedance. |  |  |  | A07¦A0705¦A070502
|  | Failure to Run on Battery | A070503 | Problem associated with the device failing to operate when not connected to a fixed power source. |  |  |  | A07¦A0705¦A070503
|  | Premature Discharge of Battery | A070504 | Battery discharging earlier than expected. |  |  |  | A07¦A0705¦A070504
| Charging Problem |  | A0706 | Problem associated with the inability of the device to successfully charge an electrical source. | MedDRA:10083923:Device charging issue |  |  | A07¦A0706
|  | Aborted Charge | A070601 | Problem associated with the premature ending of the charging process (e.g. of a battery or other charge storage device). |  |  |  | A07¦A0706¦A070601
|  | Delayed Charge Time | A070602 | Problem associated with an unexpected amount of time required to charge the device (e.g. a delay in starting charging or a longer than expected charge time). |  |  |  | A07¦A0706¦A070602
|  | Failure to Charge | A070603 | Problem associated with inability to initiate the appropriate charging process (e.g. of a battery or other charge storage device) |  |  |  | A07¦A0706¦A070603
| Failure to Discharge |  | A0707 | Problem associated with the failure of a battery or other charge storage device to appropriately discharge as intended. Does not apply to defibrillation. | MedDRA:10093158:Discharge failure of device power source |  |  | A07¦A0707
| Power Problem |  | A0708 | Problem associated with the energy to operate the device. | MedDRA:10069865:Device power source issue |  |  | A07¦A0708
|  | Complete Loss of Power | A070801 | Problem associated with the lack of power to run the device. |  |  |  | A07¦A0708¦A070801
|  | Intermittent Loss of Power | A070802 | Problem associated with an intermittent disruption to the power to run the device. |  |  |  | A07¦A0708¦A070802
|  | Failure to Power Up | A070803 | Problem associated with the inability of the device to turn on related to energy delivered to the device. |  |  |  | A07¦A0708¦A070803
|  | Unintended Power Up | A070804 | Problem associated with the device turning on when not intended. |  |  |  | A07¦A0708¦A070804
| Device Sensing Problem |  | A0709 | Problem associated with the device feature that are designed to respond to a physical stimulus (temperature, illumination, motion, cardiac rhythms) and that do not transmit a resulting signal for interpretation or measurement. | MedDRA:10075228:Device sensing failure |  |  | A07¦A0709
|  | Decreased Sensitivity | A070901 | Problem with the device being less sensitive to an input than intended or expected. |  |  |  | A07¦A0709¦A070901
|  | Increased Sensitivity | A070902 | Problem with the device being more sensitive to an input than intended or expected. |  |  |  | A07¦A0709¦A070902
|  | Failure to Analyze Signal | A070903 | Problem with the device not analyzing a signal. |  |  |  | A07¦A0709¦A070903
|  | Failure to Select Signal | A070904 | Problem associated with the failure of the device to select the appropriate input signal. |  |  |  | A07¦A0709¦A070904
|  | High Sensing Threshold | A070905 | Problem associated with the amount of input required by the device to detect a signal being higher than expected/desired. |  |  |  | A07¦A0709¦A070905
|  | Low Sensing Threshold | A070906 | Problem associated with the amount of an input required by the device to detect a signal being lower than expected/desired. |  |  |  | A07¦A0709¦A070906
|  | Loss of Threshold | A070907 | Problem associated with the loss of the minimum amount of energy, voltage, or current needed to consistently stimulate the heart muscle. |  |  |  | A07¦A0709¦A070907
|  | Failure to Sense | A070908 | Problem associated with the failure of the device designed to respond to a physical stimulus (as temperature, illumination, motion) to transmit a resulting signal for interpretation or measurement. |  |  |  | A07¦A0709¦A070908
|  | Oversensing | A070909 | Problem related to failure of the device to properly filter cardiac signals resulting in inappropriate device response. |  | Modified (editorial) | Term modified for Release Number 2026. See no. 317 of Change Log. | A07¦A0709¦A070909
|  | Undersensing | A070910 | Problem related to failure of the device to properly detect intrinsic cardiac activity and respond appropriately. |  | Modified (editorial) | Term modified for Release Number 2026. See no. 318 of Change Log. | A07¦A0709¦A070910
|  | Sensing Intermittently | A070911 | Problem with the device receiving an incoming signal on an intermittent basis when expected to be continuous. |  |  |  | A07¦A0709¦A070911
|  | Incorrect Interpretation of Signal | A070912 | Problem with the device inappropriately analyzing a signal. |  |  |  | A07¦A0709¦A070912
| Failure to Conduct |  | A0710 | Problem associated with the inability of the device to allow a current of electricity to pass or to conduct electricity continuously along an electrical path. | MedDRA:10069851:Device electrical conduction issue |  |  | A07¦A0710
| Interrogation Problem |  | A0711 | Problems associated with the device's ability to respond to signals from a system designed to interrogate its status. | MedDRA:10093226:Device electronic interrogation issue |  |  | A07¦A0711
|  | Difficult to Interrogate | A071101 | Problem associated with difficulty of a transponder system to trigger a response. |  |  |  | A07¦A0711¦A071101
|  | Failure to Interrogate | A071102 | Problem associated with the device failure to appropriately respond to signals from a system designed to interrogate its status. |  |  |  | A07¦A0711¦A071102
| Pacing Problem |  | A0712 | Problem associated with the inability of the device to generate a therapeutic simulated heart beat via electrical impulses. | MedDRA:10069864:Device pacing issue |  |  | A07¦A0712
|  | Failure to Convert Rhythm | A071201 | Failure of the device therapy or set of therapies to terminate the harmful cardiac rhythm that the therapy is meant to terminate. |  |  |  | A07¦A0712¦A071201
|  | Inaccurate Synchronization | A071202 | Problem associated with an error due to imperfect timing of two operations, e.g. signal transmission time. |  |  |  | A07¦A0712¦A071202
|  | Inappropriate Waveform | A071203 | Failure of the device to generate a correctly-shaped pacing output, e.g., a waveform that is too wide. |  |  |  | A07¦A0712¦A071203
|  | No Pacing | A071204 | Problem associated with the device ceasing to deliver paces. |  |  |  | A07¦A0712¦A071204
|  | Pacemaker Found in Back-Up Mode | A071205 | A device with a pacing function found in back-up Mode. This may be an appropriate fail-safe action (e.g. end of battery life), or be caused by device malfunction or due to operator error. |  | Retired (2023) | Term was retired on 31 January 2023, with the addition of term A0723 Device in Backup-mode. For details, see comment No. 8 of the Change Log (Release Number 2023). | A07¦A0712¦A071205
|  | Pacing Asynchronously | A071206 | Problem associated with a pacing transmission process such that between any two significant instants in the same group, there is always an integral number of unit intervals. Between two significant instants located in different groups, there are not always an integral number of unit intervals. |  |  |  | A07¦A0712¦A071206
|  | Pacing Inadequately | A071207 | Pacing voltage or pulse width is less than desired. |  |  |  | A07¦A0712¦A071207
|  | Pacing Intermittently | A071208 | Problem associated with the failure of pacing device for a limited period of time, following which the item recovers its ability to perform its required function without being subjected to any external corrective action. Note: such as failure is often recurrent. |  |  |  | A07¦A0712¦A071208
|  | Extracardiac Stimulation | A071209 | Problem associated with undesired pacemaker induced stimulation of extracardiac (muscle) tissue (e.g. diaphragm, pectoral or intercostal muscle) |  | Modified (editorial) | Term and Definition modified for Release Number 2025. Previous term was “Pocket Stimulation”. See no. 5 of Change Request Log. Term modified for Release Number 2026. See no. 216 of Change Log. | A07¦A0712¦A071209
| Defibrillation/Stimulation Problem |  | A0713 | Problem associated with the delivery of electrical shock/stimulation by a device intended to deliver shocks/stimulation. | MedDRA:10075322:Device inappropriate shock delivery |  | Technical change to the term and definition on 18 March 2020 to include stimulation and to differentiate this term with "Unintended Electrical Shock" (A0714). | A07¦A0713
|  | Failure to Deliver Shock/Stimulation | A071301 | Problem associated with the failure of the device to deliver electrical energy. |  |  | Technical change to the term on 18 March 2020 to include stimulation. | A07¦A0713¦A071301
|  | Inappropriate/Inadequate Shock/Stimulation | A071302 | A shock was delivered with incorrect energy level or at incorrect timing. |  |  | Technical change to the term and definition on 18 March 2020 to include inadequate stimulation. | A07¦A0713¦A071302
|  | Intermittent Shock/Stimulation | A071303 | Problem associated with the failure to deliver electrical energy for a limited period of time, following which the item recovers its ability to perform its required function without being subjected to any external corrective action. Note: such a failure is often recurrent. |  |  | Technical change to the term and definition on 18 March 2020 to include stimulation. | A07¦A0713¦A071303
| Unintended Electrical Shock |  | A0714 | The device delivers unintended electrical shock. Use this term for devices that are not intended to deliver shocks/stimulation. | MedDRA:10014357:Electric shock |  | Technical change to the definition on 18 March 2020 to differentiate this term with "Defibrillation/Stimulation Problem" (A0713) and lower level terms. | A07¦A0714
| Grounding Malfunction |  | A0715 | Problem associated with the inability to connect conductors of an electronic system for the purpose of controlling or impeding ground currents and voltages. | MedDRA:10065963:Device electrical finding |  |  | A07¦A0715
| Electrical Overstress |  | A0716 | Problem associated with an electrical activity that exceeded the specified threshold limit of the internal integrated circuitry. | MedDRA:10065963:Device electrical finding |  |  | A07¦A0716
| Electrostatic Discharge |  | A0717 | Problem associated with the discharge of electricity between two bodies previously electrically charged. | MedDRA:10065963:Device electrical finding |  | Term modified for Release Number 2025. Previous term was "Electro-Static Discharge". See no. 139 of Change Request Log. | A07¦A0717
| Failure to Shut Off |  | A0718 | Problem associated with the device not powering off when a shut down was requested. | MedDRA:10065963:Device electrical finding |  |  | A07¦A0718
|  | Device Remains Activated | A071801 | Problem associated with the device continuing to be in an active state after deactivation was requested. |  |  |  | A07¦A0718¦A071801
| Unexpected Shutdown |  | A0719 | Problem associated with the device unexpectedly powering down. | MedDRA:10065963:Device electrical finding |  |  | A07¦A0719
| Electromagnetic Compatibility Problem |  | A0720 | Problem associated with the ability of a system to function in its electromagnetic environment without introducing intolerable disturbances to anything in its environment. | MedDRA:10093141:Device electromagnetic compatibility issue |  |  | A07¦A0720
|  | Electromagnetic Interference | A072001 | Problem associated with a measure of electromagnetic radiation from equipment. |  |  |  | A07¦A0720¦A072001
|  | Radiofrequency Interference (RFI) | A072002 | Problem associated with the degradation of the reception of a wanted signal caused by RF disturbance. |  |  |  | A07¦A0720¦A072002
| Circuit Failure |  | A0721 | Problem associated with a failure of the internal network paths or electrical circuitry (i.e. electrical components, circuit boards, wiring) | MedDRA:10083364:Device short circuiting |  |  | A07¦A0721
|  | Capacitative Coupling | A072101 | Problem associated with the transfer of energy within an electrical network by means of the capacitance between circuit nodes. It occurs when energy is coupled from one circuit to another through an electric field. |  |  |  | A07¦A0721¦A072101
|  | Electrical Shorting | A072102 | Problem associated with an electric current traveling along an accidental path (unintended path) in a circuit. |  | Modified (editorial) | Definition modified for Release Number 2026. See no. 179 of Change Request Log. | A07¦A0721¦A072102
|  | Intermittent Continuity | A072103 | Problem associated with intermittent faults in electrical/electronic interconnections. |  |  |  | A07¦A0721¦A072103
| Impedance Problem |  | A0722 | Problem associated with electrical impedance levels between device and patient connections. | MedDRA:10069844:Device electrical impedance issue |  |  | A07¦A0722
|  | High Impedance | A072201 | Problem associated with higher than intended electrical impedance levels between device and patient connections. |  | Modified (editorial) | Term modified for Release Number 2026. See no. 217 of Change Request Log. | A07¦A0722¦A072201
|  | Low Impedance | A072202 | Problem associated with lower than intended electrical impedance levels between device and patient connections. |  | Modified (editorial) | Term modified for Release Number 2026. See no. 218 of Change Request Log. | A07¦A0722¦A072202
|  | Unstable Impedance | A072203 | The impedance varies without being consistently too high or too low. |  |  | Term added for Release Number 2025. See no. 7 of Change Request Log | A07¦A0722¦A072203
| Unexpected Mode Switch |  | A0723 | A device switches mode unexpectedly e.g. back-up mode or electrocautery mode. | MedDRA:10065963:Device electrical finding |  | Term was added on 31 January 2023. For details, see comment No. 8 of the Change Log (Release Number 2023). Term and Definition modified for Release Number 2025. Previous term was “Device in Backup-mode”. See no. 4 of Change Request Log. | A07¦A0723
Calibration Problem |  |  | A08 | Problem associated with the operation of the device, related to its accuracy, and associated with the calibration of the device. | MedDRA:10086464:Device calibration failure |  |  | A08
| Failure to Calibrate |  | A0801 | Problem associated with the failure of the device to perform a self-calibration procedure or process designed to assure the accuracy and proper performance of the device. | MedDRA:10086464:Device calibration failure |  |  | A08¦A0801
| Failure to Recalibrate |  | A0802 | Problem associated with the failure of the device which is unable to regain a standard level of accuracy when performing a calibration procedure or process designed to assure the accuracy and proper performance of the device. | MedDRA:10086464:Device calibration failure |  |  | A08¦A0802
| Imprecision |  | A0803 | Problem associated with the device providing imprecise measurements when compared to a reference standard. | MedDRA:10083645:Incorrect measurement by device |  |  | A08¦A0803
| Overcorrection |  | A0804 | Problem associated with an adjustment that surpasses a set of criteria. | MedDRA:10086464:Device calibration failure |  |  | A08¦A0804
| Undercorrection |  | A0805 | Problem associated with an adjustment that falls below a set of criteria. | MedDRA:10086464:Device calibration failure |  | Term was added on 18 March 2020 as an additional level 2 term under "Calibration Problem". | A08¦A0805
Output Problem |  |  | A09 | Problem associated with any deviation from the documented specifications of the device that relate to the end result, data, or test results provided by the device. | MedDRA:10092653:Device output issue |  |  | A09
| Audible Prompt/Feedback Problem |  | A0901 | Problem with any deviation from the documented specifications of the device that relate to audible feedback. e.g. voice prompts or beeps, but not safety related alarms which are covered under "Protective Measures Problem". | MedDRA:10093157:Device audible prompt issue |  | Editorial change to the term on 18 March 2020. | A09¦A0901
|  | Inappropriate Audible Prompt/Feedback | A090101 | Problem with audible messages which do not guide a device user to the correct action. |  |  |  | A09¦A0901¦A090101
|  | Inaudible or Unclear Audible Prompt/Feedback | A090102 | Problem associated with audible prompts which cannot be heard clearly. |  |  |  | A09¦A0901¦A090102
|  | No Audible Prompt/Feedback | A090103 | Problem associated with the device ceasing to provide audible prompts. |  |  |  | A09¦A0901¦A090103
| Display or Visual Feedback Problem |  | A0902 | Problem with any deviation from the documented specifications of the device that relate to visual feedback. e.g. the display of information, images on a screen, or output from the device. | MedDRA:10069842:Device information output issue |  |  | A09¦A0902
|  | Device Displays Incorrect Message | A090201 | Problem associated with providing incorrect display information. |  |  |  | A09¦A0902¦A090201
|  | Display Difficult to Read | A090202 | Problem associated with legibility of the display, compromising for instance the reading/interpretation of patient parameters or test results. Legibility problems can be due to color, size of font, display screen contrast or other factors. |  |  |  | A09¦A0902¦A090202
|  | Erratic or Intermittent Display | A090203 | A device does not consistently display the same message, result, reading, or image. e.g. the display might flicker, switch between readings or messages, or go completely blank for brief periods of time. |  |  |  | A09¦A0902¦A090203
|  | Image Display Error/Artifact | A090204 | Problem with image display leading to corrupted images or readouts/measurement indications. |  |  |  | A09¦A0902¦A090204
|  | Image Orientation Incorrect | A090205 | Problem associated with an incorrect image orientation on the device display. |  |  |  | A09¦A0902¦A090205
|  | No Display/Image | A090206 | Problem associated with the absence of display or image. |  |  |  | A09¦A0902¦A090206
|  | No Visual Prompts/Feedback | A090207 | Problem associated with the device ceasing to provide visual feedback. |  |  |  | A09¦A0902¦A090207
|  | Poor Quality Image | A090208 | Inadequate quality of an image or any visual representation displayed by the device, or output from the device. |  |  |  | A09¦A0902¦A090208
|  | Visual Prompts Will not Clear | A090209 | Problem with visual messages which continue to be displayed on/by the device after the appropriate action has been taken. |  | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 219 of Change log. | A09¦A0902¦A090209
| Tactile Prompts/Feedback |  | A0903 | Problem with any deviation from the documented specifications of the device that relate to tactile feedback. e.g device vibrational prompt. | MedDRA:10093151:Device tactile prompt issue |  |  | A09¦A0903
|  | Inappropriate Tactile Prompt/Feedback | A090301 | Problem with tactile feedback which does not guide a device user to the correct action. |  |  |  | A09¦A0903¦A090301
|  | No Tactile Prompts/Feedback | A090302 | Problem associated with the device ceasing to provide tactile feedback. |  |  |  | A09¦A0903¦A090302
| Energy Output Problem |  | A0904 | Problem with the device's intended output of energy. | MedDRA:10093163:Device energy output issue |  |  | A09¦A0904
|  | Energy Spectrum Incorrect | A090401 | Problem associated with the energy output from the device not being in the expected part of the spectrum. |  |  |  | A09¦A0904¦A090401
|  | Failure to Deliver Energy | A090402 | Problem associated with the failure of the device to deliver any energy. |  |  |  | A09¦A0904¦A090402
|  | Intermittent Energy Output | A090403 | Problem associated with the energy output from the device being inconsistent over time. |  |  |  | A09¦A0904¦A090403
|  | Output Above Specifications | A090404 | Device output is exceeding the documented specifications of the device. |  | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 220 of Change log. | A09¦A0904¦A090404
|  | Output Below Specifications | A090405 | Device output is below the documented specifications of the device. |  | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 221 of Change log. | A09¦A0904¦A090405
|  | Therapeutic or Diagnostic Output Failure | A090406 | Problem associated with the failure of the device to deliver the output required for treatment or identification of a disease. |  |  |  | A09¦A0904¦A090406
|  | Therapy Delivered to Incorrect Body Area | A090407 | Problem associated with the device causing unintended therapeutic action to an area of the body other than the intended area. |  |  |  | A09¦A0904¦A090407
| Radiation Output Problem |  | A0905 | Problem with the device's intended output of radiation. | MedDRA:10093161:Device radiation output issue |  |  | A09¦A0905
|  | Radiation Output Failure | A090501 | Problem associated with the absence of radiation output from radiological or diagnostic devices. |  |  |  | A09¦A0905¦A090501
|  | Radiation Overexposure | A090502 | Problem associated with excessive radiation emitted from radiological or diagnostic devices. |  |  |  | A09¦A0905¦A090502
|  | Radiation Underexposure | A090503 | Problem associated with too little radiation emitted from radiological or diagnostic devices. |  |  |  | A09¦A0905¦A090503
|  | Unexpected/Unintended Radiation Output | A090504 | Device-emitted radiation when it was not supposed to. This applies to devices which are intended to emit radiation, and the radiation being emitted from the correct part of the device, but at an incorrect time. Use "radiation leak" if the device emits radiation which should never have been emitted, or from a location from which it should never be emitted. |  |  |  | A09¦A0905¦A090504
| Gas Output Problem |  | A0906 | Problem associated with gas output. | MedDRA:10093162:Device gas output issue |  |  | A09¦A0906
| No Device Output |  | A0907 | Problem associated with no measurement outcome, value or data obtained from the device. | MedDRA:10083781:Device readings not obtained |  |  | A09¦A0907
| Incorrect, Inadequate or Imprecise Result or Readings |  | A0908 | Problem associated with a nonconforming end result, data, or test results provided by the device to its performance specifications. | MedDRA:10083739:Device incorrect result issue |  | Editorial change to the term on 16 Feburary 2021. | A09¦A0908
|  | Signal Artifact/Noise | A090801 | Problem associated with impurities or interference in a signal (e.g. ECG artifact). |  |  | Technical change to the term on 18 March 2020 to include signal noise. | A09¦A0908¦A090801
|  | Failure to Obtain Sample | A090802 | The device does not collect or transfer the sample. |  |  |  | A09¦A0908¦A090802
|  | False Negative Result | A090803 | Problem associated with the device incorrectly reporting that something has not been detected and may mislead the operator/user/HCP into not taking certain actions when action should be taken. |  |  | Definition modified for Release Number 2024. See no. 122 of Change Log | A09¦A0908¦A090803
|  | False Positive Result | A090804 | Problem associated with the device incorrectly reporting that something has been detected and may mislead the operator/user/HCP to take certain actions. |  | Modified (editorial) | Definition modified for Release Number 2024. See no. 123 of Change Log. Definition modified to correct typo for Release Number 2026. See no. 176 of Change Log. | A09¦A0908¦A090804
|  | Incorrect Measurement | A090805 | Measurement obtained from or provided by the device is obviously incorrect. |  |  |  | A09¦A0908¦A090805
|  | Non Reproducible Results | A090806 | Device results cannot be reliably reproduced. |  |  |  | A09¦A0908¦A090806
|  | High Readings | A090807 | Reading provided by the device is too high or higher than expected. |  |  |  | A09¦A0908¦A090807
|  | Low Readings | A090808 | Reading provided by the device is too low or lower than expected. |  |  |  | A09¦A0908¦A090808
|  | High Test Results | A090809 | Test results provided by the device are too high or higher than expected. |  |  |  | A09¦A0908¦A090809
|  | Low Test Results | A090810 | Test results provided by the device are too low or lower than expected. |  |  |  | A09¦A0908¦A090810
|  | Unable to Obtain Readings | A090811 | The device does not provide or display a valid reading. |  |  |  | A09¦A0908¦A090811
|  | Missing Test Results | A090812 | Problem associated with the results of a test or measurement not appearing. |  |  |  | A09¦A0908¦A090812
|  | Imprecise or Erratic Results | A090813 | Reports of erroneous/discrepant results which combine high/low and/or positive/negative results. This term is not to be selected where reports indicate consistently high or low or false positive or false negative results. |  |  | Term was added on 18 March 2020 as an additional level 3 term to "Incorrect, Inadequate or Imprecise Results or Readings". Term modified for Release Number 2025. Previous term was Erratic Results. See no. 46 of Change Request Log. | A09¦A0908¦A090813
|  | Poor Visibility | A090814 | Implanted device is difficult to detect on imaging modalities, e.g. insufficient visibility. |  |  | Term was added on 27 January 2022. For details, see comment No. 146 of the Change Log (Release Number 2022). | A09¦A0908¦A090814
| Unexpected Therapeutic Results |  | A0909 | Problem associated with the use of the device for therapeutic purposes. | MedDRA:10043412:Therapeutic effect unexpected |  |  | A09¦A0909
Temperature Problem |  |  | A10 | Problem associated with the device producing unintended temperatures. | MedDRA:10093225:Device temperature production issue |  |  | A10
| Excessive Cooling |  | A1001 | Problem associated with the device producing temperatures that are lower than specified. | MedDRA:10093230:Device cool temperature production excessive |  |  | A10¦A1001
| Excessive Heating |  | A1002 | Problem associated with the device which have a warming or heating function, producing excessive heat. | MedDRA:10093228:Device heat production excessive |  |  | A10¦A1002
| Insufficient Cooling |  | A1003 | Problem associated with the device, which has a cooling function, producing temperatures that are not as low as specified. | MedDRA:10093227:Device cool temperature production insufficient | Modified (technical) | Technical change to definition for Release Number 2026. See no. 382 of Change Log. | A10¦A1003
| Insufficient Heating |  | A1004 | Problem associated with the device or its components producing temperatures that are not as high as what is specified. | MedDRA:10093231:Device heat production insufficient |  |  | A10¦A1004
| Overheating of Device |  | A1005 | Problem associated with the device producing high temperatures, such that its operation is compromised or harm is caused (e.g. overheating that produces melting of components or automatic shutdown). | MedDRA:10083780:Device overheating |  |  | A10¦A1005
| Thermal Decomposition of Device |  | A1006 | Problems associated with a discoloration or destruction as a result of thermal decomposition of the device. | MedDRA:10083762:Device thermal decomposition |  |  | A10¦A1006
| Fire |  | A1007 | Problem associated with the combustion of the device with a steady flame. | MedDRA:10078265:Device catching fire |  |  | A10¦A1007
| Flare or Flash |  | A1008 | Problem associated with device-related burn with an unsteady flame. | MedDRA:10078265:Device catching fire |  |  | A10¦A1008
| Smoking |  | A1009 | Problem associated with a cloud of vapor or gas generated from the device, generally associated after a fire or a burn. | MedDRA:10083771:Device emitting smoke |  |  | A10¦A1009
Computer Software Problem |  |  | A11 | Problem associated with written programs, codes, and/or software system (including AI/ML) that affects device performance or communication with another device | MedDRA:10069849:Device computer software issue | Modified (technical) | Technical change to definition for Release Number 2026. See no. 116 of Change Log. | A11
| Application Network Problem |  | A1101 | Problem associated with the deviations from documented system specifications that affects overall system performance and/or the performance of an individual device connected to that system. | MedDRA:10093159:Device application network issue |  |  | A11¦A1101
| Application Program Problem |  | A1102 | Problem associated with the requirement for software to fulfill its function within an intended use or application. | MedDRA:10083782:Device software design issue |  |  | A11¦A1102
|  | Application Program Freezes or Fails to Launch | A110201 | Problem associated with an application program which freezes and/or fails to launch. |  |  | Definition was modified on 27 January 2022 to add in "/or" For details, see comment No. 27 of the Change Log (Release Number 2022). Term and Definition modified for Release Number 2025. Previous term was “Application Program Freezes, Becomes Nonfunctional”. See no. 6 of Change Request Log. | A11¦A1102¦A110201
|  | Application Program Problem: Dose Calculation Error | A110202 | Problem associated with the written program code or application software used by the device to calculate specific measurements or quantities managed by the device. |  |  |  | A11¦A1102¦A110202
|  | Application Program Problem: Medication Error | A110203 | Event in which the device software results in errors of medication preparation or administration. |  |  |  | A11¦A1102¦A110203
|  | Application Program Problem: Parameter Calculation Error | A110204 | Problem associated with the written program code or application software used by the device to calculate parameters other than those related to dose or power. |  |  |  | A11¦A1102¦A110204
|  | Application Program Problem: Power Calculation Error | A110205 | Problem associated with the written program code or application software used by the device for calculations related to device power. |  |  |  | A11¦A1102¦A110205
|  | Application Program Version or Upgrade Problem | A110206 | Problem associated with installing updates to a software system that affects the device performance or communication with another device. |  |  |  | A11¦A1102¦A110206
|  | Problem with Software Installation | A110207 | Problem associated with installing the device software in a manner that allows full functioning of the device. Source of installation could be manufacturer or user. |  |  |  | A11¦A1102¦A110207
|  | Unintended Application Program Shut Down | A110208 | Problem associated with an unintended shut down by malfunction of the application program. |  |  |  | A11¦A1102¦A110208
| Program or Algorithm Execution Problem |  | A1103 | Problem associated with execution problems relating to program or algorithm. | MedDRA:10069849:Device computer software issue |  |  | A11¦A1103
|  | Delayed Program or Algorithm Execution | A110301 | Problem associated with delayed execution relating to program or algorithm. |  |  |  | A11¦A1103¦A110301
|  | Intermittent Program or Algorithm Execution | A110302 | Problem associated with intermittent execution relating to program or algorithm. |  |  |  | A11¦A1103¦A110302
|  | Program or Algorithm Execution Failure | A110303 | Problem associated with the failure of a program or algorithm to execute. Sudden/unexpected interruption to a program's execution. |  |  |  | A11¦A1103¦A110303
| Computer Operating System Problem |  | A1104 | Problem associated with software, firmware, and/or hardware elements that control the execution of computer programs and provides such services as computer resource allocation, job control, input/output control, and file management in a computer system. | MedDRA:10093152:Device operating system issue |  |  | A11¦A1104
|  | Operating System Becomes Nonfunctional | A110401 | Problem associated with malfunction of the computer operating system as opposed to an application software problem. |  |  |  | A11¦A1104¦A110401
|  | Operating System Version or Upgrade Problem | A110402 | Problem associated with replacing an older operating system to an up-to-date operating system. |  |  |  | A11¦A1104¦A110402
| Computer System Security Problem |  | A1105 | Problem associated with unauthorized access to or modification of a software system resulting in a loss of confidentiality, integrity, or availability of written program code, application software, or data or entire device. | MedDRA:10069849:Device computer software issue |  |  | A11¦A1105
|  | Application Security Problem | A110501 | Problem associated with the acquisition of computer programming codes that can replicate and spread from one computer system to another thereby leading to damaged software, hardware and data. |  |  |  | A11¦A1105¦A110501
|  | Unauthorized Access to Computer System | A110502 | Problem associated with an access that was not permitted to the computer system that may lead to modification of program, corruption of data, or and break in network security. This concept is closely associated with computer integrity which is the degree to which a system or component prevents unauthorized access to, or modification of, computer programs or data. |  |  |  | A11¦A1105¦A110502
| Data Back-Up Problem |  | A1106 | Problems relating to a system, component, file, procedure, or person available to replace or help restore a primary item in the event of a failure or externally caused disaster. | MedDRA:10083710:Device data issue |  |  | A11¦A1106
|  | Failure to Back-Up | A110601 | Problem associated with the inability to backup or to retrieve a backed up version (corrupted file) of device data or system files. |  |  |  | A11¦A1106¦A110601
|  | Failure to Convert to Back-Up | A110602 | Problem associated with a failure to transition from a primary system, component, file, procedure to a backup in response to a failure in the primary item. |  |  |  | A11¦A1106¦A110602
| Data Problem |  | A1107 | Event in which data (charting, orders, results) is not correctly stored, transferred, updated, or displayed. | MedDRA:10083710:Device data issue |  |  | A11¦A1107
|  | Loss of Data | A110701 | Event in which data is unintentionally permanently or temporarily lost, deleted, corrupted, or overwritten. |  |  |  | A11¦A1107¦A110701
|  | Patient Data Problem | A110702 | Event in which data is accessed by the healthcare provider and either the wrong patient or the wrong data is retrieved despite correct inquiry procedures. |  |  |  | A11¦A1107¦A110702
| Date/Time-Related Software Problem |  | A1108 | Problem associated with programming of calendar dates and/or time as a factor in the operation of the device. | MedDRA:10083681:Device date-time programming issue |  |  | A11¦A1108
Connection Problem |  |  | A12 | Problem associated with linking of the device and/or the functional units set up to provide means for a transfer of liquid, gas, electricity or data. | MedDRA:10065066:Device connection issue |  |  | A12
| Blocked Connection |  | A1201 | Problem associated with linking of the device whereby their functional units set up to provide means for a transfer of fluid, gas, or data is prevented or impeded. | MedDRA:10065067:Connector block problem |  |  | A12¦A1201
| Decoupling |  | A1202 | Problem associated with the device being unassociated in such a way that fluid, gas, power or signal information may not be transferred from one to another. | MedDRA:10065066:Device connection issue |  |  | A12¦A1202
| Disconnection |  | A1203 | Problem associated with the linking of the device having a sufficient open space to prevent gas, liquid or electrical current flow between connectors. | MedDRA:10083684:Device disconnection issue |  |  | A12¦A1203
| Failure to Disconnect |  | A1204 | Problem associated with the linking of the device whereby termination of the transfer of liquid, gas, electricity, or information cannot be accomplished, or linking components do not come apart, or disconnect, when expected. | MedDRA:10083772:Device failure to disconnect |  |  | A12¦A1204
| Loose or Intermittent Connection |  | A1205 | Problem associated with the connection of the device being loose or intermittent. | MedDRA:10069848:Device connection loose |  |  | A12¦A1205
| Misconnection |  | A1206 | Problem associated with the connection of the device being improper or not in accordance with device specification, requirements or intended uses. | MedDRA:10065066:Device connection issue |  |  | A12¦A1206
| Incomplete or Inadequate Connection |  | A1207 | Problem associated with a partial linking of the device whereby device may appear to be connected however only a partial, intermittent or no transfer of liquid, gas, electricity, or information can be accomplished. | MedDRA:10065066:Device connection issue |  |  | A12¦A1207
| Fitting Problem |  | A1208 | Problem associated with the connection of the device whereby channels, switching systems, and other functional units set up to provide means for a transfer of liquid, gas, electricity, or information do not match or fit. | MedDRA:10083679:Device connection fit issue |  |  | A12¦A1208
Communication or Transmission Problem |  |  | A13 | Problem associated with the device sending or receiving signals or data. This includes transmission among internal components of the device to which the device is intended to communicate. | MedDRA:10083730:Device signal transmission issue |  |  | A13
| Failure to Read Input Signal |  | A1301 | Problem associated with a failure of the device to read a signal for interpretation or measurement. | MedDRA:10083675:Device signal detection failure |  |  | A13¦A1301
| Failure to Transmit Record |  | A1302 | Problem associated with a failure of the device to transmit a record for interpretation or measurement. | MedDRA:10083730:Device signal transmission issue |  |  | A13¦A1302
| Intermittent Communication Failure |  | A1303 | Inconsistent or lack of intended communication of data among internal components or with other external devices. | MedDRA:10093147:Device data communication failure |  |  | A13¦A1303
| Telemetry Discrepancy |  | A1304 | Problem associated with variability of the transmission of telemetry signals. | MedDRA:10083730:Device signal transmission issue |  |  | A13¦A1304
| Wireless Communication Problem |  | A1305 | Problems with the RF wireless technology characteristics and performance (e.g., frequency, output power, range, reception), wireless quality of service, wireless coexistence, security of wireless signals and data, and electromagnetic compatibility. | MedDRA:10083731:Device wireless communication issue |  |  | A13¦A1305
Infusion or Flow Problem |  |  | A14 | Problem associated with the device failing to deliver or draw liquids or gases as intended (e.g. delivering drugs at incorrect rate, problems with drawing fluid from a system). This includes vacuum collection devices and manual or mechanical pumps. | MedDRA:10070617:Device infusion issue |  | Definition was modified on 31 January 2023 (Old definition: Problem associated with the device failing to deliver or draw liquids or gases as intended (e.g. delivering drugs at incorrect rate, Problems with drawing fluid from a system). This includes vacuum collection devices.). For details, see comment No. 35 of the Change Log (Release Number 2023). | A14
| Deflation Problem |  | A1401 | Problem associated with the deflation of the device including unintended deflation or inadequate deflation. | MedDRA:10093160:Device deflation issue |  | Technical change to the definition on 18 March 2020 to accommodate additional level 3 terms. | A14¦A1401
|  | Failure to Deflate | A140101 | Problem associated with the inability to deflate a device which is intended to deflate. |  |  | Term was added on 18 March 2020 to differentiate between a failure to deflate and an unintended deflation. | A14¦A1401¦A140101
|  | Unintended Deflation | A140102 | Problem associated with the unintended deflation of a device which was not designed to deflate or not at that point in time or not to that extent. |  |  | Term was added on 18 March 2020 to differentiate between a failure to deflate and an unintended deflation. | A14¦A1401¦A140102
| Excess Flow or Over-Infusion |  | A1402 | Problem associated with a delivery overdose of therapeutic agents, such as drugs or fluids being delivered into a device or a patient. | MedDRA:10083641:Device overinfusion |  |  | A14¦A1402
| Filling Problem |  | A1403 | Problem associated with the method or amount of time associated with the delivery of a fluid. Time to delivery or amount of delivered entity may be affected. | MedDRA:10083653:Device delivery system filling issue |  |  | A14¦A1403
|  | Inability to Auto-Fill | A140301 | Complete failure to fill as part of an automated process. For insufficient filling use "Short Fill". For excessive filling use "Overfill". For inconsistent filling use "Volume Accuracy Problem". |  |  |  | A14¦A1403¦A140301
|  | Overfill | A140302 | Excessive filling of a device. For complete failure to fill use "inability to auto-fill". For insufficient filling use "Short Fill". For inconsistent filling use "Volume Accuracy Problem". |  |  |  | A14¦A1403¦A140302
|  | Short Fill | A140303 | Insufficient filling of a device. For complete failure to fill use "inability to auto-fill". For insufficient filling use "short fill". For inconsistent filling use "Volume Accuracy Problem". |  |  |  | A14¦A1403¦A140303
|  | Volume Accuracy Problem | A140304 | Inconsistent filling of a device. This describes a problem which is observed to vary between overfilling and under filling, and may be intermittent. Use "Overfill" or "Short Fill" if problem is consistent. |  |  |  | A14¦A1403¦A140304
| Filtration Problem |  | A1404 | Problem associated with the process of passing a substance through a porous medium, e.g., a blood clot filter for the removal of suspended matter. | MedDRA:10093139:Device filtration issue |  |  | A14¦A1404
|  | Inadequate Filtration Process | A140401 | Problem associated with the filter failing to remove items or substances which should have been removed. |  |  |  | A14¦A1404¦A140401
|  | Inadequate Ultra Filtration | A140402 | Problem associated with the transfer of fluid between the blood and dialysate through the dialysis membrane due to a pressure gradient (trans-membrane pressure) existing between the blood and dialysate compartments. |  |  |  | A14¦A1404¦A140402
| Improper Flow or Infusion |  | A1405 | Problem associated with the regulation and delivery of therapeutic agents (e.g. air, gas, drugs or fluids into a device or a patient under positive pressure). | MedDRA:10083686:Device delivery system improper flow |  |  | A14¦A1405
|  | Backflow | A140501 | Continuous flow of fluid (e.g. liquid, gas) against the intended flow direction. |  |  |  | A14¦A1405¦A140501
|  | Free or Unrestricted Flow | A140502 | Problem associated with uncontrolled flow of infusion of air, gas or fluids. |  |  |  | A14¦A1405¦A140502
|  | Gradient Increase | A140503 | Problem associated with the increased rate of change in temperature, pressure, or other variables as a function of distance, time, etc. |  |  |  | A14¦A1405¦A140503
|  | Inaccurate Delivery | A140504 | Delivery at endpoint not as intended; either too low or too high. |  |  |  | A14¦A1405¦A140504
|  | Inaccurate Flow Rate | A140505 | Problem associated with fluctuations in the flow volume delivered per time, even if end volume is correct, and delivered in the correct total time. |  |  |  | A14¦A1405¦A140505
|  | Intermittent Infusion or Flow | A140506 | Problem associated with the flow not being steady characterised by intermittent stoppages to the flow. |  | Modified (technical) | Change to term and definition for Release Number 2026. See no. 126 of Change Log. | A14¦A1405¦A140506
|  | Reflux Within Device | A140507 | Problem associated with partial backflow, compromising the device's flow output. |  | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 222 of Change Log. | A14¦A1405¦A140507
|  | Restricted Flow Rate | A140508 | Problem associated with flow rate. Flow volume delivered over time is not reaching intended flow rate. |  | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 223 of Change Log. | A14¦A1405¦A140508
|  | Tidal Volume Fluctuations | A140509 | Problem associated with the amount of gas that is inspired and expired during one respiratory cycle. |  |  |  | A14¦A1405¦A140509
| Inflation Problem |  | A1406 | Problem associated with the inability of the device to expand or enlarge with the intended inflation agent (e.g. saline or air). | MedDRA:10093146:Device inflation issue |  |  | A14¦A1406
| Insufficient Flow or Under Infusion |  | A1407 | Problem associated with an insufficient dose of therapeutic agents, e.g., drugs or fluids being delivered into a patient under positive pressure. | MedDRA:10083649:Device underinfusion |  |  | A14¦A1407
| No Flow |  | A1408 | Problem arising from the device failing to deliver the specified liquid or gas. | MedDRA:10083677:Device delivery system no flow |  |  | A14¦A1408
|  | Failure to Deliver | A140801 | Failure (=complete nonperformance) with regard to the intended function of delivery. |  |  |  | A14¦A1408¦A140801
|  | Failure to Infuse | A140802 | Failure (=complete nonperformance) with regard to the intended function of infusion. |  |  |  | A14¦A1408¦A140802
|  | Inability to Irrigate | A140803 | Failure (=complete nonperformance) with regard to the intended function of irrigation. |  |  |  | A14¦A1408¦A140803
| Obstruction of Flow |  | A1409 | Problem related to an obstruction or blockage within the device component (e.g. tube, opening, pipe) that results in restriction of flow (including blood clotting). | MedDRA:10064685:Device occlusion |  | Technical change to the definition on 18 March 2020 to clarify that the term includes problems related to blood clotting. | A14¦A1409
|  | Complete Blockage | A140901 | Problem related to an obstruction or blockage within the device component (e.g. tube, opening, pipe) that results in no flow. |  |  |  | A14¦A1409¦A140901
|  | Partial Blockage | A140902 | Problem related to an obstruction or blockage within the device component (e.g. tube, opening, pipe) that results in a reduction of the flow rate. |  |  |  | A14¦A1409¦A140902
| Difficult to Flush |  | A1410 | The device is difficult to flush, possibly indicating an obstruction within device. | MedDRA:10050358:Device blockage |  |  | A14¦A1410
| Pressure Problem |  | A1411 | Problem associated with the application of a force either internal or external to device that compromises the flow of fluid or gas. | MedDRA:10083760:Device delivery system pressure issue |  |  | A14¦A1411
|  | Decrease in Pressure | A141101 | Unintended decrease in pressure, compromising the device's intended function. |  |  |  | A14¦A1411¦A141101
|  | Increase in Pressure | A141102 | Unintended increase in pressure, compromising the device's intended function. |  |  |  | A14¦A1411¦A141102
|  | No Pressure | A141103 | Unintended complete loss of pressure, compromising the device's intended function. |  |  |  | A14¦A1411¦A141103
| Pumping Problem |  | A1412 | Problem associated with pump performance deviating from specifications in a way to compromise flow or infusion or dispensing. | MedDRA:10083735:Pump infusion issue |  | Definition was modified on 31 January 2023 (Old definition: Problem associated with pump performance deviating from specifications in a way to compromise flow or infusion). For details, see comment No. 35 of the Change Log (Release Number 2023). | A14¦A1412
|  | Decreased Pump Speed | A141201 | Unintended decrease in pump speed and hence, probably, flow rate, compromising the intended function of the device. |  |  |  | A14¦A1412¦A141201
|  | Increased Pump Speed | A141202 | Unintended increase in pump speed and hence, probably, flow rate, compromising the intended function of the device. |  |  |  | A14¦A1412¦A141202
|  | Failure to Pump | A141203 | Problem associated with the device which fails to start pumping. |  |  |  | A14¦A1412¦A141203
|  | Pumping Stopped | A141204 | Unexpected/unintended cessation of pump. |  |  |  | A14¦A1412¦A141204
| Suction Problem |  | A1413 | Problem associated with suction equipment, which may be a manual, electrical, vacuum or pressure source operated to evacuate and remove undesired substances (air, gas, fluid, or particulates) via tubing and collection bag. | MedDRA:10083980:Device suction event |  |  | A14¦A1413
|  | Decrease in Suction | A141301 | Problem associated with the removal of fluid or gas from a body cavity or device due to decreased suction. |  |  | Technical change to the definition on 18 March 2020 to include removal of fluid or gas from a device. | A14¦A1413¦A141301
|  | Increase in Suction | A141302 | Problem associated with the removal of fluid or gas from a body cavity or device due to increased suction. |  |  | Technical change to the definition on 18 March 2020 to include removal of fluid or gas from a device. | A14¦A1413¦A141302
|  | Suction Failure | A141303 | Problem associated with the complete inability to provide suction. |  |  |  | A14¦A1413¦A141303
| Priming Problem |  | A1414 | Problem associated with the preparation of the device to begin pumping. | MedDRA:10068944:Device failure to prime |  |  | A14¦A1414
|  | Failure to Prime | A141401 | Problem associated with the device failing to begin the priming process (i.e. the process of preparation of device for the delivery of fluids). |  |  |  | A14¦A1414¦A141401
|  | Incomplete or Inadequate Priming | A141402 | Problem associated with not adequately preparing the device. |  |  |  | A14¦A1414¦A141402
| Air/Gas in Device |  | A1415 | Problem associated with air or gas in the device (e.g. air in line). | MedDRA:10093142:Air in device issue | Modified (technical) | Term was added on 18 March 2020 as an additional level 2 term to "Infusion or Flow Problem". Technical change to definition included for Release Number 2026. See no. 52 of Change Log. | A14¦A1415
Activation, Positioning or Separation Problem |  |  | A15 | Problem associated with any deviations from the intended performance of the device that relate to the sequence of events for activation, positioning or separation of device. Note: Deployment is synonymous with activation. | MedDRA:10070691:Device deployment issue |  | Editorial change to the term on 16 Feburary 2021. Definition modified for Release Number 2025. See no. 43 of Change Request Log | A15
| Activation Problem |  | A1501 | Problem associated with the activation of the device. Note: Activation includes expansion. | MedDRA:10070691:Device deployment issue |  | Technical change to the definition on 18 March 2020 to clarify that activation includes expansion. | A15¦A1501
|  | Activation Failure | A150101 | Problem associated with the device failing to be activated including partial activation. |  |  | Editorial change to the definition on 18 March 2020. | A15¦A1501¦A150101
|  | Difficult or Delayed Activation | A150102 | Problem associated with delayed or difficult activation of the device. |  |  |  | A15¦A1501¦A150102
|  | Premature Activation | A150103 | Problem associated with early and unexpected activation of the device. |  |  |  | A15¦A1501¦A150103
|  | Self-Activation or Keying | A150104 | Problem associated with the unexpected or unintended activation of the device or a device function. For Unintended Power Up use A070804 |  |  | Definition modified for Release Number 2025. See no. 3 of Change Request Log. | A15¦A1501¦A150104
|  | Key or Button Unresponsive/not Working | A150105 | Problem associated with a device not responding to key or button inputs. |  |  | Term was added on 18 March 2020 as an additional level 3 term to "Activation Problem". | A15¦A1501¦A150105
|  | Unexpected Deactivation | A150106 | Problem associated with a device deactivating during use. |  |  | Term added for Release Number 2025. See no. 8 of Change Request Log | A15¦A1501¦A150106
| Positioning Problem |  | A1502 | Problem associated with the movement of the device to an intended location. | MedDRA:10070085:Device misdeployment (location) |  |  | A15¦A1502
|  | Positioning Failure | A150201 | Problem associated with the inability of the device to be positioned in a specified location. |  |  |  | A15¦A1502¦A150201
|  | Malposition of Device | A150202 | Problem associated with the device being positioned in a location other than intended or specified. |  |  |  | A15¦A1502¦A150202
|  | Difficult or Delayed Positioning | A150203 | Problem associated with users experiencing difficulty or delay to position the device to a specified location. |  |  |  | A15¦A1502¦A150203
|  | Failure to Advance | A150204 | Problem associated with failure to move the device to an intended location. |  |  |  | A15¦A1502¦A150204
|  | Difficult to Advance | A150205 | Problem associated with difficulty moving the device to an intended location (e.g. difficulty in advancing guide wire). |  |  |  | A15¦A1502¦A150205
|  | Difficult to Insert | A150206 | Problem associated with problems introducing or inserting the device, even if the user is operating the device in accordance with the instructions for use or labeling. |  |  |  | A15¦A1502¦A150206
|  | Difficult to Remove | A150207 | Problem associated with the user experiencing difficulty when taking out the device, even if the user is operating device in accordance with the instructions for use or labeling. |  |  | Editorial change to the definition on 18 March 2020. | A15¦A1502¦A150207
|  | Entrapment of Device | A150208 | Problem associated with the device caught within patient vasculature, tissue, or other device. |  |  |  | A15¦A1502¦A150208
| Separation Problem |  | A1503 | Problem associated with the detachment or separation of the device. | MedDRA:10065066:Device connection issue |  |  | A15¦A1503
|  | Separation Failure | A150301 | Problem associated with the device or one of its components failing to detach or separate as intended. |  |  |  | A15¦A1503¦A150301
|  | Difficult or Delayed Separation | A150302 | Problem associated with users experiencing difficulty or delay with detachment or separation of the device. |  |  |  | A15¦A1503¦A150302
|  | Premature Separation | A150303 | Problem associated with an early and unexpected detachment or separation of the device from the system. |  |  |  | A15¦A1503¦A150303
Protective Measures Problem |  |  | A16 | Problem associated with any deviations from the documented specifications of the device that relate to the implemented and inherited design features specific to devices used for reducing risks to patient or caregiver or maintaining risks within specified levels. | MedDRA:10083741:Device safety protective measures issue |  |  | A16
| Device Alarm System |  | A1601 | Problem associated with the alarm system of the device. | MedDRA:10069841:Device alarm issue |  |  | A16¦A1601
|  | Alarm not Visible | A160101 | The device does not display an alarm message when required. |  | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 224 of Change Log. | A16¦A1601¦A160101
|  | No Audible Alarm | A160102 | The device fails to emit an audible alarm. |  |  |  | A16¦A1601¦A160102
|  | Low Audible Alarm | A160103 | The audible device alarm cannot be heard clearly. |  |  |  | A16¦A1601¦A160103
|  | Delayed Alarm | A160104 | The device alarm system operates with delay. |  |  |  | A16¦A1601¦A160104
|  | False Alarm | A160105 | Problem associated with the device providing incorrect alarm warning or alert to user. |  |  |  | A16¦A1601¦A160105
|  | Defective Alarm | A160106 | The device alarm does not operate as expected and/or in agreement with device's specifications. |  |  |  | A16¦A1601¦A160106
| Fail-Safe Problem |  | A1602 | Problem associated with the feature that prevents the unsafe use of the device. | MedDRA:10083650:Device fail-safe mechanism issue |  |  | A16¦A1602
|  | Fail-Safe Did not Operate | A160201 | Problem associated with the device fail-safe mechanism, which did not function or function in a non effective way, compromising safe use of the device. |  | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 225 of Change Log. | A16¦A1602¦A160201
|  | No Fail-Safe Mechanism | A160202 | The device does not have a fail-safe mechanism, although such mechanism would be required for its appropriate and/or safe functioning. |  |  |  | A16¦A1602¦A160202
| Failure of Device to Self-Test |  | A1603 | Problem associated with the device failing to perform an internal self-diagnostic process to ensure normal operation during or prior to use. | MedDRA:10083777:Device failure to self-test |  |  | A16¦A1603
| Failure to Auto Stop |  | A1604 | Problem associated with the inability of device to turn itself off when the device is not in an operable condition. | MedDRA:10093148:Device failure to auto stop |  |  | A16¦A1604
| Reset Problem |  | A1605 | Problem associated with setting a variable, register, or other storage location back to a prescribed state. | MedDRA:10083734:Device reset issue |  |  | A16¦A1605
|  | Failure to Reset | A160501 | Problem associated with the device failing to set a variable, register, or other storage location back to a prescribed state. |  |  |  | A16¦A1605¦A160501
|  | Failure to Zero | A160502 | Problem associated with the device failing to set a variable, register, or other storage location back to zero. |  |  |  | A16¦A1605¦A160502
|  | Inappropriate or Unexpected Reset | A160503 | Problem associated with the device setting a variable, register, or other storage location to an inappropriate or unexpected state. |  |  |  | A16¦A1605¦A160503
| Premature Indicator Activation |  | A1606 | Problems with the activation of a protective measure indicator earlier than expected. | MedDRA:10090319:Safety mechanism activated prematurely |  |  | A16¦A1606
|  | Premature Elective Replacement Indicator | A160601 | Problems with the early or unexpected activation of the elective replacement indicator. |  |  |  | A16¦A1606¦A160601
|  | Premature End-of-Life Indicator | A160602 | Problem with the early or unexpected activation of the end-of-life indicator. |  |  |  | A16¦A1606¦A160602
| Shielding Failure |  | A1607 | Problem associated with the device inability to act as a barrier for absorption of radiation energy in X-rays, gamma rays, etc. | MedDRA:10056871:Device failure |  |  | A16¦A1607
Compatibility Problem |  |  | A17 | Problem associated with compatibility between device or substances (medication, body fluid, etc.) | MedDRA:10093224:Device incompatibility issue | Modified (technical) | Technical change to definition for Release Number 2026. See no. 8 of Change Log. | A17
| Component or Accessory Incompatibility |  | A1701 | Problem associated with the incompatibility of any device while being operated in the same use environment thereby leading to a dysfunction between the devices. | MedDRA:10069845:Device-device incompatibility |  |  | A17¦A1701
|  | Accessory Incompatible | A170101 | An accessory required for the intended purpose of the device appears incompatible with device, thus compromising the intended function of the device. |  |  |  | A17¦A1701¦A170101
|  | Component Incompatible | A170102 | A component required for the proper functioning of the device is not compatible with other components or subassemblies of the device, thus compromising the intended function of the device. |  |  |  | A17¦A1701¦A170102
| Device-Device Incompatibility |  | A1702 | Problem associated with the incompatibility of two or more devices while being operated in the same use environment thereby leading to a dysfunction of more than one device. | MedDRA:10069845:Device-device incompatibility |  |  | A17¦A1702
| Measurement System Incompatibility |  | A1703 | Problem associated with the incompatibility of the measurement systems between and/or within device systems that are inherent to the individual device thereby leading to miscalculated or mismatched measurements from those devices, e.g., international metric system versus U.S. measurement system. | MedDRA:10093224:Device incompatibility issue |  |  | A17¦A1703
| Unintended Compatibility |  | A1704 | Problem associated with the ability of two or more devices which are intended to be incompatible but are able to work or fit together. | MedDRA:10093140:Unintended device compatibility issue |  |  | A17¦A1704
Contamination /Decontamination Problem |  |  | A18 | Problem associated with the presence of any unexpected foreign substance found in the device, on its surface or in the package materials, which may affect performance or intended use of the device, or problem that compromise effective decontamination of the device. | MedDRA:10069173:Product contamination |  |  | A18
| Contamination |  | A1801 | Problem associated with the undesired introduction of impurities either chemical or microbiological in nature, or of foreign matter into or onto the device. | MedDRA:10069173:Product contamination |  | Technical change to the term on 18 March 2020 to separate when/where contamination occurred from type of contamination ("During Use" was removed). NOTE: This change affects all lower level terms (A180101 - A180105) which are also no longer limited to contamination during use. | A18¦A1801
|  | Biofilm Coating in/on Device | A180101 | Problem associated with the undesired introduction of a biofilm coating into or onto the device. |  | Modified (technical) | Editorial and technical change to term for Release Number 2026. See no. 56 and 227 of Change Log. | A18¦A1801¦A180101
|  | Contamination of Device Ingredient or Reagent | A180102 | Problem associated with the undesired introduction of impurities either chemical or microbiological in nature, or of foreign matter into or onto the device ingredient or reagent. |  |  |  | A18¦A1801¦A180102
|  | Device Contamination with Body Fluid | A180103 | Problem associated with the undesired presence of body fluid in/on the device, which are not part of the documented device specifications and requirements. |  |  |  | A18¦A1801¦A180103
|  | Device Contamination with Chemical or Other Material | A180104 | Problem associated with contamination of the device with chemical substance or other non biologic material. |  |  |  | A18¦A1801¦A180104
|  | Microbial Contamination of Device | A180105 | Problem associated with undesired microbial contamination of the device. |  |  |  | A18¦A1801¦A180105
| Device Contaminated During Manufacture or Shipping |  | A1802 | Problem associated with the presence of any unexpected foreign substance found on the surface or in the package materials, which may affect optimal performance for its intended use, and which occurred during manufacture or shipping. | MedDRA:10083778:Device contamination prior to use |  | Technical change to the definition on 18 March 2020 to clarify that contamination occurred during manufacture or shipping. | A18¦A1802
| Device Reprocessing Problem |  | A1803 | Problem associated with a failure during any step of reprocessing process (cleaning, disinfection, packaging, labeling, sterilization) of a used or opened from its original packaging, but unused, device. | MedDRA:10093153:Device contamination issue during reprocessing |  |  | A18¦A1803
|  | Failure to Clean Adequately | A180301 | Problem associated with the failure of the device or operator to remove any visible soil, foreign material or organism deposits on the external surfaces, crevices, and joints of the device. |  |  |  | A18¦A1803¦A180301
|  | Failure to Disinfect | A180302 | Failure to properly disinfect the device when reprocessing it. |  |  |  | A18¦A1803¦A180302
|  | Flushing Problem | A180303 | Flushing process was not executed properly. |  |  |  | A18¦A1803¦A180303
|  | Problem with Removal of Enzymatic Cleaner | A180304 | Enzymatic cleaner was not removed properly. |  |  |  | A18¦A1803¦A180304
|  | Problem with Sterilization | A180305 | Device was not sterilized properly during reprocessing. |  |  |  | A18¦A1803¦A180305
|  | Residue After Decontamination | A180306 | Problem associated with the decontamination process not adequately removing unwanted visible soil, foreign material, or organism deposits. |  |  |  | A18¦A1803¦A180306
| Device Contaminated at the User Facility |  | A1804 | Problem associated with the presence of any unexpected foreign substance found on the surface or in the package materials, which may affect optimal performance for its intended use, and which occurred at the user facility. | MedDRA:10083768:Device contamination during use |  | Term was added on 18 March 2020 for coding of contamination at the user facility which was originally included in "Contamination During Use" (A1801). | A18¦A1804
Environmental Compatibility Problem |  |  | A19 | Problem associated with the surrounding conditions in which the device is being used such as temperature, noise, lighting, ventilation, or other external factors such as power supply. | MedDRA:10083673:Device environmental compatibility issue |  |  | A19
| Ambient Noise Problem |  | A1901 | Problem associated with any undesired acoustic energy or vibration that tends to interfere with the operation of the device. | MedDRA:10083673:Device environmental compatibility issue |  |  | A19¦A1901
| Ambient Temperature Problem |  | A1902 | Problem associated with compromised device performance at the ambient temperature or the storage at an inappropriate ambient temperature. | MedDRA:10083673:Device environmental compatibility issue |  |  | A19¦A1902
| Fumes or Vapors |  | A1903 | Problem associated with the visibility, odor, or toxicity of an ambient vapor or gas. | MedDRA:10083673:Device environmental compatibility issue |  |  | A19¦A1903
| Fungus in Device Environment |  | A1904 | Problem associated with the visibility of molds, mildews, yeasts, and/or mushrooms in the immediate environment in which the device is being used. | MedDRA:10083673:Device environmental compatibility issue |  |  | A19¦A1904
| Moisture or Humidity Problem |  | A1905 | Problem associated with an unsatisfactory humidity level in the storage or use environment which affects the device performance. | MedDRA:10083683:Device environmental humidity issue |  |  | A19¦A1905
|  | Moisture Damage | A190501 | Problem associated with damage inflicted upon the device from water vapor or water in the immediate environment in which the device is being used. |  |  |  | A19¦A1905¦A190501
| Ventilation Problem in Device Environment |  | A1906 | Problem associated with the circulation of fresh air in the immediate atmosphere in which the device is being used. | MedDRA:10083673:Device environmental compatibility issue |  |  | A19¦A1906
|  | Fogging | A190601 | Problem associated with the visibility of water vapor in the immediate atmosphere in which the device is being used. |  |  |  | A19¦A1906¦A190601
| Device Unsafe to Use in Environment |  | A1907 | Problem associated with environmental condition that results in the unsafe use of the device. (e.g. electromagnetic fields, noise, vibration, microbiological contamination etc.) | MedDRA:10083673:Device environmental compatibility issue |  |  | A19¦A1907
| Environmental Particulates |  | A1908 | Problem associated with fine solids or liquid particles such as dust, smoke, fume, and/or mist suspended in the immediate atmosphere in which the device is being used. | MedDRA:10083673:Device environmental compatibility issue |  |  | A19¦A1908
| Medical Gas Supply Problem |  | A1909 | Problem associated with the facility-supplied medical gases such as medical air, oxygen, nitrous oxide, and nitrogen. | MedDRA:10083673:Device environmental compatibility issue |  |  | A19¦A1909
| Electrical Power Problem |  | A1910 | Problem associated with the quality of the facility-supplied power. | MedDRA:10083689:Device environmental power supply issue |  |  | A19¦A1910
|  | Emergency Power Failure | A191001 | Problem associated with the failure of the facility's emergency power backup system(s) including generators and/or interruptible power systems (UPS). |  |  |  | A19¦A1910¦A191001
|  | Loss of Power | A191002 | Problem associated with the failure of primary power supplied by the facility. |  |  |  | A19¦A1910¦A191002
|  | Power Conditioning Problem | A191003 | Problem associated with a momentary overpower/over voltage from the utility and electrical systems of user facilities; - Problem associated with inadequate power conditioning such as the presence of fluctuation, surges, spikes, dropouts, noise and other such undesirable transients. |  |  |  | A19¦A1910¦A191003
Installation-Related Problem |  |  | A20 | Problem associated with unsatisfactory installation, configuration, and/or setup of a specific device. | MedDRA:10083723:Device preparation issue |  |  | A20
| Misassembled During Installation |  | A2001 | Problem associated with the use of the device characterized by incorrect assembly of device components, parts or constituents. | MedDRA:10079999:Product assembly error during preparation for use |  |  | A20¦A2001
Labeling, Instructions for Use or Training Problem |  |  | A21 | Problem associated with device markings/labeling, instructions for use, training and maintenance documentation or guidelines. | MedDRA:10069289:Product label issue | Modified (editorial) | Editorial change to term and definition for Release Number 2026. See no. 180 of Change Log | A21
| Device Markings/Labeling Problem |  | A2101 | Problem associated with the written, printed or graphic material accompanying or affixed to the device or any of its packaging. This includes verbal instructions relating to identification, technical description, and usage provided by the device manufacturers. Problems can include but are not limited to this material being unclear, missing, worn out, incorrect or inaccurate. | MedDRA:10069289:Product label issue | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 181 of Change Log | A21¦A2101
|  | Expiration Date Problem | A210101 | Any problem involving the accuracy, legibility, presence, or interpretation of an expiration date on a medical device, its labeling, or associated documentation. |  | Modified (technical) | Technical change to term and definition for Release Number 2026. See no. 58 of Change Log | A21¦A2101¦A210101
|  | Illegible or Difficult to Read Information | A210102 | Problem associated with information unable to be read or deciphered. |  | Modified (technical) | Technical change to term for Release Number 2026. See no. 127 of Change Log | A21¦A2101¦A210102
|  | Inaccurate Information | A210103 | Problem associated with imprecise, incorrect inexact information. |  | Modified (editorial) | Editorial change to definition for Release Number 2026. See no. 59 of Change Log | A21¦A2101¦A210103
|  | Unclear Information | A210104 | Problem associated with ambiguous, confused information. |  |  |  | A21¦A2101¦A210104
|  | Missing Information | A210105 | Absence of information e.g. labeling, instruction for use. |  |  |  | A21¦A2101¦A210105
|  | Wrong Label | A210106 | Problem associated with an incorrect label being affixed to the device or packaging. |  |  | Term was added on 27 January 2022. For details, see comment No. 2 of the Change Log (Release Number 2022). | A21¦A2101¦A210106
|  | Device labeling look-alike | A210107 | Problem associated with medical device label that looks similar to another product leading to confusion. The other product may be another device or a different product such as a medicine or cleaning agent, or other product altogether. Use A020505 for packaging look-alike issues. |  | Modified (technical) | Term added for Release Number 2024. See no. 117 of Change Log. Technical and editorial change to term and definition for Release Number 2026. See no. 60, 61 and 228 of Change Log. | A21¦A2101¦A210107
| Lack of Maintenance Documentation or Guidelines |  | A2102 | Problem associated with user facility not receiving adequate service documentation, guidelines, or recommendations to perform preventative and corrective maintenance and performance assurance checks. | MedDRA:10069289:Product label issue |  |  | A21¦A2102
| Inadequate Instructions for Healthcare Professional |  | A2103 | Problem associated with inaccuracies in any written, printed, or graphic matter that is affixed to the device or its packaging with any matter that accompanies the device including verbal instructions related to identification, technical description and use of device provided by the device manufacturers that is intended for healthcare professionals. | MedDRA:10081132:Health care provider instructions for product use lacking |  |  | A21¦A2103
| Inadequate Instructions for Non-Healthcare Professional |  | A2104 | Problem associated with users being unclear and not able to follow any written, printed, or graphic matter that is affixed to device or its packaging with any matter that accompanies the device including verbal instructions related to identification, technical description and use of the device provided by the device manufactures that vary from the standard of medical care in a given environment. | MedDRA:10079006:Patient instructions for use in label confusing |  |  | A21¦A2104
| Inadequate or Insufficient Training |  | A2105 | Problem associated with facility not providing satisfactory initial and/or periodic user training covering operation of the device. | MedDRA:10080099:Product communication issue |  |  | A21¦A2105
Human-Device Interface Problem |  |  | A22 | Problem associated with the user experiencing difficulty operating or otherwise interacting with the device, even if the operation is being performed according to labeled instructions for use. | MedDRA:10083712:Device user interface issue |  |  | A22
| Device Difficult to Setup or Prepare |  | A2201 | Problem associated with the user experiencing difficulty in preparing the device for use, even if the operation is being performed according to labeled instructions for use. This includes difficulty attaching or connecting parts together. | MedDRA:10073159:Device difficult to handle |  |  | A22¦A2201
| Device Difficult to Program or Calibrate |  | A2202 | The device is difficult to program, calibrate or set to desired state, even by appropriately trained user/operator. | MedDRA:10083690:Device difficult to program |  |  | A22¦A2202
| Device Difficult to Maintain |  | A2203 | Problem associated with the user's ability to service the device according to the manufacturer specifications relating to the device routine maintenance, i.e., periodic inspection, failure detection, repair, and care of the device to sustain or restore acceptable operating conditions. | MedDRA:10069853:Device difficult to use |  |  | A22¦A2203
| Inadequate User Interface |  | A2204 | Problem associated with the means by which the operator and the equipment communicate or interact. | MedDRA:10083712:Device user interface issue |  |  | A22¦A2204
| Device Provides Uninformative Message |  | A2205 | The device provides a message which is unhelpful or uninformative, such as an error code without further information about the meaning or appropriate action to take. If the message does not make sense then use A2206 - Device Provides Unintelligible Message. | MedDRA:10069842:Device information output issue |  | Term added for Release Number 2024. See no. 42 of Change Log | A22¦A2205
| Device Provides Unintelligible Message |  | A2206 | The device provides an unintelligible or incomprehensible message such as an apparently random string of characters. If the message makes sense but is not helpful then use A2205 - Device Provides Uninformative Message. | MedDRA:10069842:Device information output issue |  | Term added for Release Number 2024. See no. 42 of Change Log | A22¦A2206
Use of Device Problem |  |  | A23 | Problem associated with failure to process, service, or operate the device according to the manufacturer's recommendations or recognized best practices. | MedDRA:10083996:Device operation use issue |  |  | A23
| Device Handling Problem |  | A2301 | Handling of the device not in accordance with specification, prior to use on the patient. | MedDRA:10089825:Device handling error |  |  | A23¦A2301
| Use of Incorrect Control/Treatment Settings |  | A2302 | Problem associated with the use of the device in terms of inappropriate or incorrect control settings or incorrect treatment parameters for the device's specified operation and/or intended use. | MedDRA:10083847:Use of incorrect device control settings |  | Term/Definition was modified on 27 January 2022 (Old term/definition: Use of Incorrect Control Settings: Problem associated with the use of the device in terms of inappropriate or false control settings for the device's specified operation and/or intended use). For details, see comment No. 26 of the Change Log (Release Number 2022). | A23¦A2302
| Improper or Incorrect Procedure or Method |  | A2303 | Problem associated with the use of the device in terms of nonconforming to that device's intended use, specifications, procedure and process or service instructions and information provided by the device manufacturers. | MedDRA:10083846:Improper use of device |  |  | A23¦A2303
| Off-Label Use |  | A2304 | Problem associated with the device which has been used for an unapproved indication or for an unapproved intended use. | MedDRA:10076481:Off label use of device |  |  | A23¦A2304
| Misassembled |  | A2305 | Problem associated with incorrect assembly of the device or constituents after being put into use. | MedDRA:10072878 Device use error |  |  | A23¦A2305
|  | Misassembly by Users | A230501 | Problem associated with incorrect assembly of the device or constituents by the users. |  |  |  | A23¦A2305¦A230501
|  | Misassembly During Maintenance/Repair | A230502 | Problem associated with incorrect assembly of the device or constituents during maintenance or repair. |  |  |  | A23¦A2305¦A230502
No Apparent Device Problem |  |  | A24 | An adverse event (e.g. patient harm) appears to have occurred, but there does not appear to have been a problem with the device or the way it was used. | MedDRA:10086414:No device malfunction | Modified (technical) | Technical change to term for Release Number 2026. See no. 194 of Change Log | A24
No Apparent Adverse Event |  |  | A25 | A report has been received but the description provided does not appear to relate to an adverse event. This code allows a report to be recorded for administration purposes, even if it doesn't meet the requirements for adverse event reporting. | MedDRA:10067482:No adverse event |  |  | A25
| Prophylactic Removal Due to Corrective Action |  | A2501 | Prophylactic removal of a device due to corrective action without malfunction or problem of the specific device involved | MedDRA:10052971:Medical device removal | Modified (editorial) | Term was added on 31 January 2023. For details, see comment No. 30 of the Change Log (Release Number 2023). Editorial change to term for Release Number 2026. See no. 229 of Change Log | A25¦A2501
Insufficient Device Problem Information |  |  | A26 | An adverse event appears to have occurred but there is not yet enough information available to classify the device problem. | MedDRA:10083845:Unevaluable device issue |  | Term modified for Release Number 2025. Previous term was Insufficient Information. See no. 120 of Change Request Log. | A26
Appropriate Device Problem Term/Code not Available |  |  | A27 | The device problem is not adequately described by any other term. Note: this code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. | MedDRA:10069868:Device issue | Modified (editorial) | Term modified for Release Number 2025. Previous term was "Appropriate Term/Code Not Available". See no. 107 of Change Request Log. Editorial change to term for Release Number 2026. See no. 230 of Change Log | A27

## B
Annex Name: Annex B
Annex Title: Cause Investigation - Type of Investigation
Release Number: 2026
Annex Description: Terms/codes for describing what was investigated and what kind of investigation was conducted to specify the root cause of the adverse event.
Annex Instructions: Terms/Codes in this annex should be used to answer questions such as How was the investigation performed? Did it involve testing? Did it also or only involve non-testing means (e.g. interviews, etc.) Wherever appropriate the word 'device' should be interpreted as meaning all or any of the following: medical device, in-vitro diagnostic medical device, accessory to a medical device, medical device parts and/or components. Wherever appropriate the word 'patient' should be interpreted as meaning all or any of the following: user, operator, any other person affected by the incident.
Annex Approval Date: January 2026
NOTE: Maximum field length Term: 100 characters Definition: 700 Status description: 700
Level 1 Term | Code | Definition | Non-IMDRF Code | Status | Status Description | CodeHierarchy
Testing of Actual/Suspected Device | B01 | The investigation employed relevant empirical testing of the actual device suspected in the reported adverse event in order to establish their functional and other properties and to identify possible causes for the adverse event. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. |  |  |  | B01
Testing of Device from Same Lot/Batch Retained by Manufacturer | B02 | The investigation employed relevant empirical testing of the device of the same lot or batch than that of the suspected device in the reported adverse event in order to support the identification of possible causes for the adverse event. Testing was performed using the device retained by the manufacturer (i.e. was not shipped). Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. |  |  |  | B02
Testing of Device from Same Lot/Batch Returned from User | B03 | The investigation employed relevant empirical testing of the device of the same lot or batch than that of the suspected device in the reported adverse event in order to support the identification of possible causes for the adverse event. The device was returned from the user. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. |  |  |  | B03
Testing of Device from Other Lot/Batch Retained by Manufacturer | B04 | The investigation employed relevant empirical testing of the device of another lot or batch than that of the suspected device in the reported adverse event in order to support the identification of possible causes for the adverse event. This includes devices without a lot/batch designation. Testing was performed using the device retained by the manufacturer (i.e. was not shipped). Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. |  |  |  | B04
Testing of Device from Other Lot/Batch Returned from User | B05 | The investigation employed relevant empirical testing of the device of another lot or batch than that of the suspected device in the reported adverse event in order to support the identification of possible causes for the adverse event. This includes devices without a lot/batch designation. The device was returned from the user. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the technical file or in the latest relevant standards. |  | Modified (editorial) | Term was modified for editorial consistency for Release Number 2026 See comment No. 231 of the Change Log. | B05
Testing of Model Variant | B06 | The investigation employed relevant empirical testing of a model variant of the device involved in the reported adverse event in order to support the identification of possible causes for the adverse event through plausibility reasoning. A model variant is not identical to the actual device, but shares relevant characteristics with the device involved. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. |  |  |  | B06
Testing of Raw/Starting Materials | B07 | The investigation employed relevant empirical testing of the materials used in construction of the device involved in the reported adverse event in order to support the identification of possible causes for the adverse event. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. |  |  |  | B07
Testing of Patient Sample or Reference Material Using Manufacturer's Device | B08 | The investigation employed relevant empirical testing of a patient sample or reference material using the device (usually an IVD) involved in the reported adverse event in order to support the identification of possible causes for the adverse event. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. |  | Retired (2025) | Term retired for Release Number 2025. Use B26 Testing of Patient Sample Using Manufacturer's Device or B27 Testing of Reference Material Using Manufacturer's Device. See no. 9 and 10 of Change Log | B08
Testing of Patient Sample or Reference Material Using Reference Method | B09 | The investigation employed relevant empirical testing of a patient sample or reference material using an appropriate reference method to the device (usually an IVD) involved in the reported adverse event in order to support the identification of possible causes for the adverse event. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. |  |  |  | B09
Testing of Patient Sample or Reference Material Using Competitor's Device | B10 | The investigation employed relevant empirical testing of a patient sample or reference material using a competitor's device that is comparable to the device (usually an IVD) involved in the reported adverse event in order to support the identification of possible causes for the adverse event. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. |  |  |  | B10
Historical Data Analysis | B11 | The investigation involved the analysis of historical adverse events data of the actual device involved in the adverse event and/or of products from the same and/or different batches/lots. |  |  |  | B11
Trend Analysis | B12 | The investigation involved trend analysis of adverse event of the actual device involved in the adverse event and/or of products from the same and/or different batches/lots. It should be noted that trend analysis typically is not considered sufficient as a stand-alone method, but should be used in conjunction with other investigation methods for providing for instance complementary information. |  | Retired (2026) | Term retired for Release Number 2026. Use B11 Trend Analysis. See no. 2 of Change Log | B12
Communication/Interviews | B13 | The investigation involved communication/interviews (either interpersonal or through technical means, e.g. phone, e-mail) with persons close to the adverse event, e.g. healthcare professionals (doctors, nurses etc.), the affected patient(s) or other users including, where appropriate, relatives or others engaged in caring for the affected patient. |  | Retired (2026) | Term retired for Release Number 2026. Use B15 Analysis of Information Provided by User/Third Party'. See no. 1 of Change Log | B13
Analysis of Production Records | B14 | The investigation involved the analysis of relevant production records in view of supporting the identification of possible causes for the adverse event. |  |  |  | B14
Analysis of Information Provided by User/Third Party | B15 | The investigation involved the analysis of relevant information provided by the user (e.g. healthcare professional, patient, clinical engineer) or a third party (e.g. testing facility) in view of supporting the identification of possible causes for the adverse event. This does not include information recorded by the device. |  |  | Term/definition was modified on 31 January 2023 (Old term/definition: Analysis of Data Provided by User/Third Party: The investigation involved the analysis of relevant data provided by the user (e.g. healthcare professional, patient, clinical engineer) or a third party (e.g. testing facility) in view of supporting the identification of possible causes for the adverse event.). For details, refer to comment No. 41 of the Change Log (Release Number 2023). | B15
Device not Manufactured by Reporting Manufacturer | B16 | Further information was obtained which established that the manufacturer of the device involved was not the one to which it was initially attributed. This includes mistaken identity or falsified device. Use in conjunction with C20 and D13 or D14. |  | Modified (editorial) | Term was modified for editorial consistency for Release Number 2026 See comment No. 232 of the Change Log. | B16
Device not Returned | B17 | The actual device involved in the adverse event was not returned for testing despite requests by manufacturer. |  | Modified (editorial) | Term was modified for editorial consistency for Release Number 2026 See comment No. 233 of the Change Log. | B17
Device Discarded | B18 | The actual device involved in the adverse event had been already discarded and thus irretrievably lost for testing. |  |  |  | B18
Incomplete Device Returned | B19 | The device was returned incompletely, lacking parts, components or accessories that would be required for appropriate testing and analysis of root causes. |  |  |  | B19
Device not Accessible for Testing | B20 | The actual device involved in the adverse event is not readily accessible for testing (e.g. remains implanted in patient). |  | Modified (editorial) | Term was modified for editorial consistency for Release Number 2026 See comment No. 234 of the Change Log. | B20
Type of Investigation not Yet Determined | B21 | Details to determine the type of investigation are not yet available, but are being sought. Do not use this code if the investigation is complete. |  | Modified (editorial) | Term was modified for editorial consistency for Release Number 2026 See comment No. 235 of the Change Log. | B21
Insufficient Information Available | B22 | The information available relating to the reported event is not sufficient to identify either the manufacturer, the device, or other essential information. This term indicates that no further investigation is possible. Do not use this code if further information is being sought, instead use ''Type of investigation not yet determined''. |  |  |  | B22
Specimen Requested But not Provided | B23 | A patient specimen was required to adequately investigate the issue and was requested but not provided. |  | Modified (editorial) | Term was modified for editorial consistency for Release Number 2026 See comment No. 236 of the Change Log. | B23
Event History Log Review | B24 | The investigation involved the analysis of the history log files retrieved from the device to support possible causes for the adverse event, includes device interrogation. |  |  | Definition was modified on 31 January 2023 (Old definition: The investigation involved the analysis of the history log files retrieved from the device to support possible causes for the adverse event.). For details, refer to comment No. 41 of the Change Log (Release Number 2023). | B24
Particulate Testing | B25 | Analysis of particulate or foreign material returned from user for evaluation. |  |  | Term was added on 31 January 2023. For details, refer to comment No. 42 of the Change Log (Release Number 2023). | B25
Testing of Patient Sample Using Manufacturer's Device | B26 | The investigation employed relevant empirical testing of a patient sample using the device (usually an IVD) involved in the reported adverse event or an identical device. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. |  |  | Term added for Release Number 2025. See no. 9 of Change Request Log | B26
Testing of Reference Material Using Manufacturer's Device | B27 | The investigation employed relevant empirical testing of a reference material using the device (usually an IVD) involved in the reported adverse event or an identical device. Relevant testing would typically be based on test methods used for evaluating safety and performance as described in the latest relevant standards. |  |  | Term added for Release Number 2025. See no. 10 of Change Request Log | B27
Device Lost in Transit | B28 | The actual device involved in the adverse event has been returned by the user but not received for testing. |  |  |  | B28
Device Data not Available for Required Analysis | B29 | Device data such as software logs are not available for analysis by the manufacturer |  |  |  | B29
Analysis of Images | B30 | The investigation involved review of images of the device, and/or procedural or diagnostic images. |  |  | Term added for Release Number 2025. See no. 36 of Change Request Log | B30
Labeling Review | B31 | A review of the reported information against the product labeling, including instructions for use was conducted to determine if the complaint resulted from inadequacies in the product labeling or a failure by the user to adhere to the labeling. Note: labeling review alone is not sufficient as an investigation for the incident. |  | Modified (editorial) | Term was modified for editorial consistency for Release Number 2026 See comment No. 182 of the Change Log. | B31
Analysis of Service/Maintenance Records | B32 | The investigation involved the analysis of the service history or maintenance records for the device involved. |  |  | Term added for Release Number 2025. See no. 89 of Change Request Log | B32
Incorrect Manufacturer Attribution | B33 | Further information was obtained which established that the manufacturer of the device involved was not the one to which it was initially attributed. This includes mistaken identity or falsified device. Use in conjunction with C20 and D13 or D14. |  | New | Term added for Release Number 2026. See no. 11 of Change Request Log | B33
Software Code Analysis | B34 | systematic review and testing of source code, through manual and automated methods, to detect errors, vulnerabilities, and deviations from requirements |  | New | Term added for Release Number 2026. See no. 37 of Change Request Log | B34

## C
Annex Name: Annex C
Annex Title: Cause Investigation - Investigation Findings
Release Number: 2026
Annex Description: Terms/codes for describing the findings in the specific investigation that are the keys to identify the root cause. This annex has hierarchical levels, allowing jurisdictions to choose the level of coding to use.
Annex Instructions: Terms/Codes in this annex should be used to answer the question What were the results of the investigation? IMDRF Terms and Definitions use American spelling. Wherever appropriate the word 'device' should be interpreted as meaning all or any of the following: medical device, in-vitro diagnostic medical device, accessory to a medical device, medical device parts and/or components. Wherever appropriate the word 'patient' should be interpreted as meaning all or any of the following: user, operator, any other person affected by the incident.
Annex Approval Date: January 2026
NOTE: Maximum field length Term: 100 characters Definition: 700 Status description: 700
Level 1 Term | Level 2 Term | Level 3 Term | Code | Definition | Non-IMDRF Code/Term | Status | Status Description | CodeHierarchy
Biological Problem Identified |  |  | C01 | Problems relating to, caused by or affecting biological processes or living organisms. |  |  |  | C01
| Biocompatibility Problem Identified |  | C0101 | The device causes cellular or tissue responses that elicit an undesirable local or systemic effect in the recipient or beneficiary of that therapy. (See ISO 10993) |  |  |  | C01¦C0101
| Biological Contamination |  | C0102 | The undesirable presence of living organisms such as bacteria, fungi, or viruses or their products (enzymes or toxins). |  |  |  | C01¦C0102
|  | Endotoxin Contamination | C010201 | The undesirable presence of toxins associated with certain bacteria (e.g. gram negative bacteria). |  |  |  | C01¦C0102¦C010201
|  | Microbial Contamination | C010202 | The undesirable presence of microorganisms or microbes such as bacteria and fungi (yeasts and molds). |  |  |  | C01¦C0102¦C010202
| Material or Material Leachate Pyrogenic Problem |  | C0103 | The undesirable presence of pyrogens or fever-producing organisms caused by materials that permeate through the device. |  |  |  | C01¦C0103
| Cytotoxicity Problem Identified |  | C0104 | The device was found to have an undesirable level of toxicity to living cells. |  |  |  | C01¦C0104
| Genotoxicity Problem Identified |  | C0105 | The device's ability to cause damage to genetic material (e.g. leading to malignant tumors). (See ISO 10993) |  |  |  | C01¦C0105
|  | Carcinogenic Problem | C010501 | The device's ability to trigger development of cancer. |  |  |  | C01¦C0105¦C010501
|  | Mutagenic Problem | C010502 | The device's ability to change genetic information (usually DNA) of an organism and thus increasing the frequency of mutations. |  |  |  | C01¦C0105¦C010502
| Hematological Problem Identified |  | C0106 | The device affects or impacts the blood or its components. (See ISO 10993 all parts) |  |  |  | C01¦C0106
|  | Agglutination Problem | C010601 | The device affects the ability of the blood to clot which may be induced by chemical, mechanical, or thermal properties of the device. |  |  |  | C01¦C0106¦C010601
|  | Complement Activation Problem | C010602 | The device affects the body's ability to activate the complement system of the immune system, thereby interfering with the ability to clear pathogens. This may be caused by an interaction of the device with chemicals or materials. |  |  |  | C01¦C0106¦C010602
|  | Platelet Activation Problem | C010603 | The device affects the body's ability to activate platelet formation. |  |  |  | C01¦C0106¦C010603
|  | Problem Due to Thrombosis Activation | C010604 | The device causes the formation of blood clots in or along blood vessels resulting in disturbed or disrupted blood flow. |  | Modified (Editorial) | Term modified for editorial consistency in Release Number 2026. See no. 239 of Change Request Log | C01¦C0106¦C010604
| Unintended Presence of Allergens |  | C0107 | Unintended or unexpected presence of allergens in the device. If the presence of the allergen is expected but not adequately labeled, then use C05 or child terms. |  | Modified (Editorial) | Definition modified for editorial consistency. Release Number 2026. See no. 183 and 195 of Change Request Log | C01¦C0107
| Reproductive Toxicity Problem Identified |  | C0108 | The device affects reproductive function, embryo development (teratogenicity), and prenatal and early postnatal development. (ISO 10993 part 3) |  |  |  | C01¦C0108
| Undesirable Presence of Endogenous Materials |  | C0109 | Undesirable presence of endogenous substances, body fluids, cellular tissues, etc. |  |  | Term was added on 31 January 2023. For details, see comment No. 61 and No. 64 of the Change Log (Release Number 2023). | C01¦C0109
Electrical Problem Identified |  |  | C02 | Events associated with an electrically powered device where an electrical malfunction results in a device problem (e.g. electrical circuitry, contact or component failed) even if the problem is intermittent. |  |  |  | C02
| Electrical/Electronic Component Problem Identified |  | C0201 | The performance of an electrical or electronic component was found to be inadequate. |  |  |  | C02¦C0201
| Hardware Timing Problem Identified |  | C0202 | Problems that results from improper sequential activation of components. |  |  |  | C02¦C0202
| Impedance Problem Identified |  | C0203 | Problems due to insufficient or excessive resistance to current flow either by the device or circuit. |  |  |  | C02¦C0203
| Insulation Problem Identified |  | C0204 | Problems due to inadequate or incorrect electrical insulation material. |  |  |  | C02¦C0204
| Open Circuit |  | C0205 | Problem due to an electrical circuit that does not conduct current because a switch is open, a wire is broken, etc. |  |  |  | C02¦C0205
| Current Leakage Problem |  | C0206 | Problems related to leakage currents which may cause electric shock. These currents usually flow through the protective ground conductor. In its absence, these currents could flow from the device to the ground via the human body. |  |  |  | C02¦C0206
| Power Source Problem Identified |  | C0207 | Problems related to the source that provides electrical power to the device. |  |  |  | C02¦C0207
|  | Energy Storage System Problem | C020701 | Problems related to the energy storage system (e.g. the rechargeable battery, charging system, or capacitor) and includes problems such as premature power source depletion and battery explosions. |  |  |  | C02¦C0207¦C020701
|  | Loss of Power Identified | C020702 | A device that experienced problems due to a loss in the power supply. |  |  | Term modified for Release Number 2025. Previous term was "Loss of Power". See no. 126 of Change Request Log. | C02¦C0207¦C020702
|  | Power Fluctuation | C020703 | The device failed due to fluctuations within the power supply (e.g. transient power, power spike, power dip, or power sequencing). |  |  |  | C02¦C0207¦C020703
| Short Circuit |  | C0208 | Problems due to an unintentionally low-resistance connection between two points in an electric circuit, resulting in either excessive current flow that often causes damage or in a new shorter circuit that draws current away from the original pathways and components. |  |  |  | C02¦C0208
| Signal Loss |  | C0209 | Problems due to the loss or weakening of an electrical signal or signals. |  |  |  | C02¦C0209
| Corrupted Memory |  | C0210 | Malfunction that resulted from a device memory corruption, including bit flips or single event errors. |  | Modified (Editorial) | Term modified for editorial consistency in Release Number 2026. See no. 240 of Change Request Log | C02¦C0210
Electromagnetic Compatibility Problem Identified |  |  | C03 | Device-to-device or device-environment problem resulting from electromagnetic disturbances. |  |  |  | C03
| Conducted Interference |  | C0301 | Problems related to electromagnetic interference (EMI) by physical contact with conductors (e.g. wires, resistors, terminals) as opposed to radiated EMI which is caused by induction (without physical contact of the conductors). |  |  |  | C03¦C0301
| Electrostatic Discharge Identified |  | C0302 | Problems due to sudden and momentary bursts of electrical current flowing between two objects at different electrical potentials. |  |  | Term modified for Release Number 2025. Previous term was "Electrostatic Discharge". See no. 140 of Change Request Log. | C03¦C0302
| Inadequate Immunity |  | C0303 | Problems related to immunity or capabilities to resist electromagnetic interference (EMI). |  |  |  | C03¦C0303
| Unintended Emission |  | C0304 | Problems due to unintended emission of electromagnetic energy by the device. |  |  |  | C03¦C0304
| Radiofrequency Interference (RFI) Identified |  | C0305 | Problems due to radiofrequency interference. RFI is a disturbance that affects an electrical circuit due to either electromagnetic conduction or electromagnetic radiation emitted from an external source. |  |  | Term modified for Release Number 2025. Previous term was "Radiofrequency Interference (RFI)". See no. 132 of Change Request Log. | C03¦C0305
Interoperability Problem Identified |  |  | C04 | Problems with the mechanical, electrical, or communication interface between two or more separate devices. |  |  |  | C04
| Communications Problem Identified |  | C0401 | Devices that do not send or receive adequate signals (this speaks to the interoperability between devices). |  |  |  | C04¦C0401
|  | Wired Communication Problem | C040101 | Communications problems between devices within a wired system. |  |  |  | C04¦C0401¦C040101
|  | Wireless Communication Problem Identified | C040102 | Communications problems between devices within a wireless system. |  |  | Term modified for Release Number 2025. Previous term was "Wireless Communication Problem". See no. 138 of Change Request Log. | C04¦C0401¦C040102
|  | Network Communication Problem | C040103 | Communications problems between devices within a network system. |  |  |  | C04¦C0401¦C040103
| Incompatible Component/ Accessory |  | C0402 | A device that malfunctions due to a component(s)/accessory that does not operate correctly and according to the device's specifications. |  |  |  | C04¦C0402
| Device not Compatible With Another Device |  | C0403 | A device that malfunctions due to being used in combination with, or in the presence of, another device. |  | Modified (editorial) | Term modified for editorial consistency in Release Number 2026. See no. 241 of Change Request Log | C04¦C0403
| Unintended Compatibility Identified |  | C0404 | The device was confirmed to be compatible with another device with which the device is intended to be incompatible. |  |  | Term modified for Release Number 2025. Previous term was "Unintended Compatibility". See no. 136 of Change Request Log. | C04¦C0404
Labeling and Instructions for Use/Maintenance |  |  | C05 | Insufficient, inadequate, or incorrect information provided on a device's label or documentation regarding e.g. its intended use, directions for use, and characteristics of the device, including its maintenance. |  |  |  | C05
| Inadequate Labeling and/or Instructions for Use |  | C0501 | Inadequate information on the labels or in the instructions for use e.g. steps that are difficult to follow or that are missing. |  | Modified (Editorial) | Term modified for editorial consistency in Release Number 2026. See no. 184 of Change Request Log | C05¦C0501
| Incorrect Labeling and/or Instructions for Use |  | C0502 | Missing, incorrect, or inappropriate information on the labels e.g. mislabeled contents or device labeling characteristics or package contents. |  |  |  | C05¦C0502
| Inadequate or Incorrect Instructions for Maintenance |  | C0503 | Inadequate or incorrect information in the instructions for maintenance. |  |  |  | C05¦C0503
Material and/or Chemical Problem Identified |  |  | C06 | Problems with the device materials or how its materials react to other elements either within the device or within the environment. |  |  |  | C06
| Degradation Problem Identified |  | C0601 | Problems that occur when the device becomes worn, weakened, corroded, or broken down due to processes such as aging, permeation, and corrosion. |  |  |  | C06¦C0601
| Inappropriate Material |  | C0602 | Problems that occur due to undesired presence, composition or structure of a material. |  | Modified (technical) | Definition modified for Release Number 2026. See no. 136 of Change Request Log. | C06¦C0602
|  | Improper Composition/ Concentration | C060201 | Problems associated with the improper combination of materials or elements present in the device (e.g. improper composition of the materials of a capacitor). |  |  |  | C06¦C0602¦C060201
|  | Improper Physical Structure | C060202 | Problems related to the incorrect or inadequate arrangement of the parts, components, elements, or materials. |  |  |  | C06¦C0602¦C060202
|  | Molecular Structure Problem | C060203 | Problems related to the presence of an inappropriate molecular geometry somewhere in the device (i.e. the spatial arrangement of atoms in a molecule and the chemical bonds that hold the atoms together). |  |  |  | C06¦C0602¦C060203
| Inadequate Physicochemical Properties |  | C0603 | Problems that occur due to the physicochemical properties. |  |  |  | C06¦C0603
| Incompatible Material |  | C0604 | Problems that occur due to the incompatibility of materials that co-exist simultaneously as part of the device. |  |  |  | C06¦C0604
| Reactivity Problem Identified |  | C0605 | Problems that occur due to the reactivity of materials (e.g. over-react or under-react). |  |  |  | C06¦C0605
| Tolerance Stack-Up |  | C0606 | Problems that result from a combination of specification variances of the components. |  |  |  | C06¦C0606
Mechanical Problem Identified |  |  | C07 | Problems that result from internal or external forces including fluids, other objects, or environmental or physiologic influences. |  |  |  | C07
| Device Migration |  | C0701 | A device that has moved from its original location due to external forces (e.g. stent or lead movement). |  |  |  | C07¦C0701
| Friction Problem Identified |  | C0702 | Problems caused by its surface coming in contact with another surface or fluid. |  |  |  | C07¦C0702
| Leakage/Seal |  | C0703 | Problems caused by inadequate/broken seal within the device. |  |  |  | C07¦C0703
| Lubrication Problem Identified |  | C0704 | Problems that occurred because of the presence of either too much or too little lubricant where required (e.g. connectors, leading to failure mechanisms such as corrosion). |  |  |  | C07¦C0704
| Stiffness Problem Identified |  | C0705 | Problems that occurred when its material is either too flexible/pliable or inflexible/rigid when in contact by an applied force. |  |  |  | C07¦C0705
| Stress Problem Identified |  | C0706 | Problems caused by either excessive or inadequate physical force exerted on it by another object resulting in problems e.g. wear, bending, deformation, fracture, fatigue. |  |  |  | C07¦C0706
|  | Deformation/Distortion Problem | C070601 | Problems caused by changes in the shape or size of the device due to an applied force. This can be a result of tensile forces (pulling), compressive forces, shear, bending, knotting, or torsion. |  |  | Term and Definition modified for Release Number 2025. Previous term was "Deformation Problem". See no. 13 of Change Request Log. | C07¦C0706¦C070601
|  | Fatigue Problem | C070602 | Problems due to the weakening or breakdown of its material when subjected to stress or a series of repeated stresses. |  |  |  | C07¦C0706¦C070602
|  | Separation Problem | C070603 | Problems caused by the cut, tear, or fracture of a component, object, or material into two or more pieces including shear. |  |  | Term and Definition modified for Release Number 2025. Previous term was "Fracture Problem". See no. 15 of Change Request Log. | C07¦C0706¦C070603
|  | Mechanical Shock Problem | C070604 | Problems caused by the sudden violent blow or collision to the whole device (e.g. by dropping). |  |  |  | C07¦C0706¦C070604
|  | Vibration Problem | C070605 | Problems caused by the constant rhythmic motion of the device, or something in the environment to which the device is exposed. |  |  |  | C07¦C0706¦C070605
|  | Wear Problem | C070606 | Problems due to the premature or expected erosion of its material by use, deterioration, or change. |  |  |  | C07¦C0706¦C070606
| Incorrect Dimension |  | C0707 | Problems caused by incorrect physical dimensions of the device or one of its parts |  |  |  | C07¦C0707
| Blockage Identified |  | C0708 | Problems that occurred due to an obstruction or blockage. |  |  | Term was added on 31 January 2023. For details, see comment No. 45 of the Change Log (Release Number 2023). | C07¦C0708
Optical Problem Identified |  |  | C08 | Problems related to the optical properties of a device. |  |  |  | C08
| Optical Transmission Problem Identified |  | C0801 | Problems with the device's ability to pass light energy. |  |  |  | C08¦C0801
| Light Source Problem Identified |  | C0802 | Problems with the optical properties of a device such as diopter, glare, and irradiance or glistening. |  |  |  | C08¦C0802
Clinical Imaging Problem Identified |  |  | C09 | Problems that occur with devices used for radiographic or imaging procedures e.g. CT scanners, magnetic resonance imaging. |  |  |  | C09
| Gradient Induced Field Problem |  | C0901 | Problems that result from the gradient-induced fields generated during radiologic procedures e.g. magnetic resonance imaging. |  |  |  | C09¦C0901
| Image Artifact |  | C0902 | The unacceptable distortion of an image due to signal loss that may occur during a radiologic procedure such as magnetic resonance imaging. |  |  |  | C09¦C0902
| Magnetically-Induced Movement |  | C0903 | Problems due to unintended or excessive movement created by the application of magnetic fields. |  |  |  | C09¦C0903
| Radiofrequency Induced Overheating |  | C0904 | Problems due to unintended radiofrequency-induced temperature increase that can occur in the vicinity of the device. |  |  |  | C09¦C0904
Software Problem Identified |  |  | C10 | Problems related to the device software. |  |  |  | C10
| Configuration Issue |  | C1001 | Problems due to change control or incorrect version, including regional requirements. |  |  |  | C10¦C1001
| Design Error |  | C1002 | The device had faulty (incomplete or incorrect) software design. |  |  |  | C10¦C1002
|  | Data Compression Error | C100201 | Data was lost or corrupted during the operation of reducing storage space or communication bandwidth. |  |  |  | C10¦C1002¦C100201
|  | Incorrect Algorithm | C100202 | The device software was found to implement an incorrect sequence of steps for a specific computation. |  |  |  | C10¦C1002¦C100202
|  | Incorrect Data Definition | C100203 | The device software was found to contain errors in specifying or manipulating data items. |  |  |  | C10¦C1002¦C100203
|  | Interface Design Error | C100204 | The device software was found to contain errors in the user interface (including usability problems) or the interfaces with other systems. |  |  |  | C10¦C1002¦C100204
|  | Non-Functional Defect | C100205 | The device software contained software errors that did not impact its operation. |  |  |  | C10¦C1002¦C100205
|  | Software Timing Problem | C100206 | Problems that results from the incorrect sequencing or activation of software modules. |  |  |  | C10¦C1002¦C100206
| Software Maintenance Problem Identified |  | C1003 | The device software was not maintained/updated properly. |  |  |  | C10¦C1003
| Software Installation Problem Identified |  | C1004 | The device software was not installed as per the specifications or failed to properly install. |  |  |  | C10¦C1004
| Software Requirement Error |  | C1005 | The software requirements for the device are either incomplete, inadequate, or in conflict. |  |  |  | C10¦C1005
| Software Runtime Error |  | C1006 | The device software failed during operation as a result of a coding error. |  |  |  | C10¦C1006
| Software Security Vulnerability |  | C1007 | The device software failed to provide adequate authorization, access control, protection and accountability features. |  |  |  | C10¦C1007
| Erroneous Data Transfer |  | C1008 | The device software fails to transfer the expected data within a system or to another device. |  |  |  | C10¦C1008
| Data Storage or Loss of Data |  | C1009 | Storage of data was unsuccessful in total or in part. |  |  |  | C10¦C1009
| Artificial Intelligence Training/Validation Problem Identified |  | C1010 | Problem associated with the training and/or validation of artificial intelligence, including machine learning algorithms. |  |  | Term was added on 27 January 2022. For details, see comment No. 11 of the Change Log (Release Number 2022). | C10¦C1010
| Corrupt Software |  | C1011 | The device software becomes corrupted and results in device malfunction. |  |  | Term was added on 27 January 2022. For details, see comment No. 30 of the Change Log (Release Number 2022). | C10¦C1011
Thermal Problem |  |  | C11 | Problems related to the temperature of the device. Note: For problems related to environmental temperature use ''Environment Problem Identified''. |  |  |  | C11
| Overheating Problem Identified |  | C1101 | The device was found to become hotter than expected during operation. This applies to devices which are not intended to deliver heat. Use "Excessive heating identified" for devices which are intended to deliver heat during operation. Use "Inadequate cooling identified" if the overheating was related to a problem with a cooling system. |  |  |  | C11¦C1101
| Excessive Heating Identified |  | C1102 | The device delivered more heat than intended or expected during operation. This applies to devices which are intended to deliver heat. Use "Overheating problem identified" for devices which are not intended to deliver heat during operation. |  |  |  | C11¦C1102
| Excessive Cooling Identified |  | C1103 | The device cooled the patient or another device more than intended or expected during operation. |  |  |  | C11¦C1103
| Inadequate Cooling Identified |  | C1104 | The device did not sufficiently cool the patient or another device during operation. |  |  |  | C11¦C1104
Protective System Problem Identified |  |  | C12 | Problems related to the system(s) designed to prevent or warn about unsafe operation of the device. |  |  |  | C12
| Fail-Safe Problem Identified |  | C1201 | A system intended to prevent unsafe operation of the device did not operate correctly. |  | Modified (editorial) | Term modified for editorial consistency in Release Number 2026. See no. 242 of Change Request Log | C12¦C1201
| Alarm System Problem Identified |  | C1202 | A system intended to warn of a potentially unsafe condition did not operate correctly. |  |  |  | C12¦C1202
| Problem of Device to Self-Test |  | C1203 | Malfunction of the device's self-test system. |  |  |  | C12¦C1203
| Problem to Auto Stop |  | C1204 | An auto stop function of a device did not operate correctly. |  |  |  | C12¦C1204
| Premature Indicator Activation Identified |  | C1205 | A system intended to indicate the device status was triggered prematurely. |  |  | Term modified for Release Number 2025. Previous term was "Premature Indicator Activation". See no. 130 of Change Request Log. | C12¦C1205
| Reset Problem Identified |  | C1206 | The device does not reset properly. |  |  | Term modified for Release Number 2025. Previous term was "Reset Problem". See no. 134 of Change Request Log. | C12¦C1206
| Shielding Problem |  | C1207 | Inadequate shielding of/by the device. |  |  |  | C12¦C1207
| Missing or Inadequate Safety Measures Identified |  | C1208 | Safety measures are inadequately applied or missing. |  |  | Term modified for Release Number 2025. Previous term was "Missing or Inadequate Safety Measures". See no. 127 of Change Request Log. | C12¦C1208
Operational Problem Identified |  |  | C13 | Problems that occur during the performance, use, or functioning of the device. |  |  |  | C13
| Device Incorrectly Reprocessed |  | C1301 | Problems associated with the failure to properly and adequately reprocess the device. |  |  |  | C13¦C1301
|  | Device Incorrectly Cleaned During Reprocessing | C130101 | The cleaning procedure is not followed correctly or used inappropriate cleaning materials. |  |  |  | C13¦C1301¦C130101
|  | Device Incorrectly Disinfected/Sterilised During Reprocessing | C130102 | The disinfection/sterilization process was incorrect and/or the wrong products for disinfection/sterilization were used. |  |  |  | C13¦C1301¦C130102
|  | Device Incorrectly Assembled During Reprocessing | C130103 | Incorrect assembly of the device following reprocessing. |  |  |  | C13¦C1301¦C130103
| Failure to Calibrate or Used Out of Calibration |  | C1302 | A device that cannot calibrate (establish the relationship between a measuring device and the units of measure) or is used out of calibration to ensure accurate readings. |  |  | Term/definition was modified on 31 January 2023 (Old term/definition: Failure to Calibrate: A device that cannot calibrate (establish the relationship between a measuring device and the units of measure) to ensure accurate readings). For details, see comment No. 58 of the Change Log (Release Number 2023). | C13¦C1302
| Device Difficult to Operate |  | C1303 | Problems including set-up, operation, and disassembly of equipment. Not including reprocessing. |  |  |  | C13¦C1303
| Incorrect Interpretation of Results/Data |  | C1304 | Problems resulting from the incorrect interpretation by the user of the results or data provided by the device. |  |  |  | C13¦C1304
Patient Sample Problem |  |  | C14 | Problems that occurred due to endogenous or exogenous interferent in the sample, or unexpected variation in the target analyte/marker. |  |  |  | C14
| New or Unknown Interferent |  | C1401 | New or unknown endogenous or exogenous interferent (sample) identified. |  |  |  | C14¦C1401
| Known Interferent |  | C1402 | Known interferent in the sample identified. |  |  |  | C14¦C1402
| Problem Related to Variant/Mutant |  | C1403 | Device performance affected by variant/mutant which is not covered in the labeling. |  | Modified (Editorial) | Definition modified for editorial consistency in Release Number 2026. See no. 185 of Change Request Log | C14¦C1403
| Pre-Analytical Handling Problem |  | C1404 | Incorrect pre-analytical handling of patient's sample by the user. |  | Modified (Editorial) | Term modified for editorial consistency in Release Number 2026. See no. 243 of Change Request Log | C14¦C1404
Environment Problem Identified |  |  | C15 | Problems that occurred due to factors within the environment e.g. dust, dirt, humidity, temperature. |  |  |  | C15
| Environmental Temperature Problem Identified |  | C1501 | Device performance was affected by the temperature, or changes in temperature, of the environment in which it was used. |  |  |  | C15¦C1501
| Contamination of Device by Foreign Material from Environment |  | C1502 | A device that experienced problems due to ingress or coating of dust, dirt or other foreign material |  |  | Term and Definition modified for Release Number 2025. Previous term was "Dust or Dirt Problem Identified". See no. 14 of Change Request Log. | C15¦C1502
| Contamination of Environment by Device |  | C1503 | Operation of the device results in contamination of the nearby environment e.g. dust, dirt, smoke, heat or biological material. |  |  |  | C15¦C1503
| Environmental Pressure Problem Identified |  | C1504 | Device performance was affected by the pressure, or changes in pressure, of the environment in which it was used. |  |  |  | C15¦C1504
| Ambient Light Problem Identified |  | C1505 | Device performance was affected by ambient light. This term applies to the direct effects of ambient light on the device, and to the user's ability to operate the device (e.g. to read device output). |  |  |  | C15¦C1505
| Environmental Humidity Problem Identified |  | C1506 | Device performance was affected by the humidity, or changes in humidity, of the environment in which it was used. |  |  |  | C15¦C1506
Manufacturing Process Problem Identified |  |  | C16 | Problems with a device that can be traced to a problem in the manufacturing and/or production process. |  |  |  | C16
| Assembly Problem Identified |  | C1601 | Problems that occurred because the device was assembled incorrectly. |  |  |  | C16¦C1601
| Sterilization Problem Identified |  | C1602 | Problems that occurred during terminal sterilization by the manufacturer. |  |  |  | C16¦C1602
| Installation Problem Identified |  | C1603 | A device that malfunctions because it was incorrectly installed, set-up, or configured (e.g. misconfiguration of an "automatic" defibrillator to "semi-automatic", thereby leading to failure). |  |  |  | C16¦C1603
| Maintenance of Manufacturing Machinery |  | C1604 | Problems caused by failure to maintain manufacturing equipment used to produce the device. |  |  |  | C16¦C1604
| Packaging Problem Identified |  | C1605 | Problems that occurred because of the device packaging. |  |  |  | C16¦C1605
|  | Packaging Compromised | C160501 | Problems that occurred because of a compromised packaging of the device (e.g. broken or incomplete seal). |  |  |  | C16¦C1605¦C160501
|  | Packaging Materials Problem | C160502 | Problems that occurred because of the composition or type of packaging materials was inappropriate for the device. |  |  |  | C16¦C1605¦C160502
|  | Packaging Contains Unintended Material | C160503 | Problems that occurred because unintended material was packaged with the device. |  |  |  | C16¦C1605¦C160503
|  | Packaging Contains Incorrect or Incomplete Device | C160504 | Problems that occurred because the packaging contained an incorrect or incomplete device (missing components). |  |  | Term/definition was modified on 31 January 2023 (Old term/definition: Packaging Contains Incorrect Device: Problems that occurred because the packaging contained an incorrect device.). For details, see comment No. 51 of the Change Log (Release Number 2023). | C16¦C1605¦C160504
Maintenance Problem Identified |  |  | C17 | A device malfunction or problem that occurs after production because the device was not properly maintained according to the instructions (e.g. maintenance may be performed by user facility, distributor, or service provider). |  |  |  | C17
| Misadjustment/Misalignment Identified |  | C1701 | Problems caused by incorrect adjustment of device components. |  | Modified (Editorial) | Term modified for editorial consistency in Release Number 2026. See no. 244 of Change Request Log | C17¦C1701
Transport/Storage Problem Identified |  |  | C18 | Problems was caused by transport or storage conditions. |  |  |  | C18
| Transport Problem Identified |  | C1801 | Problems traced to how the device was transported e.g. temperature of shipping compartment or method of transportation. |  |  |  | C18¦C1801
| Storage Problem Identified |  | C1802 | Problems that result from storing the device in an uncontrolled or improper environment (e.g. moisture sensitive devices stored in a humid environment). |  |  |  | C18¦C1802
No Device Problem Found |  |  | C19 | Reported problem cannot be reproduced during investigation. |  |  | Definition of this term was clarified on 21 January 2021. For details, refer to comment No. 17 and 29 of the Change Log. (Old definition: "The device either functioned as intended or a problem was not found.") | C19
| Unable to Exclude Device Problem |  | C1901 | The investigation findings do not clearly confirm the presence or absence of the reported problem with the device. |  | New | Level 2 D terms moved to C19 for Release Number 2026. See No. 3 of Change request Log. Minor editorial change to term for Release Number 2026. See no. 246 of Change Request Log. | C19¦C1901
| Device Problem Excluded |  | C1902 | The investigation findings clearly confirm that there was no problem with the device. |  | New | Level 2 D terms moved to C19 for Release Number 2026. See No. 4 of Change request Log. Minor editorial change to term for Release Number 2026. See no. 246 of Change Request Log. | C19¦C1902
No Findings Available |  |  | C20 | Use when no investigation is performed or when the investigation is insufficient to reasonably exclude a device malfunction with adequate certainty. This code should typically be used in conjunction with B16 and D13. |  | Modified (Editorial) | Definition modified for Release Number 2026. See no. 10 of Change Request Log | C20
Results Pending Completion of Investigation |  |  | C21 | Investigation is ongoing and results are not yet available. Do not use this code if the investigation is complete. |  |  |  | C21
Appropriate Investigation Findings Term/Code not Available |  |  | C22 | Problems is not adequately described by any other term. Note: This code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. |  | Modified (editorial) | Term modified for editorial consistency in Release Number 2026. See no. 245 of Change Request Log | C22
Usage Problem Identified |  |  | C23 | Problems that occurred related to the actions of a healthcare professional, patient, or other device user. |  |  | Term was added on 21 January 2021 for coding device usage problems. For details, refer to comment No. 32 of the Change Log. | C23
| Insufficient Sample Volume |  | C2301 | Not having a sufficient volume of sample to perform a test. |  |  | Term was added on 31 January 2023. For details, see comment No. 53 of the Change Log (Release Number 2023). | C23¦C2301
| Use of Non-Validated Controls Identified |  | C2302 | Problems that occur related to the use of non-validated controls. |  |  | Term was added on 31 January 2023. For details, see comment No. 57 of the Change Log (Release Number 2023). | C23¦C2302
Malfunction Observed Without Conclusive Finding |  |  | C24 | Malfunction was verified but no conclusive finding is available. |  |  | Term was added on 31 January 2023. For details, see comment No. 62 of the Change Log (Release Number 2023). | C24

## D
Annex Name: Annex D
Annex Title: Cause Investigation – Investigation Conclusion
Release Number: 2026
Annex Description: Terms/codes for describing the conclusion of the device involved in the reported event.
Annex Instructions: Terms/Codes in this annex should be used to answer the question What is the conclusion regarding the root cause? (based on the results noted in the Investigation Findings and the way they were obtained noted in the Type of Investigation.). IMDRF Terms and Definitions use American spelling. Wherever appropriate the word 'device' should be interpreted as meaning all or any of the following: medical device, in-vitro diagnostic medical device, accessory to a medical device, medical device parts and/or components. Wherever appropriate the word 'patient' should be interpreted as meaning all or any of the following: user, operator, any other person affected by the incident.
Annex Approval Date: January 2026
NOTE: Maximum field length Term: 100 characters Definition: 700 Status description: 700
Level 1 Term | Level 2 Term | Code | Definition | Non-IMDRF Code | Status | Status Description | CodeHierarchy
Cause Traced to Device Design |  | D01 | Problems traced to the design specifications (e.g. in the requirements, testing processes, hazard analysis, implementation strategy). |  |  |  | D01
| Design Inadequate for Purpose | D0101 | Problems traced to design/design features of the device that do not support or interfere with the intended purpose of the device. |  |  |  | D01¦D0101
| Human Factors Engineering - Device Difficult to Operate | D0102 | Problems traced to inappropriate and/or inadequate assessment and engineering design of the device to accommodate how or where the device will be used. |  |  |  | D01¦D0102
| Human Factors Engineering - Device Difficult to Assemble | D0103 | Problems traced to inadequate design of the component parts and/or assembly steps resulting in the device not being able to be assembled correctly. |  |  |  | D01¦D0103
| Human Factors Engineering - Device Difficult to Reprocess | D0104 | Problems traced to inadequate design of the reprocessing steps and/or the device resulting in the device remaining unclean. |  |  |  | D01¦D0104
| Cause Traced to Missing or Inadequate Safety Measures | D0105 | Problems traced to inadequate design or complete lack of safety measures leading to device malfunction or unintended properties of the device including possible hazards for persons using the device. |  |  | Editorial change to term for Release Number 2025. See no. 128 of Change Log. | D01¦D0105
| Design Change Validation Inadequate | D0106 | Problems traced to inadequate or lack of validation of design changes of the device leading to malfunction or unintended properties of the device including possible hazards for persons using the device. |  |  |  | D01¦D0106
| Inadequate Cybersecurity | D0107 | Problems traced to inadequate protection of computers, servers, mobile devices, electronic systems, networks, programs, and data from malicious attacks, damage, data leakage, or unauthorized access. |  | Modified (technical) | Technical change to definition for Release Number 2026. See no. 118 of Change Log. | D01¦D0107
| Inadequate Software Design | D0108 | Cause traced to inadequate or inappropriate design of the software (including AI/ML) component of a medical device, or Software as a Medical Device. |  | Modified (technical) | Technical change to definition for Release Number 2026. See no. 119 of Change Log. | D01¦D0108
Cause Traced to Component Failure |  | D02 | Expected or random component failure without any design or manufacturing issue. |  |  |  | D02
Cause Traced to Manufacturing |  | D03 | A defect in the processes or systems used in the manufacture of the device. Examples include problems within the change control, production, or quality control processes. |  |  |  | D03
| Manufacturing Deficiency | D0301 | Problems traced to manufacturing process. |  |  |  | D03¦D0301
| Quality Control Deficiency | D0302 | Problems traced to the failure to maintain or establish techniques for controlling and verifying the product specifications identified by the manufacturer. Includes inadequacy of components or raw materials supplied to the manufacturer. |  |  | Definition modified for Release Number 2024. See no. 55 of Change Log | D03¦D0302
Cause Traced to Transport/Storage |  | D04 | Problems traced to the inappropriate transport or storage of the device. |  |  |  | D04
Cause Traced to Infrastructure |  | D05 | Problems traced to underlying framework, systems, and processes at the healthcare facility or other point of use (e.g. as building power supply, network, oxygen systems). |  |  |  | D05
Cause Traced to Environment |  | D06 | Problems caused by exposure to environmental conditions outside the expected range. |  |  |  | D06
Cause Traced to Maintenance |  | D07 | Problems traced to improper routine or preventative maintenance. |  |  |  | D07
Cause Traced to Training |  | D08 | Problems caused by inadequate training. |  |  |  | D08
Cause Traced to Labeling |  | D09 | Problems that occur as the result of problems with the labeling (including package inserts, instruction manuals, instructions for use). |  |  |  | D09
Cause Traced to Non-Device Related Factors |  | D10 | The adverse event that occurred is confirmed to have been caused by something other than the device or usage. |  |  | The term and definition were modified on 21 January 2021. (Old term/definition: Cause Cannot be Traced to Device/The adverse event that occurred is not attributable to a device.). For details, refer to comment No. 28 of the Change Log. | D10
| Adverse Event Related to Patient Anatomy and/or Condition | D1001 | The patient's condition, disease or anatomy is demonstrably responsible for the adverse event and use of the device has neither caused nor otherwise influenced the adverse event. |  |  | Term and Definition modified for Release Number 2025. Previous term was "Adverse Event Related to Patient Condition". See no. 53 of Change Request Log. | D10¦D1001
| Adverse Event Related to Procedure | D1002 | The adverse event occurred during the procedure and the device had no influence on event. |  |  |  | D10¦D1002
| Adverse Event Related to Commutability | D1003 | The adverse event occurred because the material being used to calibrate or assess IVD performance did not have similar properties to those of human samples, leading to inappropriate bias and erroneous results. Examples of materials include: reference materials, calibrators, proficiency testing samples. |  |  |  | D10¦D1003
Cause Traced to User |  | D11 | The adverse event caused partially or wholly by the user of the device including sample handling. |  |  |  | D11
| Failure To Follow Instructions | D1101 | Problems traced to the user not following the manufacturer's instructions. |  |  |  | D11¦D1101
| Unintended Use Error Caused or Contributed to Event | D1102 | The interaction between the user and device, or sample, caused or contributed to the error. This includes unintended inappropriate use of the device and incorrect sample preparation. |  |  |  | D11¦D1102
| Cause Traced to Intentional Off-Label, Unapproved, or Contraindicated Use | D1103 | Problems traced to the intentional use of the device in an unapproved procedure, for an unapproved patient, or for which it is contraindicated, or not listed on the label. |  |  |  | D11¦D1103
| Shelf Life/Expiration Date Exceeded | D1104 | A device that has exceeded the period of time/date recommended by the manufacturer for storing the device without a degradation in quality. |  |  |  | D11¦D1104
| End of Life Problem Identified | D1105 | Problems traced to the device reaching the end of its useful life. |  |  |  | D11¦D1105
| Reuse of Single Use Device | D1106 | Problems traced to the use of the device more than once when it is designed for only one use. |  |  |  | D11¦D1106
| Sabotage or Intentional Misuse | D1107 | Problems occurred because of willful damage or reckless misuse. This term indicates that there was an apparent intent to cause damage or injury. |  |  | Editorial change to the definition on 16 February 2021. | D11¦D1107
| Reasonably Foreseeable Misuse | D1108 | Problems related to the use of a device in a way not intended in the instructions for use, but which may result from readily predictable human behavior. |  |  | Term was added on 27 January 2022. For details, see comment No. 10 of the Change Log (Release Number 2022). | D11¦D1108
Known Inherent Risk of Device |  | D12 | Reported adverse event known and documented in the labeling and all reasonable mitigation steps have been taken (including both short or long term known complications or adverse reactions). |  |  | The definition was modified on 21 January 2021 (addition of phrase "and all reasonable mitigation steps have been taken."). For details, refer to comment No. 31 of the Change Log. | D12
Falsified Device |  | D13 | Devices that deliberately and/or fraudulently misrepresent their identity, composition or source. |  |  |  | D13
No Problem Detected |  | D14 | Either it was not confirmed that there was a problem with the device, or it was confirmed that there was no problem with the device. |  |  | Definition modified for Release Number 2024. See no. 53 of Change Log | D14
| Unable to exclude device problem | D1401 | The investigation findings do not clearly confirm the presence or absence of the reported problem with the device. |  | Retired (2026) | Term added for Release Number 2024. See no. 54 of Change Log. Term retired in Release Number 2026, term moved to Annex C (C1901). See No. 3 Change Request Log. | D14¦D1401
| Device problem excluded | D1402 | The investigation findings clearly confirm that there was no problem with the device. |  | Retired (2026) | Term added for Release Number 2024. See no. 54 of Change Log. Term retired in Release Number 2026, term moved to Annex C (C1902). See No. 4 Change Request Log. | D14¦D1402
Cause not Established |  | D15 | The investigation findings do not lead to a clear conclusion about the cause of the reported adverse event. |  | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 248 of Change Log. | D15
| Cause Linked to Device but Unable to Trace More Specifically | D1501 | Problems traced to the device, but the investigation could not identify the actual root cause (e.g. Design, Manufacturing, Component). |  |  | New term/definition added on 21 January 2021. for details, refer to comment No. 30 of the Change Log. | D15¦D1501
Conclusion not Yet Available |  | D16 | A conclusion has yet to be established as the investigation is incomplete. Do not use this code if the investigation is complete. |  | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 249 of Change Log. | D16
Appropriate Investigation Conclusion Term/Code not Available |  | D17 | The concluded cause is not adequately described by any other term. Note: This code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. |  | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 250 of Change Log. | D17
Cause Traced to Software Coding |  | D18 | Problems traced to an error, flaw or fault in a computer program or system that causes it to produce an incorrect or unexpected result, or to behave in unintended ways. |  |  | New term/definition added on 21 January 2021. for details, refer to comment No. 27 of the Change Log. | D18
Cause Traced to Artificial Intelligence Training/Validation Process |  | D19 | Problem traced to the process or dataset(s) used for training and/or validating artificial intelligence, including machine learning algorithms. |  |  | Term was added on 27 January 2022. For details, see comment No. 12 of the Change Log (Release Number 2022). | D19
Cause Traced to Another Device |  | D20 | Problem is traced to be caused by another device. |  |  | Term was added on 31 January 2023. For details, see comment No. 75 of the Change Log (Release Number 2023). | D20
Cause Traced to Health Disparity |  | D21 | Problems traced to device performing inadequately or inconsistently in people with certain equality characteristics (e.g. sex, racial or ethnic origin). |  |  | Term was added on 31 January 2023. For details, see comment No. 142 of the Change Log (Release Number 2023). | D21

## E
Annex Name: Annex E
Annex Title: Health Effects - Clinical Signs and Symptoms or Conditions
Release Number: 2026
Annex Description: Terms/codes for describing the clinical signs and symptoms or conditions of the affected person appearing as a result of the medical device adverse event/incident. The MedDRA® trademark is registered by ICH. This information/material is provided "as is" without warranty of any kind. In no event shall the ICH or the authors of the original information/material be liable for any claim, damages or other liability arising from or in connection with the use of this information/material.
Annex Instructions: These terms allow capturing of the patient signs and symptoms observed related to the medical device adverse event through observational language without using diagnostic specifics. These terms are largely based on a subset of MedDRA terms. This annex is organized along organ systems as well as physiological problems. The Level 1 term is mainly for organizational purposes and is expected to be too high level for coding purposes. IMDRF Terms and Definitions use American spelling. Wherever appropriate the word 'device' should be interpreted as meaning all or any of the following: medical device, in-vitro diagnostic medical device, accessory to a medical device, medical device parts and/or components. Wherever appropriate the word 'patient' should be interpreted as meaning all or any of the following: user, operator, any other person affected by the incident.
NOTE: Maximum field length Term: 100 characters Definition: 700 Status description: 700
Annex Approval Date: January 2026
Level 1 Term | Level 2 Term | Level 3 Term | Code | Definition | Non-IMDRF Code | Primary Category | Secondary Category | Status | Status Description | CodeHierarchy
Nervous System |  |  | E01 | Nervous System |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E01
| Balance Problems |  | E0101 | A feeling of falling down which can occur whether the person is standing, sitting or lying down. | MedDRA:10049848:Balance disorder |  |  |  |  | E01¦E0101
| Brain Injury |  | E0102 | Damage to the brain. | MedDRA:10060690:Traumatic brain injury | Nervous System | Injury |  |  | E01¦E0102
|  | Encephalocele | E010201 | Hernia of brain substance and meninges through a congenital or traumatic opening of the skull. | MedDRA:10014617:Encephalocele | Nervous System | Injury |  |  | E01¦E0102¦E010201
| Cerebral Edema |  | E0103 | A swelling in the brain caused by the presence of excessive fluid. | MedDRA:10008107:Cerebral edema | Nervous System | Generalized Disorders |  |  | E01¦E0103
| Cerebral Hyperperfusion Syndrome |  | E0104 | Unexpected increase in cerebral blood flow after carotid endarterectomy (CEA) or carotid artery stenting (CAS). | MedDRA:10064730:Cerebral hyperperfusion syndrome | Nervous System | Vascular System |  |  | E01¦E0104
| Cerebral Ventriculomegaly |  | E0105 | Abnormal enlargement of the cerebral ventricles. | MedDRA:10068352:Cerebral ventriculomegaly |  |  |  | Term was modified on 21 January 2021 (old term: Cerebral Ventriculomeglia). | E01¦E0105
| Cerebrospinal Fluid Leakage |  | E0106 | The loss of cerebrospinal fluid into the surrounding tissues or outside the body. | MedDRA:10008164:Cerebrospinal fluid leakage |  |  |  | Definition modified for Release Number 2024. See no. 6 of Change Log | E01¦E0106
| Cognitive Changes |  | E0107 | Changes in perception, thinking, or remembering. | MedDRA:10057668:Cognitive disorder |  |  |  |  | E01¦E0107
|  | Confusion/ Disorientation | E010701 | A mental state characterized by a lack of clear and orderly thought and behavior. | MedDRA:10010300:Confusion |  |  |  |  | E01¦E0107¦E010701
|  | Dementia | E010702 | Loss of intellectual abilities interfering with an individual's social and occupational functions. | MedDRA:10012267:Dementia |  |  |  |  | E01¦E0107¦E010702
| Concussion |  | E0108 | Traumatic brain injury as a result of the action of a mechanical force on the head. | MedDRA:10010254:Concussion | Nervous System | Injury |  |  | E01¦E0108
| Convulsion/Seizure |  | E0109 | Sudden, involuntary skeletal muscular contractions of cerebral or brain stem origin. | MedDRA:10010904:Convulsion |  |  |  |  | E01¦E0109
|  | Convulsion, Clonic | E010901 | A convulsion marked by alternating contracting and relaxing of the muscles. | MedDRA:10053398:Clonic convulsion |  |  |  |  | E01¦E0109¦E010901
|  | Convulsion, Tonic | E010902 | A convulsion marked by prolonged contraction of the muscles. | MedDRA:10043994:Tonic convulsion |  |  |  |  | E01¦E0109¦E010902
|  | Epilepsy | E010903 | Epilepsy caused or apparently caused by device. Do not use when epilepsy is a pre-existing condition. | MedDRA:10015037:Epilepsy |  |  |  |  | E01¦E0109¦E010903
|  | Status Epilepticus | E010904 | A life-threatening condition characterised by a single prolonged seizure or a series of seizures without intervening full recovery of consciousness. | MedDRA:10041962:Status epilepticus |  |  |  | Editorial change to the definition on 16 February 2021. | E01¦E0109¦E010904
| Hyporesponsive to Stimuli |  | E0110 | Decreased sensitivity affecting one or more senses | MedDRA:10071552:Hyporesponsive to stimuli |  |  | Modified (editorial) | Term and Definition modified for Release Number 2025. Previous term was "Decreased Sensitivity". See no. 115 of Change Request Log. Editorial change to term for Release Number 2026. See no. 251 of Change Log. | E01¦E0110
| Hyperresponsive to Stimuli |  | E0111 | Increased sensitivity affecting one or more senses | MedDRA:10082489:Hyperresponsive to stimuli |  |  | Modified (editorial) | Term and Definition modified for Release Number 2025. Previous term was "Increased Sensitivity". See no. 119 of Change Request Log. Editorial change to term for Release Number 2026. See no. 252 of Change Log. | E01¦E0111
| Dizziness |  | E0112 | A sensation of lightheadedness, unsteadiness, turning, spinning or rocking. | MedDRA:10013573:Dizziness |  |  |  |  | E01¦E0112
| Dysphasia |  | E0113 | Impairment of verbal communication skills, often resulting from brain damage. | MedDRA:10013951:Dysphasia |  |  |  |  | E01¦E0113
|  | Aphonia | E011301 | Inability to speak. | MedDRA:10002953:Aphonia |  |  |  |  | E01¦E0113¦E011301
| Encephalitis |  | E0114 | An acute inflammatory process affecting the brain parenchyma. | MedDRA:10014581:Encephalitis | Nervous System | Infections |  |  | E01¦E0114
| Encephalopathy |  | E0115 | A functional and/or structural disorder of the brain that is acquired or congenital and characterised by mental and neurological symptoms. | MedDRA:10014625:Encephalopathy |  |  |  | Editorial change to the definition on 16 February 2021. | E01¦E0115
| Headache |  | E0116 | Pain in various parts of the head, not confined to the area of distribution of any nerve. | MedDRA:10019211:Headache | Nervous System | Generalized Disorders |  |  | E01¦E0116
| Hydrocephalus |  | E0117 | A disorder characterized by an abnormal increase of cerebrospinal fluid in the ventricles of the brain. | MedDRA:10020508:Hydrocephalus |  |  |  |  | E01¦E0117
| Intracranial Hemorrhage |  | E0118 | Bleeding within the cranium. | MedDRA:10022763:Intracranial hemorrhage |  |  |  |  | E01¦E0118
| Loss of Consciousness |  | E0119 | A level of awareness that can be described as consistently not responsive to stimuli. | MedDRA:10024855:Loss of consciousness |  |  | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 253 of Change Log. | E01¦E0119
|  | Coma | E011901 | A state of profound unconsciousness associated with markedly depressed cerebral activity. | MedDRA:10010071:Coma |  |  |  |  | E01¦E0119¦E011901
|  | Presyncope | E011902 | Extremely weak; threatened with syncope. | MedDRA:10036653:Presyncope |  |  |  |  | E01¦E0119¦E011902
|  | Syncope/Fainting | E011903 | A spontaneous, transient, loss of consciousness caused by insufficient blood supply to the brain. | MedDRA:10042772:Syncope |  |  |  |  | E01¦E0119¦E011903
| Memory Loss/Impairment |  | E0120 | Systematic and extensive loss or impairment of memory caused by organic or psychological factors. This may be temporary or permanent, and may involve old or recent memories. | MedDRA:10027176:Memory loss |  |  |  |  | E01¦E0120
| Meningitis |  | E0121 | A disorder characterized by acute inflammation of the meninges of the brain and/or spinal cord. | MedDRA:10027199:Meningitis | Nervous System | Infections |  |  | E01¦E0121
| Movement Disorder |  | E0122 | Neurological conditions resulting in abnormal voluntary or involuntary movement, which may impact the speed, fluency, quality and ease of movement. | MedDRA:10028035:Movement disorder |  |  |  |  | E01¦E0122
|  | Dyskinesia | E012201 | Difficulty moving; distortion or impairment of voluntary movement, as in tic, spasm, or myoclonus. | MedDRA:10013916:Dyskinesia |  |  |  |  | E01¦E0122¦E012201
|  | Paralysis | E012202 | Partial or complete loss of function of one or more muscles. It is usually caused by damage to the nervous system. | MedDRA:10033799:Paralysis |  |  |  |  | E01¦E0122¦E012202
|  | Paraplegia | E012203 | Paralysis of the legs and lower part of the body. | MedDRA:10033892:Paraplegia |  |  |  |  | E01¦E0122¦E012203
|  | Paresis | E012204 | A slight or incomplete paralysis. | MedDRA:10033985:Paresis |  |  |  |  | E01¦E0122¦E012204
|  | Quadriplegia | E012205 | Paralysis of all four limbs. | MedDRA:10037714:Quadriplegia |  |  |  |  | E01¦E0122¦E012205
|  | Shaking/Tremors | E012206 | The shaking movement of the whole body or just a certain part of it, often caused by problems of the neurons responsible for muscle action. | MedDRA:10040528:Shaking |  |  |  |  | E01¦E0122¦E012206
|  | Twitching | E012207 | The occurrence of a single contraction or a series of contractions of a muscle. | MedDRA:10045198:Twitching |  |  |  |  | E01¦E0122¦E012207
|  | Restless Legs Syndrome | E012208 | A condition that is characterized by an irresistible urgency to move the legs to obtain relief from a strange and uncomfortable sensation in the legs. | MedDRA:10058920:Restless legs syndrome |  |  |  | Term was added on 27 January 2022. For details, see comment No. 85 of the Change Log (Release Number 2022). | E01¦E0122¦E012208
|  | Facial Paralysis | E012209 | A partial or complete loss of function of one or more of the facial muscles resulting in paralysis of the facial regions. Includes both central and peripheral facial paralysis or palsy. | MedDRA:10016062:Facial paralysis |  |  |  | Term was added on 27 January 2022. For details, see comment No. 242 / 243 / 244 / 245 of the Change Log (Release Number 2022). | E01¦E0122¦E012209
|  | Dystonia | E012210 | Movement disorder characterized by involuntary muscle contractions. | MedDRA:10013983:Dystonia |  |  | New | New term added for Release Number 2026. See no.112 of Change Request Log. | E01¦E0122¦E012210
|  | Apraxia | E012211 | Inability to perform particular purposive actions | MedDRA:10003062:Apraxia |  |  | New | New term added for Release Number 2026. See no.113 of Change Request Log. | E01¦E0122¦E012211
| Nerve Damage |  | E0123 | Injury to nervous tissue. | MedDRA:10029177:Nerve damage | Nervous System | Injury |  |  | E01¦E0123
|  | Optical Nerve Damage | E012301 | Injury to optical nerve. | MedDRA:10030938:Optic nerve injury | Nervous System | Injury |  |  | E01¦E0123¦E012301
| Nervous System Injury |  | E0124 | Injury to the central or peripheral nervous system. | MedDRA:10080898:Nervous system injury | Nervous System | Injury |  |  | E01¦E0124
|  | Spinal Cord Injury | E012401 | Traumatic damage of the spinal cord. | MedDRA:10041552:Spinal cord injury | Nervous System | Injury |  |  | E01¦E0124¦E012401
| Neuralgia |  | E0125 | Intense painful sensation along a nerve or group of nerves. | MedDRA:10029223:Neuralgia |  |  |  |  | E01¦E0125
| Neuropathy |  | E0126 | A disorder affecting the cranial nerves or the peripheral nervous system. It manifests with pain, tingling, numbness, and muscle weakness. | MedDRA:10029328:Neuropathy |  |  |  |  | E01¦E0126
| Numbness |  | E0127 | The loss of the sensation of feeling in an area of the body. | MedDRA:10029829:Numbness |  |  |  |  | E01¦E0127
| Peripheral Nervous Injury |  | E0128 | Injury to the peripheral nervous system. | MedDRA:10034586:Peripheral nerve injury | Nervous System | Injury |  |  | E01¦E0128
| Sensitivity of Teeth |  | E0129 | Pain or discomfort caused by exposure of the dentin layer of tooth to thermal, tactile, or other stimuli. | MedDRA:10040012:Sensitivity of teeth |  |  |  |  | E01¦E0129
| Sleep Dysfunction |  | E0130 | A change from the patient's baseline sleeping pattern, in the hours slept and/or an alteration/dysfunction in the stages of sleep. | MedDRA:10040984:Sleep disorder |  |  |  |  | E01¦E0130
| Speech Disorder |  | E0131 | A term referring to disorders characterized by the disruption of normal speech. It includes stuttering, lisps, dysarthria and voice disorders. | MedDRA:10041466:Speech disorder |  |  |  |  | E01¦E0131
| Spinal Arachnoiditis |  | E0132 | A chronic adhesive arachnoiditis in the spinal arachnoid, with root and spinal cord symptoms similar to those caused by pressure from a tumor. | MedDRA:10068160:Spinal arachnoiditis |  |  |  |  | E01¦E0132
| Stroke/CVA |  | E0133 | A sudden loss of neurological function secondary to hemorrhage or ischemia in the brain parenchyma due to a vascular event. | MedDRA:10042244:Stroke |  |  |  |  | E01¦E0133
|  | Hemorrhagic Stroke | E013301 | An acute neurological deficit caused by bleeding within the brain or between the brain and the skull. | MedDRA:10048863:Hemorrhagic stroke |  |  |  |  | E01¦E0133¦E013301
|  | Ischemia Stroke | E013302 | An ischemic condition of the brain, producing a persistent focal neurological deficit in the area of distribution of the cerebral arteries. | MedDRA:10055221:Ischemic stroke |  |  |  |  | E01¦E0133¦E013302
| Tactile Disorders |  | E0134 | conditions characterized by an alteration in the ability to or perception of touch. | MedDRA:10062872:Dysesthesia |  |  |  |  | E01¦E0134
|  | Hyperesthesia | E013401 | Increased sensitivity to tactile stimulation. | MedDRA:10020622:Hyperesthesia |  |  |  |  | E01¦E0134¦E013401
|  | Hypoesthesia | E013402 | Impairment of tactile sensitivity manifesting as partial loss of sensitivity to sensory stimuli. | MedDRA:10020976:Hypoesthesia |  |  |  |  | E01¦E0134¦E013402
|  | Paresthesia | E013403 | Abnormal touch sensations, such as burning or prickling, that occur without an outside stimulus. | MedDRA:10033987:Paresthesia |  |  |  |  | E01¦E0134¦E013403
| Taste Disorder |  | E0135 | Conditions characterized by an alteration in taste function or perception | MedDRA:10082490:Taste disorder |  |  |  | Editorial change to the definition on 16 February 2021. | E01¦E0135
|  | Ageusia | E013501 | Inability to taste. | MedDRA:10001480:Ageusia |  |  |  |  | E01¦E0135¦E013501
|  | Dysgeusia | E013502 | Distortion of taste. | MedDRA:10013911:Dysgeusia |  |  |  |  | E01¦E0135¦E013502
| Tics/Tremor |  | E0136 | An involuntary, compulsive, repetitive stereotyped movement. | MedDRA:10066770:Motor tic |  |  |  |  | E01¦E0136
| Transient Ischemic Attack |  | E0137 | A brief attack (from a few minutes to an hour) of cerebral dysfunction of vascular origin, with no persistent neurological deficit. | MedDRA:10072760:Transient ischemic attack |  |  |  |  | E01¦E0137
| Undesired Nerve Stimulation |  | E0138 | Unintended reaction in nerve system caused by electrical stimulation. | MedDRA:10068145:Undesired device stimulation of nerve |  |  | Retired (2025) | Term retired for Release Number 2025. Use E2131 - Inappropriate Device Stimulation of Tissue. See no. 17 of Change Request Log | E01¦E0138
| Unspecified Nervous System Problem |  | E0139 | The report describes a non-specific problem with the nervous system. Note: Please use E2402 if the report describes a specific injury but the relevant term does not exist. |  |  |  | Modified (editorial) | Term was added on 18 March 2020 for cases where the category of a health effect was reported but with no further details. Editorial change to definition for Release Number 2026. See no. 196 of Change Log. | E01¦E0139
| Intracranial Pressure Increased |  | E0140 | Pressure within the craniospinal compartment is increased. | MedDRA:10022773:Intracranial pressure increased |  |  |  | Term was added on 27 January 2022. For details, see comment No. 86 of the Change Log (Release Number 2022). | E01¦E0140
| Pneumocephalus |  | E0141 | The presence of air or gas within the cranial cavity. | MedDRA:10048736:Pneumocephalus |  |  |  | Term was added on 27 January 2022. For details, see comment No. 212 of the Change Log (Release Number 2022). | E01¦E0141
| Idiopathic Intracranial Hypertension |  | E0142 | A condition that mimics the presence of a tumor (Pseudotumour Cerebri) or hydrocephalus, caused by chronic increase in intracranial pressure. | MedDRA:10078904:Idiopathic intracranial hypertension |  |  | Modified (editorial) | Term was added on 27 January 2022. For details, see comment No. 219 of the Change Log (Release Number 2022). Editorial change to definition for Release Number 2026. See no. 186 of Change Log. | E01¦E0142
| Altered State of Consciousness |  | E0143 | A level of awareness and arousal other than normal. | MedDRA:10001854:Altered state of consciousness |  |  |  | Term was added on 27 January 2022. For details, see comment No. 77 of the Change Log (Release Number 2022). | E01¦E0143
|  | Increased Level of Consciousness | E014301 | A neurological state characterised by increased ability to perceive and respond. | MedDRA:10021683:Increased state of alertness |  |  |  | Term was added on 27 January 2022. For details, see comment No. 77 of the Change Log (Release Number 2022). | E01¦E0143¦E014301
|  | Decreased Level of Consciousness | E014302 | A neurological state characterised by decreased ability to perceive and respond. | MedDRA:10012373:Depressed level of consciousness |  |  |  | Term was added on 27 January 2022. For details, see comment No. 77 of the Change Log (Release Number 2022). | E01¦E0143¦E014302
| Dural Tear |  | E0144 | Injury to the dura mater, with or without damage to any other layer of the meninges. | MedDRA:10063395:Dural tear | Nervous System | Procedural Complications |  | Term added for Release Number 2024. See no. 22 of Change Log | E01¦E0144
| Olfactory Changes |  | E0145 | Conditions characterized by an alteration in the function or perception of smell, including diminished, distorted, or lost olfactory sensation. | MedDRA:10001853:Altered smell sensation |  |  | New | Term added for Release Number 2026. See no. 111 of Change Log. | E01¦E0145
| Hyporeflexia |  | E0146 | Diminished or absent reflex responses | MedDRA:10021089:Hyporeflexia | Nervous System | Musculoskeletal | New | Term added for Release Number 2026. See no. 14 of Change Log. | E01¦E0146
| Acoustic Shock |  | E0901 | Neurophysiological and psychological symptoms caused by sudden and unexpected loud noise. | MedDRA:10068150:Acoustic shock | Ear and Labyrinth | Nervous System |  |  | E01¦E0901
| Tinnitus |  | E0904 | A noise in the ears, such as ringing, buzzing, roaring or clicking in the absence of auditory stimulation. | MedDRA:10043882:Tinnitus | Ear and Labyrinth | Nervous System |  |  | E01¦E0904
| Perinatal Brain Injury |  | E1513 | Injury to the brain of a newborn infant occurring around the time of birth. | MedDRA:10056392:Perinatal brain damage | Pregnancy, Childbirth and the Puerperium | Nervous System |  |  | E01¦E1513
| Muscle Weakness/Atrophy |  | E1621 | A reduction in the strength and/or mass of one or more muscles. | MedDRA:10028350:Muscle weakness | Musculoskeletal System | Nervous System |  | Term/Definition was modified on 27 January 2022 (Old term/definition: Muscle Weakness: A reduction in the strength of one or more muscles.) For details, see comment No. 120 of the Change Log (Release Number 2022). | E01¦E1621
| Complex regional pain syndrome |  | E2125 | Condition where a person experiences persistent severe and debilitating pain, involving an extremity or other body region, hyperesthesia, and localized autonomic dysfunction following injury to soft tissue or nerve | MedDRA:10064332:Complex regional pain syndrome | Procedural Complications | Nervous System |  | Term added for Release Number 2024. See no. 56 of Change Log. Editorial change to term for Release Number 2026. See no. 295 of Change Log. | E01¦E2125
Mental, Emotional and Behavioral Disorders |  |  | E02 | Mental, Emotional and Behavioral Disorders |  |  |  | Not selectable. Modified (editorial) | Please use more detailed term within the hierarchy. Editorial amendment to term and definition for Release Number 2026. See no. 187 of Change Log. | E02
| Bruxism |  | E0201 | Excessive clenching of the jaw and grinding of the teeth. | MedDRA:10006514:Bruxism |  |  |  |  | E02¦E0201
| Emotional Changes |  | E0202 | Change from the affected person's usual emotional state. | MedDRA:10014551:Emotional disorder |  |  |  |  | E02¦E0202
|  | Anxiety | E020201 | Apprehension or uneasiness which may be accompanied by restlessness, tension, tachycardia, and dyspnea. | MedDRA:10002855:Anxiety |  |  |  |  | E02¦E0202¦E020201
|  | Depression | E020202 | A condition marked by ongoing feelings of sadness, despair, loss of energy, diminished interest/pleasure and difficulty dealing with normal daily life. | MedDRA:10012378:Depression |  |  |  |  | E02¦E0202¦E020202
|  | Distress | E020203 | A state of physiological or psychological stress that cannot be compensated for by normal adaptive measures. | MedDRA:10049119:Emotional distress |  |  |  |  | E02¦E0202¦E020203
|  | Malaise | E020204 | A feeling of general discomfort or uneasiness, an out-of-sorts feeling. | MedDRA:10025482:Malaise |  |  |  |  | E02¦E0202¦E020204
|  | Irritability | E020205 | A feeling of agitation. Tending to become frustrated or upset easily. | MedDRA:10022998:Irritability |  |  |  |  | E02¦E0202¦E020205
| Hallucination |  | E0203 | A false sensory perception in the absence of an external stimulus. | MedDRA:10019063:Hallucination |  |  |  |  | E02¦E0203
| Lethargy |  | E0204 | Characterized by a lack of vitality or energy. | MedDRA:10024264:Lethargy |  |  |  |  | E02¦E0204
| Suicidal Ideation |  | E0205 | Thoughts of taking one's own life. | MedDRA:10042458:Suicidal ideation |  |  |  |  | E02¦E0205
| Unspecified Mental, Emotional or Behavioral Problem |  | E0206 | The report describes a non-specific Mental, Emotional or Behavioral problem. Note: Please use E2402 if the report describes a specific problem but the relevant term does not exist. |  |  |  | Modified (editorial) | Term was added on 18 March 2020 for cases where the category of a health effect was reported but with no further details. Editorial changes to term and definition for Release Number 2026. See no. 188 and 197 of Change Log. | E02¦E0206
| Delirium |  | E0207 | A usually reversible condition characterized by the acute and sudden development of confusion, illusions, movement changes, inattentiveness, agitation and hallucinations. | MedDRA:10012218:Delirium |  |  |  | Term was added on 27 January 2022. For details, see comment No. 36 of the Change Log (Release Number 2022). | E02¦E0207
| Psychosis |  | E0208 | A mental disorder characterized by personality change, impaired functioning and loss of touch with reality. | MedDRA:10037234:Psychosis |  |  |  | Term was added on 27 January 2022. For details, see comment No. 37 of the Change Log (Release Number 2022). | E02¦E0208
| Sleepiness/Drowsiness/Somnolence |  | E0209 | The sensation of struggling to remain awake. | MedDRA:10041349:Somnolence |  |  |  | Term was added on 31 January 2023. For details, see comment No. 77 of the Change Log (Release Number 2023). | E02¦E0209
| Paranoia |  | E0210 | Excessive mistrust or suspicion of people -thinking incorrectly, for example, that someone is trying to hurt you, that people are doing things to deliberately annoy you or that there is a conspiracy against you. | MedDRA:10033864:Paranoia |  |  | New | Term added for Release Number 2026. See no. 65 of Change Log. | E02¦E0210
Blood and Lymphatic System |  |  | E03 | Blood and Lymphatic System |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E03
| Anemia |  | E0301 | A reduction in the number of red blood cells, the amount of hemoglobin, and/or the volume of packed red blood cells. | MedDRA:10002272:Anemia |  |  |  |  | E03¦E0301
|  | Hemolytic Anemia | E030101 | Anemia resulting from the premature destruction of the peripheral blood red cells. | MedDRA:10019493:Hemolytic anemia |  |  |  |  | E03¦E0301¦E030101
| Coagulation Disorder |  | E0302 | A condition in which there is a deviation from or interruption of the normal coagulation properties of the blood. | MedDRA:10009731:Coagulation disorder |  |  |  |  | E03¦E0302
|  | Disseminated Intravascular Coagulation (DIC) | E030201 | A pathological process where the blood starts to coagulate throughout the whole body. This depletes the body of its platelets and coagulation factors, and there is an increased risk of hemorrhage. | MedDRA:10013442:Disseminated intravascular coagulation |  |  |  |  | E03¦E0302¦E030201
|  | Thrombocytopenia | E030202 | Low blood platelet count. | MedDRA:10043554:Thrombocytopenia |  |  |  |  | E03¦E0302¦E030202
| Hemolysis |  | E0303 | Disruption of the integrity of the erythrocyte membrane causing release of hemoglobin. | MedDRA:10019491:Hemolysis |  |  |  |  | E03¦E0303
| Hypervolemia |  | E0304 | An abnormally high circulating blood volume. | MedDRA:10020920:Hypervolemia |  |  |  | Definition was modified on 21 January 2021 (Old definition: Too much fluid in the blood.). For details, see comment No. 2 of the Change Log. | E03¦E0304
| Hypovolemia |  | E0305 | An abnormally low circulating blood volume. | MedDRA:10021139:Hypovolemia |  |  |  |  | E03¦E0305
| Sepsis |  | E0306 | The presence of pathogenic microorganisms in the blood stream causing a rapidly progressing systemic reaction that may lead to shock. Symptoms include fever, chills, tachycardia, and increased respiratory rate. It is a medical emergency that requires urgent medical attention. | MedDRA:10040047:Sepsis | Blood and Lymphatic System | Infections |  |  | E03¦E0306
| Seroma |  | E0307 | A tumor-like collection of serum in the tissues. | MedDRA:10040102:Seroma |  |  |  |  | E03¦E0307
| Swollen Lymph Nodes/Glands |  | E0308 | The presence of an enlarged lymph node. | MedDRA:10067201:Swollen lymph nodes |  |  |  |  | E03¦E0308
| Low White Blood Cell Count |  | E0309 | The white blood cell count is below the normal range, including neutropenia. | MedDRA:10051617:White blood cell count low |  |  |  | Definition was modified on 27 January 2022 (Old definition:The white blood cell count is below the normal range.). MedDRA mapping also updated from White blood cell count decreased (10047942) to White blood cell count low (10051617). For details, see comment No. 73 of the Change Log (Release Number 2022). | E03¦E0309
| Unspecified Blood or Lymphatic Problem |  | E0310 | The report describes a non-specific problem with the blood or lymphatic system. Note: Please use E2402 if the report describes a specific problem but the relevant term does not exist. |  |  |  | Modified (editorial) | Term was added on 18 March 2020 for cases where the category of a health effect was reported but with no further details. Term and definition modified for Release Number 2026. See no. 198 and 254 of Change log. | E03¦E0310
| High White Blood Cell Count |  | E0311 | The white blood cell count is above the normal range. | MedDRA:10051771:White blood cell count high |  |  |  | Term was added on 27 January 2022. For details, see comment No. 79 of the Change Log (Release Number 2022). | E03¦E0311
| Transfusion Reaction |  | E0312 | Adverse events associated with the transfusion of whole blood or one of its components, most commonly a hemolytic reaction to cross-typed blood. | MedDRA:10044359:Transfusion reaction |  |  |  | Term was added on 27 January 2022. For details, see comment No. 202 of the Change Log (Release Number 2022). | E03¦E0312
| Blood Protein Abnormal |  | E0313 | One or more of the blood protein levels being higher or lower than normal (e.g. Albumin, C-reactive Protein). |  |  |  |  | Term was added on 27 January 2022. For details, see comment No. 108 / 111 of the Change Log (Release Number 2022). | E03¦E0313
| Electrolyte Imbalance |  | E1202 | Higher or lower than normal values for the serum electrolytes; usually affecting Na, K, Cl, Ca, CO2, BUN. | MedDRA:10014418:Electrolyte imbalance | Endocrine, Metabolism and Nutrition | Blood and Lymphatic System |  |  | E03¦E1202
|  | Hypernatremia | E120201 | Higher than normal levels of sodium in the circulating blood. | MedDRA:10020680:Hypernatremia | Endocrine, Metabolism and Nutrition | Blood and Lymphatic System |  |  | E03¦E1202¦E120201
|  | Hyponatremia | E120202 | Lower than normal levels of sodium in the circulating blood. | MedDRA:10021038:Hyponatremia | Endocrine, Metabolism and Nutrition | Blood and Lymphatic System |  |  | E03¦E1202¦E120202
Immune System |  |  | E04 | Immune System |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E04
| Autoimmune Disorder |  | E0401 | A disorder resulting from loss of function or tissue destruction of an organ or multiple organs, arising from humoral or cellular immune responses of the individual to his own tissue constituents. | MedDRA:10061664:Autoimmune disorder |  |  |  |  | E04¦E0401
| Hypersensitivity/Allergic Reaction |  | E0402 | A local or general reaction of an organism following contact with a specific allergen to which it has been previously exposed and to which it has become sensitized. | MedDRA:10020751:Hypersensitivity |  |  | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 255 of Change Log. | E04¦E0402
|  | Anaphylactic Shock | E040201 | Progression of anaphylaxis into systemic shock, which is characterized by significant reduction in tissue perfusion which leads to hypoxia and hypotension, and can lead to death if untreated. | MedDRA:10002199:Anaphylactic shock | Immune System | Generalized Disorders |  |  | E04¦E0402¦E040201
|  | Anaphylactoid / Anaphylaxis | E040202 | Acute immune response that occurs from exposure to an allergen. For conditions which have progressed to systemic shock, use E040201 - Anaphylactic Shock. | MedDRA:10002218:Anaphylaxis | Immune System | Generalized Disorders |  | Definition was modified on 31 January 2023 (Old definition: Resembling anaphylaxis.) MedDRA mapping also updated from Anaphylactoid reaction (10002216) to Anaphylaxis (10002218). For details, see comment No. 78 of the Change Log (Release Number 2023). | E04¦E0402¦E040202
|  | Local Reaction | E040203 | A regionally-limited response to an antigen, which may include inflammation, induration, erythema, pruritus or pain. | MedDRA:10024769:Local reaction |  |  |  |  | E04¦E0402¦E040203
| Immunodeficiency |  | E0403 | A deficiency of immune response or a disorder characterized by deficient immune response. | MedDRA:10061598:Immunodeficiency |  |  |  |  | E04¦E0403
| Positive Antinuclear Antibodies (ANA) |  | E0404 | Indicates that antibodies directed against nuclear antigens have been detected. Almost invariably found in systemic lupus erythematosus and are frequently found in rheumatoid arthritis, scleroderma, Sjogren's Syndrome and mixed connective tissue disease. | MedDRA:10060055:Antinuclear antibody positive |  |  | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 256 of Change Log. | E04¦E0404
| Unspecified Immune System Problem |  | E0405 | The report describes a non-specific problem with the immune system. Note: Please use E2402 if the report describes a specific problem but the relevant term does not exist. |  |  |  | Modified (editorial) | Term was added on 18 March 2020 for cases where the category of a health effect was reported but with no further details. Editorial change to definition for Release Number 2026. See no. 199 of Change Log. | E04¦E0405
Vascular System |  |  | E05 | Vascular System |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E05
| Cerebral Hyperperfusion Syndrome |  | E0104 | Unexpected increase in cerebral blood flow after carotid endarterectomy (CEA) or carotid artery stenting (CAS). | MedDRA:10064730:Cerebral hyperperfusion syndrome | Nervous System | Vascular System |  |  | E05¦E0104
| Aneurysm |  | E0501 | Bulging or ballooning in an area of an artery secondary to arterial wall weakening. | MedDRA:10002329:Aneurysm |  |  |  |  | E05¦E0501
|  | Ruptured Aneurysm | E050101 | An aneurysm that has burst, resulting in hemorrhage into adjacent tissues. | MedDRA:10048380:Aneurysm ruptured |  |  |  |  | E05¦E0501¦E050101
|  | Aneurysm Expansion | E050102 | Expansion of aneurysm. | MedDRA:10069506:Aneurysm enlargement |  |  | Modified (editorial) | Term was added on 27 January 2022. For details, see comment No. 185 of the Change Log (Release Number 2022). Editorial change to term for Release Number 2026. See no. 257 of Release Log. | E05¦E0501¦E050102
| Arteriosclerosis/ Atherosclerosis |  | E0502 | A vascular disorder characterized by thickening and hardening of the walls of the arteries. Build-up of fatty material and calcium deposition in the arterial wall resulting in partial or complete occlusion of the arterial lumen. | MedDRA:10003210:Arteriosclerosis |  |  |  |  | E05¦E0502
| Embolism/Embolus |  | E0503 | The blockage of a blood vessel lumen by air or solid material such as blood clot or other tissues (e.g., adipose tissue, cancer cells) that have migrated from another anatomic site. | MedDRA:10061169:Embolism |  |  |  |  | E05¦E0503
|  | Air Embolism | E050301 | The presence of bubbles of air in the vascular system; occurrence is related to the entry of air into the venous circulation following trauma or surgery. | MedDRA:10001526:Air embolism |  |  |  |  | E05¦E0503¦E050301
|  | Foreign Body Embolism | E050302 | Type of embolism when lumen of vessel or ductus is occluded by part of device. | MedDRA:10075247:Foreign body embolism |  |  |  |  | E05¦E0503¦E050302
|  | Pulmonary Embolism | E050303 | The obstruction of the pulmonary artery or one of its branches by an embolus, sometimes associated with infarction of the lung. | MedDRA:10037377:Pulmonary embolism | Vascular System | Respiratory System |  |  | E05¦E0503¦E050303
|  | Thromboembolism | E050304 | Occlusion of the lumen of a vessel by a thrombus that has migrated from a distal site via the blood stream. | MedDRA:10043566:Thromboembolism |  |  |  |  | E05¦E0503¦E050304
|  | Fat Embolism | E050305 | The blockage of a blood vessel by fat particles in the circulation. | MedDRA:10016246:Fat embolism |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 57 of Change Log. Editorial change to term for Release Number 2026. See no. 258 of Release Log. | E05¦E0503¦E050305
| Extravasation |  | E0504 | The leakage of body fluid or cells from the vascular system to the surrounding tissues. This process occurs either by force or as a result of a pathologic reaction (e.g., inflammation). | MedDRA:10015866:Extravasation |  |  |  |  | E05¦E0504
| Hematoma |  | E0505 | A localized collection of blood, usually clotted, in an organ, space, or tissue, due to a break in the wall of a blood vessel. | MedDRA:10019428:Hematoma |  |  |  |  | E05¦E0505
| Hemorrhage/Blood Loss/Bleeding |  | E0506 | The loss of blood from the circulation, including blood loss from a device connected to a patient. | MedDRA:10019524:Hemorrhage |  |  |  | Term and definition was modified on 27 January 2022 (Old term/definition: Hemorrhage/Bleeding: The loss of blood from a ruptured blood vessel.). For details, see comment No. 87 / 199 of the Change Log (Release Number 2022). | E05¦E0506
|  | Ecchymosis | E050601 | A flat discoloration of the skin or a mucous membrane, larger in area than a petechia, caused by the extravasation of blood. Discolouration under the skin caused by a bleed underneath. Typically a bruise. | MedDRA:10014080:Ecchymosis |  |  |  |  | E05¦E0506¦E050601
|  | Exsanguination | E050602 | Extensive loss of blood due to internal or external hemorrhage. | MedDRA:10015719:Exsanguination |  |  |  |  | E05¦E0506¦E050602
| High Pulmonary Arterial Wedge Pressure |  | E0507 | The measurement of the mean left atrial pressure, as measured by a catheter introduced into the distal pulmonary artery, is high. | MedDRA:10037328:Pulmonary arterial wedge pressure high | Vascular System | Investigations and Diagnostic Tests |  |  | E05¦E0507
| Low Pulmonary Arterial Wedge Pressure |  | E0508 | The measurement of the mean left atrial pressure, as measured by a catheter introduced into the distal pulmonary artery, is low. | MedDRA:10037330:Pulmonary arterial wedge pressure low | Vascular System | Investigations and Diagnostic Tests |  | Definition was modified on 21 January 2021 (Old definition: The measurement of the mean left pulmonary atrial pressure, as measured by a catheter introduced into the distal pulmonary artery, is low.). For details, see comment No. 3 of the Change Log. | E05¦E0508
| Ischemia |  | E0509 | Lack of oxygen in tissues, commonly due to impaired blood supply to an area of the body, resulting in impairment of tissue oxygenation. | MedDRA:10054438:Ischemia |  |  |  |  | E05¦E0509
| Neovascularization |  | E0510 | The formation of new blood vessels. | MedDRA:10029114:Neovascularization |  |  |  |  | E05¦E0510
| Perforation of Vessels |  | E0511 | Damage to a blood vessel resulting from a full thickness disruption in the integrity of the vessel wall. | MedDRA:10068149:Vessel perforation | Vascular System | Procedural Complications |  |  | E05¦E0511
|  | Great Vessel Perforation | E051101 | Damage to the venae cavae, pulmonary artery, pulmonary veins or aorta resulting from a full thickness disruption in the integrity of the vessel wall. | MedDRA:10080896:Perforation of great vessels | Vascular System | Procedural Complications |  |  | E05¦E0511¦E051101
| Vasculitis |  | E0512 | Inflammation of any one of the blood vessels, including the arteries, veins, and rest of the vasculature system in the body. | MedDRA:10047115:Vasculitis |  |  | Modified (editorial) | Term and Definition and MedDRA map modified for Release Number 2024. Previous term was Phlebitis. Previous MedDRA map was 10034879:Phlebitis. See no. 19 of Change Log. Editorial change to definition for Release Number 2026. See no. 175 of Release Log. | E05¦E0512
| Pseudoaneurysm |  | E0513 | The creation of hematoma outside the arterial wall following bleeding due to puncture of the arterial wall. Pseudoaneurysms can also occur in the heart chambers following myocardial infarction with rupture of the wall of the ventricle. | MedDRA:10048977:Pseudoaneurysm |  |  |  |  | E05¦E0513
| Thrombosis/Thrombus |  | E0514 | The formation of a blood clot in the lumen of a vessel or heart chamber. | MedDRA:10043607:Thrombosis |  |  |  |  | E05¦E0514
| Vascular Dissection |  | E0515 | Spontaneous or induced tear within the wall of a blood vessel. | MedDRA:10070693:Vascular dissection |  |  |  |  | E05¦E0515
|  | Aortic Dissection | E051501 | Spontaneous or induced tear within the wall of the aorta | MedDRA:10002895:Aortic dissection |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 83 of Change Log. Editorial change to term for Release Number 2026. See no. 259 of Release Log. | E05¦E0515¦E051501
| Vasoconstriction |  | E0516 | Vasoconstriction involves narrowing of a blood vessel lumen due to smooth muscle contraction in the vessel wall and leads to decreased blood flow. | MedDRA:10047139:Vasoconstriction |  |  |  |  | E05¦E0516
| Vasodilatation |  | E0517 | Widening of a blood vessel lumen due to smooth muscle relaxation in the vessel wall and leads to increased blood flow. | MedDRA:10047141:Vasodilatation |  |  |  |  | E05¦E0517
| Unspecified Vascular Problem |  | E0518 | The report describes a non-specific problem with the vascular system. Note: Please use E2402 if the report describes a specific problem but the relevant term does not exist. |  |  |  | Modified (editorial) | Term was added on 18 March 2020 for cases where the category of a health effect was reported but with no further details. Editorial change to definition for Release Number 2026. See no. 200 of Release Log. | E05¦E0518
| Reduced Blood Flow |  | E0519 | Inadequate blood flow to a particular area of the body. | MedDRA:10054880:Vascular insufficiency |  |  |  | Term was added on 27 January 2022. For details, see comment No. 190 of the Change Log (Release Number 2022). | E05¦E0519
| Reperfusion Injury |  | E0520 | Tissue damage caused when blood supply returns to tissue after a period of ischemia or lack of oxygen. The restoration of circulation results in inflammation and oxidative damage. | MedDRA:10063837:Reperfusion injury |  |  |  | Term was added on 31 January 2023. For details, see comment No. 125 of the Change Log (Release Number 2023). | E05¦E0520
| Coronary Obstruction/Occlusion |  | E0521 | Partial or total occlusion or obstruction of the coronary ostia or artery. | MedDRA:10011103:Coronary occlusion | Vascular System | Heart | Modified (editorial) | Term added for Release Number 2024. See no. 82 of Change Log. Editorial change to term for Release Number 2026. See no. 260 of Release Log. | E05¦E0521
| Vasospasm |  | E0522 | Contraction of muscular walls of blood vessels | MedDRA:10047163:Vasospasm |  |  |  | Term added for Release Number 2025. See no. 35 of Change Log | E05¦E0522
| False Lumen Perfusion |  | E0523 | Persistent blood flow in the false lumen of a dissection following an endovascular aortic repair (EVAR) procedure. | MedDRA:10074841:Continued perfusion of a false lumen |  |  |  | Term added for Release Number 2025. See no. 66 of Change Log | E05¦E0523
| Vasoplegia |  | E0524 | A condition characterized by abnormally low systemic vascular resistance, often presenting with hypotension and preserved or elevated cardiac output; may be referred to as vasodilatory shock or vascular tone loss | MedDRA:10067080:Vasoplegia syndrome |  |  | New | Term added for Release Number 2026. See no. 17 of Change Log | E05¦E0524
| Central Venous Pressure Abnormal |  | E0525 | A deviation from normal central venous pressure values | MedDRA:10007978:Central venous pressure abnormal |  |  | New | Term added for Release Number 2026. See no. 26 of Change Log | E05¦E0525
| Telangiectasia |  | E0526 | Telangiectasia, also known as Spider veins or thread veins, are prominent clusters of small, damaged blood vessels visible in the skin. They appear as thin red, blue, or purple lines. | MedDRA:10043189:Telangiectasia | Vascular System | Skin | New | Term added for Release Number 2026. See no. 67 of Change Log | E05¦E0526
| Cardiovascular Dysfunction/Insufficiency |  | E0607 | Circulatory dysfunction/insufficiency due to a decrease in the force of the cardiac contractions and / or the tone of the vascular walls. Includes diastolic and systolic dysfunction. | MedDRA:10065929:Cardiovascular insufficiency | Heart | Vascular System |  | Term and definition was modified on 31 January 2023 (Old term/definition: Cardiovascular Insufficiency: Circulatory insufficiency due to a decrease in the force of the cardiac contractions and / or the tone of the vascular walls.). For details, see comment No. 81 of the Change Log (Release Number 2023). Error in secondary CodeHierarchy corrected for Release No. 2024 | E05¦E0607
| Bronchial Hemorrhage |  | E0707 | Bleeding from the bronchial wall and/or lung parenchyma. | MedDRA:10065757:Bronchial hemorrhage | Respiratory System | Vascular System |  |  | E05¦E0707
| Conjunctival Hemorrhage |  | E0850 | Blood extravasation in the conjunctiva. | MedDRA:10010720:Conjunctival hemorrhage | Eye | Vascular System |  | Term was added on 27 January 2022. For details, see comment No. 136 of the Change Log (Release Number 2022). | E05¦E0850
| Retinal Hemorrhage |  | E0851 | Blood extravasation in the retina. | MedDRA:10038870:Retinal hemorrhage | Eye | Vascular System |  | Term was added on 27 January 2022. For details, see comment No. 137 of the Change Log (Release Number 2022). | E05¦E0851
| Choroidal Hemorrhage |  | E0852 | Blood extravasation in the choroid. | MedDRA:10055239:Choroidal hemorrhage | Eye | Vascular System |  | Term was added on 27 January 2022. For details, see comment No. 138 of the Change Log (Release Number 2022). | E05¦E0852
| Optic Nerve Sheath Hemorrhage |  | E0853 | Blood extravasation in the Optic Nerve Sheath. | MedDRA:10054519:Optic nerve sheath hemorrhage | Eye | Vascular System |  | Term was added on 27 January 2022. For details, see comment No. 139 of the Change Log (Release Number 2022). | E05¦E0853
| Compartment syndrome |  | E1644 | Condition in which a closed muscle compartment's pressure increases to such an extent that the circulation is diminished, e.g. due to decrease in a compartment volume or an increase in the contents of a compartment, or both. | MedDRA:10010121:Compartment syndrome | Musculoskeletal System | Vascular System |  | Term added for Release Number 2024. See no. 74 of Change Log | E05¦E1644
| Hemarthrosis |  | E1645 | Bleeding into the joints | MedDRA:10019409:Hemarthrosis | Musculoskeletal System | Vascular System |  | Term added for Release Number 2024. See no. 75 of Change Log | E05¦E1645
| Clinical Endoleak |  | E2121 | The observed patient condition is persistent blood flow/pressure in the aneurysm sac following an endovascular aneurysm repair (EVAR) procedure. Where other patient signs and symptoms indicate the presence of an endoleak those signs and symptoms should be coded. | MedDRA:10063824:Endoleak | Procedural Complications | Vascular System |  | Term was added on 27 January 2022. For details, see comment No. 181 / 207 of the Change Log (Release Number 2022). Term and Definition modified for Release Number 2025. Previous term was "Endoleaks". See no. 117 of Change Request Log. | E05¦E2121
| Stenosis |  | E2337 | Narrowing or stricture of a vessel, duct or canal. | MedDRA:10076711:Stenosis | Generalized Disorders | Vascular System |  | Secondary category was added on 27 January 2022. For details, see comment No. 234 / 235 of the Change Log (Release Number 2022). Error in CodeHierarchy corrected for Release No. 2024 | E05¦E2337
Heart |  |  | E06 | Heart |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E06
| Coronary obstruction/occlusion |  | E0521 | Partial or total occlusion or obstruction of the coronary ostia or artery. | MedDRA:10011103:Coronary occlusion | Vascular System | Heart |  | Term added for Release Number 2024. See no. 82 of Change Log. Editorial change to term for Release Number 2026. See no. 260 of Release Log. | E06¦E0521
| Arrhythmia |  | E0601 | Any variation from the normal rate or rhythm (which may include the origin of the impulse and/or its subsequent propagation) in the heart. | MedDRA:10003119:Arrhythmia |  |  |  |  | E06¦E0601
|  | Asystole | E060101 | An electrocardiographic finding showing a state of cardiac standstill with no cardiac output and no ventricular depolarization. | MedDRA:10003586:Asystole |  |  |  |  | E06¦E0601¦E060101
|  | Atrial Fibrillation | E060102 | Irregular heart rhythm, usually associated with a rapid rate, that is caused by abnormal electrical activity within the atrial. | MedDRA:10003658:Atrial fibrillation |  |  |  |  | E06¦E0601¦E060102
|  | Atrial Flutter | E060103 | A rapid heart rate, associated with a regular rhythm, that is caused by abnormal electrical activity within the atria. | MedDRA:10003662:Atrial flutter |  |  |  |  | E06¦E0601¦E060103
|  | Bradycardia | E060104 | An abnormally slow heart rate. Thresholds for different age, gender, and patient populations exist. | MedDRA:10006093:Bradycardia |  |  |  |  | E06¦E0601¦E060104
|  | Ectopic Heartbeat | E060105 | A cardiac beat originating from an abnormal site. These changes lead to extra or dropped heartbeats. | MedDRA:10014159:Ectopic heartbeats |  |  |  |  | E06¦E0601¦E060105
|  | Heart Block | E060106 | An interference with the normal conduction of electrical impulses that control activity of the heart muscle. | MedDRA:10019252:Heart block |  |  |  |  | E06¦E0601¦E060106
|  | Idioventricular Rhythm | E060107 | An electrocardiographic finding of three or more consecutive complexes of ventricular origin. The QRS complexes are wide and have an abnormal morphology. (CDISC) | MedDRA:10039111:Rhythm idioventricular |  |  |  |  | E06¦E0601¦E060107
|  | Irregular Pulse | E060108 | A pulse that beats at irregular intervals. | MedDRA:10022994:Irregular pulse |  |  |  |  | E06¦E0601¦E060108
|  | Tachycardia | E060109 | Abnormally high heart rate. Thresholds for different age, gender, and patient populations exist. | MedDRA:10043071:Tachycardia |  |  |  |  | E06¦E0601¦E060109
|  | Ventricular Fibrillation | E060110 | Arrhythmia characterized by chaotic contractions of the ventricular muscle. | MedDRA:10047290:Ventricular fibrillation |  |  |  |  | E06¦E0601¦E060110
| Cardiac Arrest |  | E0602 | The sudden cessation of productive cardiac activity in an individual who becomes unresponsive, without normal breathing and no signs of circulation. | MedDRA:10007515:Cardiac arrest |  |  |  |  | E06¦E0602
| Cardiac Enzyme Elevation |  | E0603 | Increase from the normal levels of cardiac enzymes. | MedDRA:10007548:Cardiac enzymes increased | Heart | Investigations and Diagnostic Tests |  |  | E06¦E0603
| Cardiac Perforation |  | E0604 | The presence of an acquired hole in the heart. | MedDRA:10058039:Cardiac perforation | Heart | Procedural Complications |  |  | E06¦E0604
| Cardiac Tamponade |  | E0605 | Acute compression of the heart caused by increased intrapericardial pressure due to the accumulation of blood or fluid in the pericardium from rupture of the heart or great vessels, penetrating trauma, or progressive effusion. | MedDRA:10007610:Cardiac tamponade |  |  |  |  | E06¦E0605
| Cardiomyopathy |  | E0606 | A disease of the heart muscle. | MedDRA:10007636:Cardiomyopathy |  |  |  |  | E06¦E0606
| Cardiovascular Dysfunction/Insufficiency |  | E0607 | Circulatory dysfunction/insufficiency due to a decrease in the force of the cardiac contractions and / or the tone of the vascular walls. Includes diastolic and systolic dysfunction. | MedDRA:10065929:Cardiovascular insufficiency | Heart | Vascular System |  | Term and definition was modified on 31 January 2023 (Old term/definition: Cardiovascular Insufficiency: Circulatory insufficiency due to a decrease in the force of the cardiac contractions and / or the tone of the vascular walls.). For details, see comment No. 81 of the Change Log (Release Number 2023). Error in secondary CodeHierarchy corrected for Release No. 2024 | E06¦E0607
| Cusp Tear |  | E0608 | A tear located on a leaflet of a heart valve that prevents its complete closing, which can result in valvular dysfunction. | MedDRA:10068167:Cusp tear |  |  |  |  | E06¦E0608
| Diminished Pulse Pressure |  | E0609 | Weak or absent palpable pulse due to decreased arterial pulse pressure. | MedDRA:10037480:Pulse pressure decreased |  |  |  |  | E06¦E0609
| Endocarditis |  | E0610 | Inflammation of the endocardium. | MedDRA:10014665:Endocarditis | Heart | Infections |  |  | E06¦E0610
| Heart Failure/Congestive Heart Failure |  | E0611 | Inability of the heart to pump blood adequately to fill tissue metabolic requirements or the ability to do so only at an elevated filling pressure. | MedDRA:10019279:Heart failure |  |  |  |  | E06¦E0611
| Ischemic Heart Disease |  | E0612 | A disorder of cardiac function caused by insufficient blood flow to the muscle tissue of the heart. The decreased blood flow may be due to narrowing of the coronary arteries, to obstruction by a thrombus, or less commonly, to diffuse narrowing of arterioles and other small vessels within the heart. Severe interruption of the blood supply to the myocardial tissue may result in necrosis of cardiac muscle (myocardial infarction). | MedDRA:10055218:Ischemic heart disease |  |  |  |  | E06¦E0612
|  | Angina | E061201 | Chest pain resulting from inadequate oxygen delivery for the needs of the myocardium. | MedDRA:10002383:Angina pectoris |  |  |  |  | E06¦E0612¦E061201
|  | Myocardial Infarction | E061202 | A condition which occurs when blood flow decreases or stops to a part of the heart, causing damage to the heart muscle. | MedDRA:10028596:Myocardial infarction |  |  |  |  | E06¦E0612¦E061202
| Low Cardiac Output |  | E0613 | A below normal volume of blood pumped from the right or left ventricle of the heart per unit time. | MedDRA:10007595:Cardiac output decreased |  |  |  |  | E06¦E0613
| Mitral Perforation |  | E0614 | A rupture in the mitral valve tissue due to traumatic or pathologic processes. | MedDRA:10068138:Mitral perforation | Heart | Procedural Complications |  |  | E06¦E0614
| Myocardial Contusion |  | E0615 | A bruise to the heart. | MedDRA:10068147:Myocardial contusion | Heart | Procedural Complications |  |  | E06¦E0615
| Myocardial Hypertrophy |  | E0616 | Thickening of the myocardium often due to chronic pressure overload. | MedDRA:10072733:Myocardial hypertrophy |  |  |  |  | E06¦E0616
| Myocarditis |  | E0617 | Inflammation of the muscle tissue of the heart. | MedDRA:10028606:Myocarditis |  |  |  |  | E06¦E0617
| Non Specific EKG/ECG Changes |  | E0618 | An electrocardiographic finding of a change in cardiac electrical activity. | MedDRA:10061116:Electrocardiogram change |  |  | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 261 of Change Log. | E06¦E0618
| Pericardial Effusion |  | E0619 | Fluid collection within the pericardial sac, usually due to inflammation. | MedDRA:10034474:Pericardial effusion |  |  |  | Error in MedDRA mapping corrected for Release No. 2024 | E06¦E0619
| Pericarditis |  | E0620 | Inflammation of the pericardium. | MedDRA:10034484:Pericarditis |  |  |  |  | E06¦E0620
| Valvular Insufficiency/ Regurgitation |  | E0621 | Dysfunction of one of the cardiac valves, with incomplete valve closure or damage to a valve leaflet resulting in valvular backflow. | MedDRA:10058086:Heart valve insufficiency |  |  |  |  | E06¦E0621
|  | Aortic Valve Insufficiency/ Regurgitation | E062101 | Incomplete closure or damage to a valve leaflet of the aortic valve, resulting in aortic valve backflow. | MedDRA:10052839:Aortic valve insufficiency |  |  |  |  | E06¦E0621¦E062101
|  | Mitral Valve Insufficiency/ Regurgitation | E062102 | Incomplete closure or damage to a valve leaflet of the mitral valve, resulting in mitral valve backflow. | MedDRA:10074862:Mitral valve insufficiency |  |  |  |  | E06¦E0621¦E062102
|  | Pulmonary Valve Insufficiency/ Regurgitation | E062103 | Incomplete closure or damage to a valve leaflet of the pulmonary valve, resulting in pulmonary valve backflow. | MedDRA:10052842:Pulmonary valve insufficiency |  |  |  |  | E06¦E0621¦E062103
|  | Tricuspid Valve Insufficiency/ Regurgitation | E062104 | Incomplete closure or damage to a valve leaflet of the tricuspid valve, resulting in tricuspid valve backflow. | MedDRA:10044635:Tricuspid insufficiency |  |  |  |  | E06¦E0621¦E062104
| Valvular Stenosis |  | E0622 | Obstruction of flow across a heart valve due to immobility of the leaflets and/or a restricted orifice. | MedDRA:10061996:Heart valve stenosis |  |  |  |  | E06¦E0622
|  | Aortic Valve Stenosis | E062201 | Obstruction of flow across a aortic heart valve due to immobility reduced elasticity of the leaflets and/or a restricted orifice. | MedDRA:10002918:Aortic valve stenosis |  |  |  |  | E06¦E0622¦E062201
|  | Mitral Valve Stenosis | E062202 | Obstruction of flow across a mitral heart valve due to immobility reduced elasticity of the leaflets and/or a restricted orifice. | MedDRA:10027733:Mitral valve stenosis |  |  |  |  | E06¦E0622¦E062202
|  | Pulmonary Valve Stenosis | E062203 | Obstruction of flow across a pulmonary heart valve due to immobility reduced elasticity of the leaflets and/or a restricted orifice. | MedDRA:10037450:Pulmonary valve stenosis |  |  |  |  | E06¦E0622¦E062203
|  | Tricuspid Valve Stenosis | E062204 | Obstruction of flow across a tricuspid heart valve due to immobility reduced elasticity of the leaflets and/or a restricted orifice. | MedDRA:10044642:Tricuspid valve stenosis |  |  |  |  | E06¦E0622¦E062204
| Unspecified Heart Problem |  | E0623 | The report describes a non-specific problem with the heart. Note: Please use E2402 if the report describes a specific problem but the relevant term does not exist. |  |  |  | Modified (editorial) | Term was added on 18 March 2020 for cases where the category of a health effect was reported but with no further details. Editorial change to definition for Release Number 2026. See no. 201 of Release Log. | E06¦E0623
| Intraoperative Cardiac Valve Injury |  | E0624 | An intraoperative event that results in damage to, or abnormal function of, a heart valve. E.g. a tear in the leaflet or the chordae tendinae. | MedDRA:10088741:Intraoperative cardiac valve injury | Heart | Procedural Complications |  | Term was added on 31 January 2023. For details, see comment No. 80 / 114 of the Change Log (Release Number 2023). Definition modified for Release Number 2025. See no. 62 of Change Request Log | E06¦E0624
| Cardiac Septal Defect Residual Shunt |  | E0625 | Irregular blood flow in the heart after surgical correction of a septal defect. | MedDRA:10069121:Cardiac septal defect residual shunt |  |  |  | Term was added on 31 January 2023. For details, see comment No. 82 of the Change Log (Release Number 2023). | E06¦E0625
| Transvalvular Pressure Gradient Increased |  | E0626 | Increased pressure across the heart valve | MedDRA:10079128:Transvalvular pressure gradient increased |  |  |  | Term was added on 31 January 2023. For details, see comment No. 112 of the Change Log (Release Number 2023). | E06¦E0626
| Intraoperative Cardiac Injury |  | E0627 | An intraoperative event that results in damage to, or abnormal function of, the heart. For valvular injury, use E0624 - Intraoperative Cardiac Valve Injury | MedDRA:10065843:Intraoperative cardiac injury |  |  |  | Term was added on 31 January 2023. For details, see comment No. 118 of the Change Log (Release Number 2023). Definition modified for Release Number 2025. See no. 62 of Change Request Log. | E06¦E0627
| Fetal Distress |  | E1507 | Signs or symptoms which may be indicative of distress to a fetus, which may include repetitive variable decelerations, fetal tachycardia or bradycardia, late decelerations, or low biophysical profile. | MedDRA:10016486:Fetal distress | Pregnancy, Childbirth and the Puerperium | Heart |  |  | E06¦E1507
| Pneumopericardium |  | E0628 | A collection of air or gas in the pericardial cavity. | MedDRA:10048731:Pneumopericardium |  |  |  | Term added for Release Number 2025. See no. 32 of Change Request Log | E06¦E0628
| Palpitations |  | E0629 | The feeling of having a fast-beating, fluttering or pounding heart. | MedDRA:10033557:Palpitations |  |  |  | Term added for Release Number 2025. See no. 60 of Change Request Log | E06¦E0629
| Abnormal Precordial Findings |  | E0630 | Unusual physical findings on precordial examination, including palpable vibrations, thrills, or ventricular heaves. | MedDRA:10077162:Abnormal precordial movement |  |  | New | Term added for Release Number 2026. See no.15 of Change Request Log | E06¦E0630
Respiratory System |  |  | E07 | Respiratory System |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E07
| Pulmonary Embolism |  | E050303 | The obstruction of the pulmonary artery or one of its branches by an embolus, sometimes associated with infarction of the lung. | MedDRA:10037377:Pulmonary embolism | Vascular System | Respiratory System |  |  | E07¦E050303
| Adult Respiratory Distress Syndrome |  | E0701 | Progressive and life-threatening pulmonary distress in the absence of an underlying pulmonary condition, usually following major trauma or surgery. Cases of neonatal respiratory distress syndrome are not included in this definition. | MedDRA:10001409:Adult respiratory distress syndrome |  |  |  |  | E07¦E0701
| Airway Obstruction |  | E0702 | Blockage of air flow in the respiratory tract. | MedDRA:10001539:Airways obstruction |  |  |  |  | E07¦E0702
| Apnea |  | E0703 | Transient cessation of respiration. | MedDRA:10002972:Apnea |  |  |  |  | E07¦E0703
| Aspiration/Inhalation |  | E0704 | The entry of foreign material past the trachea, and into the lower respiratory tract. | MedDRA:10003504:Aspiration |  |  |  |  | E07¦E0704
| Aspiration Pneumonitis |  | E0705 | Inflammation of the lungs due to the inhalation of solid or liquid material. | MedDRA:10003526:Aspiration pneumonitis |  |  |  |  | E07¦E0705
| Asthma |  | E0706 | A chronic respiratory disease manifested as difficulty breathing due to the narrowing of bronchial passageways. | MedDRA:10003553:Asthma |  |  |  |  | E07¦E0706
| Bronchial Hemorrhage |  | E0707 | Bleeding from the bronchial wall and/or lung parenchyma. | MedDRA:10065757:Bronchial hemorrhage | Respiratory System | Vascular System |  |  | E07¦E0707
| Bronchitis |  | E0708 | An acute or chronic inflammatory process affecting the bronchi. | MedDRA:10006451:Bronchitis |  |  |  |  | E07¦E0708
| Bronchospasm |  | E0709 | Sudden contraction of the smooth muscles of the bronchial wall. | MedDRA:10006482:Bronchospasm |  |  |  |  | E07¦E0709
| Choking |  | E0710 | Interference with respiration by obstruction of the larynx or trachea. | MedDRA:10008589:Choking | Respiratory System | Injury |  | Editorial change to the definition on 16 February 2021. | E07¦E0710
| Chronic Obstructive Pulmonary Disease (COPD) |  | E0711 | A chronic and progressive lung disorder characterized by the loss of elasticity of the bronchial tree and the air sacs, destruction of the air sacs wall, thickening of the bronchial wall, and mucous accumulation in the bronchial tree. The pathologic changes result in the disruption of the air flow in the bronchial airways. Signs and symptoms include shortness of breath, wheezing, productive cough, and chest tightness. The two main types of chronic obstructive pulmonary disease are chronic obstructive bronchitis and emphysema. | MedDRA:10009033:Chronic obstructive pulmonary disease |  |  |  |  | E07¦E0711
|  | Pulmonary Emphysema | E071101 | A condition characterized by increase beyond normal in the size of air spaces distal to the terminal bronchioles, either from dilatation of the alveoli or from destruction of their walls. | MedDRA:10014563:Emphysema pulmonary |  |  |  |  | E07¦E0711¦E071101
| Cough |  | E0712 | A sudden, often repetitive, spasmodic contraction of the thoracic cavity, resulting in violent release of air from the lungs, and usually accompanied by a distinctive sound. | MedDRA:10011224:Cough |  |  |  |  | E07¦E0712
| Decreased Peak Expiratory Flow Rate |  | E0713 | Decrease in the maximum flow rate generated during a forceful exhalation, starting from full lung inflation. | MedDRA:10034195:Peak expiratory flow rate decreased | Respiratory System | Investigations and Diagnostic Tests | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 262 of Change Log. | E07¦E0713
| Increased Peak Expiratory Flow Rate |  | E0714 | An increase of the greatest rate of flow that can be achieved during forced expiration beginning with the lungs fully inflated. | MedDRA:10034196:Peak expiratory flow rate increased | Respiratory System | Investigations and Diagnostic Tests | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 263 of Change Log. | E07¦E0714
| Decreased Respiratory Rate |  | E0715 | Abnormal decrease of rate of breathing. | MedDRA:10038710:Respiratory rate decreased | Respiratory System | Investigations and Diagnostic Tests |  |  | E07¦E0715
| Increased Respiratory Rate |  | E0716 | Abnormal increase of rate of breathing. | MedDRA:10038712:Respiratory rate increased | Respiratory System | Investigations and Diagnostic Tests |  |  | E07¦E0716
| Dyspnea |  | E0717 | An uncomfortable sensation of difficulty breathing. It may present as an acute or chronic sign of an underlying respiratory or heart disorder. | MedDRA:10013963:Dyspnea |  |  |  |  | E07¦E0717
| Epistaxis |  | E0718 | Nose bleed. | MedDRA:10015090:Epistaxis |  |  |  |  | E07¦E0718
| Forced Expiratory Volume Decreased |  | E0719 | A decrease in the fraction of the forced vital capacity that is exhaled in a specific number of seconds. | MedDRA:10016987:Forced expiratory volume decreased | Respiratory System | Investigations and Diagnostic Tests |  |  | E07¦E0719
| Forced Expiratory Volume Increased |  | E0720 | An increase in the fraction of the forced vital capacity that is exhaled in a specific number of seconds. | MedDRA:10016989:Forced expiratory volume increased | Respiratory System | Investigations and Diagnostic Tests |  |  | E07¦E0720
| Hemoptysis |  | E0721 | Coughing up blood from the respiratory tract. | MedDRA:10019523:Hemoptysis |  |  |  |  | E07¦E0721
| Hemothorax |  | E0722 | A collection of blood in the pleural cavity. | MedDRA:10019614:Hemothorax |  |  |  |  | E07¦E0722
| Hiccups |  | E0723 | An involuntary contraction of the diaphragm against closed vocal cords, producing the hic sound. | MedDRA:10020039:Hiccups |  |  |  |  | E07¦E0723
| Hyperventilation |  | E0724 | Abnormally prolonged, rapid, and deep breathing. | MedDRA:10020910:Hyperventilation |  |  |  |  | E07¦E0724
| Hypoventilation |  | E0725 | A state in which there is a reduced amount of air entering the pulmonary alveoli. | MedDRA:10021133:Hypoventilation |  |  |  |  | E07¦E0725
| Hypoxia |  | E0726 | A deficiency in the amount of oxygen reaching the tissues. | MedDRA:10021143:Hypoxia |  |  |  |  | E07¦E0726
| Mediastinal Shift |  | E0727 | Displacement of structures in the thoracic cavity in one direction due to increased pleural pressure on the contralateral side. | MedDRA:10066239:Mediastinal shift |  |  |  |  | E07¦E0727
| Nasal Obstruction |  | E0728 | Blockage of the nasal passages. | MedDRA:10028748:Nasal obstruction |  |  |  |  | E07¦E0728
| Overinflation of Lung |  | E0729 | Abnormal permanent enlargement of the lung air spaces distal to terminal bronchiole not resulted from wall destruction, e.g. due to loss of opposite lung. | MedDRA:10059487:Lung hyperinflation |  |  |  |  | E07¦E0729
| Pharyngitis |  | E0730 | Inflammation of the throat most often caused by viral and bacterial infections. | MedDRA:10034835:Pharyngitis |  |  |  |  | E07¦E0730
| Pleural Effusion |  | E0731 | Increased amounts of fluid within the pleural cavity. Symptoms include shortness of breath, cough, and chest pain. | MedDRA:10035598:Pleural effusion |  |  |  |  | E07¦E0731
| Pleural Empyema |  | E0732 | An accumulation of pus in the pleural space | MedDRA:10090758:Pleural empyema |  |  |  |  | E07¦E0732
| Pneumonia |  | E0733 | An acute and/or chronic inflammation focally or diffusely affecting the lung parenchyma. | MedDRA:10035664:Pneumonia | Respiratory System | Infections |  |  | E07¦E0733
|  | Bronchopneumonia | E073301 | Acute inflammation of the walls of the terminal bronchioles. | MedDRA:10006469:Bronchopneumonia | Respiratory System | Infections |  |  | E07¦E0733¦E073301
| Pneumothorax |  | E0734 | Abnormal presence of air in the pleural cavity. | MedDRA:10035759:Pneumothorax |  |  |  |  | E07¦E0734
| Pulmonary Dysfunction |  | E0735 | Impaired or abnormal functioning of the pulmonary system causing multiple pulmonary issues. | MedDRA:10037386:Pulmonary function impairment |  |  |  |  | E07¦E0735
| Pulmonary Edema |  | E0736 | Accumulation of fluid in the lung tissues causing disturbance of the gas exchange that may lead to respiratory failure. | MedDRA:10037375:Pulmonary edema | Respiratory System | Generalized Disorders |  |  | E07¦E0736
| Pulmonary Hypertension |  | E0737 | Increased pressure within the pulmonary circulation due to lung or heart disorder. | MedDRA:10037400:Pulmonary hypertension |  |  |  |  | E07¦E0737
| Pulmonary Infarction |  | E0738 | Localized necrosis of lung tissue caused by obstruction of the arterial blood supply, most often due to pulmonary embolism. | MedDRA:10037410:Pulmonary infarction |  |  |  |  | E07¦E0738
| Respiratory Acidosis |  | E0739 | A condition in which the blood pH is less than normal, secondary to impaired gas exchange. | MedDRA:10038661:Respiratory acidosis | Respiratory System | Endocrine, Metabolism and Nutrition |  |  | E07¦E0739
| Respiratory Alkalosis |  | E0740 | A condition in which the blood pH is greater than normal, secondary to impaired gas exchange. | MedDRA:10038664:Respiratory alkalosis | Respiratory System | Endocrine, Metabolism and Nutrition |  |  | E07¦E0740
| Respiratory Arrest |  | E0741 | Cessation of breathing function. | MedDRA:10038669:Respiratory arrest |  |  |  |  | E07¦E0741
| Respiratory Failure |  | E0742 | The significant impairment of gas exchange within the lungs resulting in hypoxia, hypercarbia, or both, to the extent that organ tissue perfusion is severely compromised. | MedDRA:10038695:Respiratory failure |  |  |  |  | E07¦E0742
| Respiratory Insufficiency |  | E0743 | Impairment of gas exchange within the lungs secondary to a disease process, neoplasm, or trauma, possibly resulting in hypoxia, hypercarbia, or both. | MedDRA:10038701:Respiratory insufficiency |  |  |  |  | E07¦E0743
| Respiratory Tract Infection |  | E0744 | An infectious disease affecting the respiratory tract. | MedDRA:10062352:Respiratory tract infection | Respiratory System | Infections |  |  | E07¦E0744
| Sinus Perforation |  | E0745 | A rupture in the sinus tissue due to traumatic or pathologic processes. | MedDRA:10040748:Sinus perforation | Respiratory System | Procedural Complications |  |  | E07¦E0745
| Sneezing |  | E0746 | A symptom consisting of the involuntary expulsion of air from the nose. | MedDRA:10041232:Sneezing |  |  |  |  | E07¦E0746
| Sore Throat |  | E0747 | Any kind of inflammatory process of the tonsils, pharynx, or/and larynx characterized by pain in swallowing. | MedDRA:10041367:Sore throat |  |  |  |  | E07¦E0747
| Stacking Breaths |  | E0748 | Condition where there are a greater number of inhalations relative to the number of exhalations. | MedDRA:10068523:Breath stacking |  |  |  |  | E07¦E0748
| Strangulation |  | E0749 | Interference with respiration by compression of the larynx, laryngopharynx or trachea, as well as of the carotid arteria or jugular veins. | MedDRA:10042169:Strangulation | Respiratory System | Injury |  | Editorial change to the definition on 16 February 2021. | E07¦E0749
| Ventilator Dependent |  | E0750 | A condition that requires mechanical life support due to inability to breath effectively to maintain normal levels of oxygen and carbon dioxide in the body. | MedDRA:10080927:Dependence on ventilator |  |  |  |  | E07¦E0750
| Wheezing |  | E0751 | A symptom and finding during physical examination, characterized by high-pitched, whistling sound during breathing. It results from narrowing or obstruction of the respiratory airways. Also known as stridor. | MedDRA:10047924:Wheezing |  |  |  | Definition modified for Release Number 2024. See no. 60 of Change Log | E07¦E0751
| Unspecified Respiratory Problem |  | E0752 | The report describes a non-specific problem with the respiratory system. Note: Please use E2402 if the report describes a specific problem but the relevant term does not exist. |  |  |  | Modified (editorial) | Term was added on 18 March 2020 for cases where the category of a health effect was reported but with no further details. Editorial change to definition for Release Number 2026. See no. 202 of Release Log. | E07¦E0752
| Upper Respiratory Tract Injury |  | E0753 | Mechanical injuries to the upper and/or lower airways. For chemical injuries or exposure use E170401 - Caustic/Chemical Burns or E2003 - Chemical Exposure |  |  |  | Retired (2025) | Term was added on 27 January 2022. For details, see comment No. 103 of the Change Log (Release Number 2022). Term retired for Release Number 2025. For mechanical injury use E0759 Mechanical Respiratory Tract Injury. For chemical injuries or exposure use E170401 - Caustic/Chemical Burns or E2003 - Chemical Exposure. See no. 86 of Change Request Log | E07¦E0753
| Upper Respiratory Tract Irritation |  | E0754 | Irritation and swelling of the upper airways | MedDRA:10070841:Upper respiratory tract irritation |  |  |  | Term added for Release Number 2024. See no. 1 of Change Log | E07¦E0754
| Snoring |  | E0755 | A particular type of noisy sound made by air passage across relaxed tissues during sleep. | MedDRA:10041235:Snoring |  |  |  | Term added for Release Number 2024. See no. 2 of Change Log | E07¦E0755
| Nasal Congestion |  | E0756 | The blockage of the nasal passage due to mucosal edema. It is usually the result of a viral infection or an allergic reaction. | MedDRA:10028735:Nasal congestion |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 21 of Change Log. Editorial change to term for Release Number 2026. See no. 264 of Change Log. | E07¦E0756
| Breathing-Related Sleep Disorder |  | E0757 | Clinical spectrum with frequent episodes of either complete or partial obstruction of the upper airway during sleep. Includes sleep apnea. Also known as Sleep disordered breathing (SDB). | MedDRA:10006344:Breathing-related sleep disorder |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 61 of Change Log. Editorial change to term for Release Number 2026. See no. 265 of Change Log. | E07¦E0757
| Rhinitis |  | E0758 | Pattern of symptoms as a result of nasal inflammation and/or dysfunction of the nasal mucosa. There are various subtypes, such as allergic, infectious and non-allergic (e.g. atrophic) rhinitis. | MedDRA:10039083:Rhinitis |  |  |  | Term added for Release Number 2024. See no. 62 of Change Log | E07¦E0758
| Mechanical Respiratory Tract Injury |  | E0759 | Mechanical injuries to the upper and/or lower airways. For chemical injuries or exposure use E170401 - Caustic/Chemical Burns or E2003 - Chemical Exposure | MedDRA:10081319:Respiratory tract procedural complication |  |  |  | Term added for Release Number 2025. See no. 86 of Change Request Log. | E07¦E0759
| Atelectasis |  | E0760 | Incomplete expansion of a lung or a portion of a lung; or portion of the lung collapses or becomes airless, often due to a blockage in the airway. | MedDRA:10003598:Atelectasis |  |  |  | Term added for Release Number 2025. See no. 88 of Change Request Log | E07¦E0760
| Nasal Dryness |  | E0761 | Lack of moisture in the nasal mucosa | MedDRA:10028740:Nasal dryness |  |  |  | Term added for Release Number 2025. See no. 101 of Change Request Log | E07¦E0761
| Airway Burns |  | E0762 | Burn to upper respiratory tract from inhalation of heated gases. | MedDRA:10067204:Airway burns | Respiratory System | Injury | New | Term added for Release Number 2026. See no. 68 of Change Request Log | E07¦E0762
| Respiratory Distress Syndrome of Newborns |  | E1516 | A condition of the newborn marked by dyspnea with cyanosis. | MedDRA:10038692:Respiratory distress syndrome in newborn | Pregnancy, Childbirth and the Puerperium | Respiratory System |  |  | E07¦E1516
| High Oxygen Saturation |  | E2202 | A high level of the degree to which oxygen is bound to hemoglobin given as a percentage calculated by dividing the maximum oxygen capacity into the actual oxygen content and multiplying by 100. Oxygen saturation usually is measured using pulse oximetry. | MedDRA:10033319:Oxygen saturation high | Investigations and Diagnostic Tests | Respiratory System |  |  | E07¦E2202
| Low Oxygen Saturation |  | E2203 | A low level of the degree to which oxygen is bound to hemoglobin given as a percentage calculated by dividing the maximum oxygen capacity into the actual oxygen content and multiplying by 100. Oxygen saturation usually is measured using pulse oximetry. | MedDRA:10033321:Oxygen saturation low | Investigations and Diagnostic Tests | Respiratory System |  |  | E07¦E2203
Eye |  |  | E08 | Eye |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E08
| Capsular Bag Tear |  | E0801 | A tear in the capsular bag housing the lens. | MedDRA:10063941:Posterior capsule tear |  |  |  |  | E08¦E0801
| Cataract |  | E0802 | Partial or complete opacity of the crystalline lens of one or both eyes that decreases visual acuity and eventually results in blindness. Some cataracts appear in infancy or in childhood, but most develop in older individuals. (Sternberg Diagnostic Surgical Pathology, 3rd ed.) | MedDRA:10007739:Cataract |  |  |  |  | E08¦E0802
| Chemosis |  | E0803 | Edema of the mucous membrane of the eyeball and eyelid lining. | MedDRA:10008435:Chemosis | Eye | Infections |  |  | E08¦E0803
| Conjunctivitis |  | E0804 | Inflammation of the conjunctiva of the eye. | MedDRA:10010741:Conjunctivitis |  |  |  |  | E08¦E0804
| Corneal Clouding/Hazing |  | E0805 | Clouding of the middle area of the cornea, usually as a result of scarring due to infection, injury, or inflammation. | MedDRA:10052116:Corneal clouding |  |  |  |  | E08¦E0805
| Corneal Decompensation |  | E0806 | Inability to maintain corneal integrity; or corneal edema resulting from failure of the corneal endothelium to maintain detumescence. | MedDRA:10052117:Corneal decompensation |  |  |  |  | E08¦E0806
| Corneal Edema |  | E0807 | Hazy, swollen cornea. | MedDRA:10011007:Corneal edema | Eye | Generalized Disorders |  |  | E08¦E0807
|  | Corneal Stromal Edema | E080701 | Fluid from the aqueous humor enters the corneal stroma and causes swelling. | MedDRA:10054772:Corneal stromal edema | Eye | Generalized Disorders |  |  | E08¦E0807¦E080701
|  | Microcystic Edema | E080702 | Corneal edema of inflammatory origin associated with contact lens wear and morphologically characterized by presence of epithelial microcysts. | MedDRA:10068529:Corneal microcystic edema | Eye | Generalized Disorders |  |  | E08¦E0807¦E080702
| Corneal Epithelial Microcysts |  | E0808 | Very small, round vesicles containing fluid and cellular debris observed on the surface of the cornea under slit lamp examination in some types of corneal dystrophy and in wearers of extended wear lenses. | MedDRA:10068144:Corneal epithelial microcysts |  |  |  |  | E08¦E0808
| Corneal Infiltrates |  | E0809 | Discrete, small lesions present in the cornea as a result of corneal inflammation and, in some cases, after soft contact lens wear especially extended-wear lenses. | MedDRA:10011022:Corneal infiltrates |  |  |  |  | E08¦E0809
| Corneal Pannus |  | E0810 | Superficial vascularization of the cornea with infiltration of granulation tissue. | MedDRA:10033682:Pannus (corneal) |  |  |  |  | E08¦E0810
| Corneal Perforation |  | E0811 | A rupture in the corneal tissue due to traumatic or pathologic processes. | MedDRA:10011039:Corneal perforation | Eye | Procedural Complications |  |  | E08¦E0811
| Corneal Scar |  | E0812 | Replacement of corneal tissue with scar tissue as a result of injury to the deeper layers of the cornea. | MedDRA:10011044:Corneal scar |  |  |  |  | E08¦E0812
| Corneal Touch |  | E0813 | Complications of intraocular lens surgery; usually refers to intraocular lens contact with cornea, can be intermittent or chronic. | MedDRA:10011055:Corneal touch |  |  |  |  | E08¦E0813
| Corneal Ulcer |  | E0814 | Area of epithelial tissue loss from corneal surface; associated with inflammatory cells in the cornea and anterior chamber. | MedDRA:10048492:Corneal ulcer |  |  |  |  | E08¦E0814
| Dry Eye(s) |  | E0815 | A syndrome characterized by dryness of the cornea and conjunctiva. It is usually caused by a deficiency in tear production. Symptoms include a feeling of burning eyes and a possible foreign body presence in the eye. | MedDRA:10013774:Dry eye |  |  |  |  | E08¦E0815
| Endophthalmitis |  | E0816 | Inflammation of the intraocular fluids or tissues usually with microbial organisms. | MedDRA:10014801:Endophthalmitis | Eye | Infections |  |  | E08¦E0816
| Excessive Tear Production |  | E0817 | Profuse lacrimation. | MedDRA:10015591:Excess tears |  |  |  |  | E08¦E0817
| Eye Infections |  | E0818 | An infection of any part of the eye, including the eyelid. | MedDRA:10015929:Eye infection | Eye | Infections |  |  | E08¦E0818
|  | Intraocular Infection | E081801 | Infection within the eye. | MedDRA:10054762:Eye infection intraocular | Eye | Infections |  |  | E08¦E0818¦E081801
| Eye Injury |  | E0819 | Damage to the eye. | MedDRA:10061128:Eye injury | Eye | Injury |  |  | E08¦E0819
|  | Corneal Abrasion | E081901 | Injury to the epithelium of the cornea. | MedDRA:10010984:Corneal abrasion |  |  |  |  | E08¦E0819¦E081901
|  | Eye Burn | E081902 | Burn injury, including chemical burn, to any part of the eye. | MedDRA:10015911:Eye burns | Eye | Injury |  |  | E08¦E0819¦E081902
|  | Retinal Detachment | E081903 | Separation of the inner retina layers from the underlying pigment epithelium. | MedDRA:10038848:Retinal detachment |  |  |  |  | E08¦E0819¦E081903
|  | Retinal Injury | E081904 | Damage to the retina due to toxicity, exposure to bright lights, or trauma. | MedDRA:10057430:Retinal injury | Eye | Injury |  |  | E08¦E0819¦E081904
|  | Retinal Tear | E081905 | A usually small tearing of the retina occurring when the vitreous separates from the retina. It may lead to retinal detachment. | MedDRA:10038897:Retinal tear |  |  |  |  | E08¦E0819¦E081905
|  | Vitreous Detachment | E081906 | Detachment of the vitreous humor from the retina. | MedDRA:10047650:Vitreous detachment | Eye | Procedural Complications |  |  | E08¦E0819¦E081906
| Eye Pain |  | E0820 | Painful sensation in the eye. | MedDRA:10015958:Eye pain | Eye | Generalized Disorders |  |  | E08¦E0820
| Foreign Body Sensation in Eye |  | E0821 | Feeling of grittiness or having something in the eye; frequently caused by a foreign body. | MedDRA:10051116:Foreign body sensation in eyes |  |  |  |  | E08¦E0821
| Glaucoma |  | E0822 | Increased pressure in the eyeball due to obstruction of the outflow of aqueous humor. | MedDRA:10018304:Glaucoma |  |  |  |  | E08¦E0822
| Halo |  | E0823 | Hazy ring around bright lights seen by some patients with refractive error or optical defects, e.g. cataracts, or corneal swelling. | MedDRA:10019099:Halo vision |  |  |  |  | E08¦E0823
| Hyphema |  | E0824 | Hemorrhage within the anterior chamber of the eye. | MedDRA:10020924:Hyphema |  |  |  | Definition was modified on 31 January 2023 to remove "bloodshot". For details, see comment No. 139 of the Change Log (Release Number 2023). | E08¦E0824
| Hypopyon |  | E0825 | An accumulation of pus in the anterior chamber of the eye. | MedDRA:10021086:Hypopyon |  |  |  |  | E08¦E0825
| Intraocular Pressure Decreased |  | E0826 | Decrease of pressure of the intraocular fluid in the eye. | MedDRA:10022804:Intraocular pressure decreased |  |  |  |  | E08¦E0826
| Intraocular Pressure Increased |  | E0827 | Increase of pressure of the intraocular fluid in the eye, also known as ocular hypertension. | MedDRA:10022806:Intraocular pressure increased |  |  |  | Definition was modified on 31 January 2023 (Old definition: Increase of pressure of the intraocular fluid in the eye.). For details, see comment No. 108 of the Change Log (Release Number 2023). | E08¦E0827
| Iritis |  | E0828 | Inflammation of the iris. | MedDRA:10022955:Iritis |  |  |  |  | E08¦E0828
| Keratitis |  | E0829 | Inflammation of the cornea. | MedDRA:10023332:Keratitis |  |  |  |  | E08¦E0829
|  | Infective Keratitis | E082901 | Keratitis due to infection, including infection by acanthamoeba. | MedDRA:10076672:Infective keratitis |  |  |  | Term/Definition was modified on 27 January 2022 (Old term/definition: Acanthamoeba Keratitis: Keratitis due to infection by acanthamoeba.) MedDRA mapping also updated from Acanthameba keratitis(10074137) to Infective keratitis(10076672). For details, see comment No. 252 of the Change Log (Release Number 2022). | E08¦E0829¦E082901
| Macular Edema |  | E0830 | Accumulation of intraretinal fluid and protein in the macula, which may result in swelling and decreased central vision. | MedDRA:10054467:Macular edema | Eye | Generalized Disorders |  |  | E08¦E0830
| Overwear Syndrome |  | E0831 | Ocular pain which may be very intense, accompanied by corneal epithelium damage, conjunctival injection, lacrimation, blepharospasm, photophobia and hazy vision following corneal edema caused by overwear of contact lenses, principally the PMMA type. | MedDRA:10080905:Overwear syndrome |  |  |  |  | E08¦E0831
| Pupillary Block |  | E0832 | An obstruction of the pupil. | MedDRA:10037520:Pupillary block |  |  |  |  | E08¦E0832
| Red Eye(s) |  | E0833 | Lay term applied to any condition with dilation of conjunctival or ciliary blood vessels; innumerable causes. | MedDRA:10038189:Red eye |  |  |  |  | E08¦E0833
| Retinal Degeneration |  | E0834 | Degeneration of the retina. | MedDRA:10038845:Retinal degeneration |  |  |  |  | E08¦E0834
| Toxic Anterior Segment Syndrome (TASS) |  | E0835 | An acute postoperative inflammatory reaction in which a noninfectious substance enters the anterior segment and induces toxic damage to the intraocular tissues. | MedDRA:10066366:Toxic anterior segment syndrome |  |  |  | Editorial change to the definition on 16 February 2021. | E08¦E0835
| UGH (Uveitis Glaucoma Hyphema) Syndrome |  | E0836 | Clinical triad of uveitis, glaucoma, and hyphema (UGH), especially associated with the insertion of an anterior chamber intraocular lens (IOL). | MedDRA:10068156:Uveitis-glaucoma-hyphema syndrome |  |  |  |  | E08¦E0836
| Uveitis |  | E0837 | Inflammation of the uvea. | MedDRA:10046851:Uveitis |  |  |  |  | E08¦E0837
| Visual Disturbances |  | E0838 | An interference to normal eyesight without necessarily affecting visual acuity. | MedDRA:10047545:Visual disturbances |  |  |  |  | E08¦E0838
|  | Flashers | E083801 | The perception of sudden or brief bursts of light. | MedDRA:10047569:Visual flashes |  |  |  |  | E08¦E0838¦E083801
|  | Phosphene Visualization | E083802 | An objective visual sensation that appears with the eyes closed and in the absence of visual light. | MedDRA:10034936:Phosphenes | Eye | Generalized Disorders |  |  | E08¦E0838¦E083802
|  | Vitreous Floaters | E083803 | Spots or stringy shapes in vision, often caused by deposits, aggregates, or cell fragments in the vitreous humor or lens. | MedDRA:10047654:Vitreous floaters |  |  |  |  | E08¦E0838¦E083803
|  | Photophobia | E083804 | Increased sensitivity of the eyes to light causing discomfort or pain. | MedDRA:10034960:Photophobia |  |  |  | Term was added on 27 January 2022. For details, see comment No. 251 of the Change Log (Release Number 2022). | E08¦E0838¦E083804
| Visual Impairment |  | E0839 | Decreased ability to see to a degree that requires corrective measures | MedDRA:10047571:Visual impairment |  |  |  |  | E08¦E0839
|  | Blurred Vision | E083901 | The loss of visual acuity (sharpness of vision). | MedDRA:10005886:Blurred vision |  |  |  |  | E08¦E0839¦E083901
|  | Loss of Vision | E083902 | Severely reduced visual acuity that cannot be corrected, or total loss of vision. | MedDRA:10024881:Loss of vision |  |  |  |  | E08¦E0839¦E083902
|  | Diplopia/Double Vision | E083903 | Perception of multiple images from one object. | MedDRA:10013036:Diplopia |  |  |  | Term was added on 27 January 2022. For details, see comment No. 130 of the Change Log (Release Number 2022). | E08¦E0839¦E083903
| Vitreous Hemorrhage |  | E0840 | Blood extravasation in the vitreous humor. | MedDRA:10047656:Vitreous hemorrhage |  |  |  |  | E08¦E0840
| Vitreous Loss |  | E0841 | loss or prolapse of vitreous gel. | MedDRA:10063942:Vitreous loss |  |  |  |  | E08¦E0841
| Vitritis |  | E0842 | Inflammatory intraocular reaction with clouding and cells in vitreous; often accompanies inflammation of ciliary body, iris, choroid, or retina. | MedDRA:10047663:Vitritis |  |  |  |  | E08¦E0842
| Zonular Dehiscence |  | E0843 | Rupture of the fibrous strands connecting the ciliary body with the crystalline lens of the eye (usually at the time of cataract surgery). | MedDRA:10068351:Ciliary zonular dehiscence |  |  |  | Editorial change to the definition on 16 February 2021. | E08¦E0843
| Unspecified Eye / Vision Problem |  | E0844 | The report describes a non-specific problem with the eye(s) or vision. Note: Please use E2402 if the report describes a specific problem but the relevant term does not exist. |  |  |  | Modified (editorial) | Term was added on 18 March 2020 for cases where the category of a health effect was reported but with no further details. Editorial change to definition for Release Number 2026. See no. 203 of Release Log. | E08¦E0844
| Eye Irritation |  | E0845 | A generic term used to describe abnormal sensations that bother the eyes like dryness, itchiness, burning, grittiness and excessive blinking. | MedDRA:10015946:Eye irritation |  |  |  | Definition was modified on 31 January 2023 to add "abnormal". For details, see comment No. 103 of the Change Log (Release Number 2023). Definition modified for Release Number 2025. See no. 102 of Change Request Log | E08¦E0845
| Corneal Thinning |  | E0846 | Progressive deformation of the corneal structure, characterized by thinning of the cornea. | MedDRA:10054760:Corneal thinning |  |  |  | Term was added on 27 January 2022. For details, see comment No. 126 of the Change Log (Release Number 2022). | E08¦E0846
| Pigment Dispersion Syndrome |  | E0847 | Condition where iris pigment is released into the anterior chamber and deposited in the trabecular meshwork, due to a procedure or device implantation. | MedDRA:10052134:Pigment dispersion syndrome | Eye | Procedural Complications |  | Term was added on 27 January 2022. For details, see comment No. 127 of the Change Log (Release Number 2022). | E08¦E0847
| Corneal Endothelial Cell Loss |  | E0848 | Decrease in corneal endothelial cell density. | MedDRA:10071681:Corneal endothelial cell loss | Eye | Procedural Complications |  | Term was added on 27 January 2022. For details, see comment No. 129 of the Change Log (Release Number 2022). | E08¦E0848
| Posterior Capsule Opacification |  | E0849 | Remnant lens epithelial cell proliferation and migration to the posterior capsule. | MedDRA:10036346:Posterior capsule opacification | Eye | Procedural Complications |  | Term was added on 27 January 2022. For details, see comment No. 131 of the Change Log (Release Number 2022). | E08¦E0849
| Conjunctival Hemorrhage |  | E0850 | Blood extravasation in the conjunctiva. | MedDRA:10010720:Conjunctival hemorrhage | Eye | Vascular System |  | Term was added on 27 January 2022. For details, see comment No. 136 of the Change Log (Release Number 2022). | E08¦E0850
| Retinal Hemorrhage |  | E0851 | Blood extravasation in the retina. | MedDRA:10038870:Retinal hemorrhage | Eye | Vascular System |  | Term was added on 27 January 2022. For details, see comment No. 137 of the Change Log (Release Number 2022). | E08¦E0851
| Choroidal Hemorrhage |  | E0852 | Blood extravasation in the choroid. | MedDRA:10055239:Choroidal hemorrhage | Eye | Vascular System |  | Term was added on 27 January 2022. For details, see comment No. 138 of the Change Log (Release Number 2022). | E08¦E0852
| Optic Nerve Sheath Hemorrhage |  | E0853 | Blood extravasation in the Optic Nerve Sheath. | MedDRA:10054519:Optic nerve sheath hemorrhage | Eye | Vascular System |  | Term was added on 27 January 2022. For details, see comment No. 139 of the Change Log (Release Number 2022). | E08¦E0853
| Mydriasis |  | E0854 | Dilated pupils | MedDRA:10028521:Mydriasis |  |  |  | Term was added on 31 January 2023. For details, see comment No. 107 of the Change Log (Release Number 2023). | E08¦E0854
| Anterior Capsule Opacification |  | E0855 | Remnant lens epithelial cell proliferation of the anterior capsule. | MedDRA:10088740:Anterior capsule opacification |  |  |  | Term was added on 31 January 2023. For details, see comment No. 119 / 121 of the Change Log (Release Number 2023). | E08¦E0855
| Pupillary Disorder |  | E0856 | Condition which affects the structures or function of the pupil of the eye. | MedDRA:10037521:Pupillary disorder |  |  |  | Term was added on 31 January 2023. For details, see comment No. 120 of the Change Log (Release Number 2023). | E08¦E0856
| Corneal Warpage |  | E0857 | A change in the shape of the corneal surface presenting as abnormal radial symmetry. | MedDRA:10075404:Corneal warpage |  |  |  | Term was added on 31 January 2023. For details, see comment No. 129 of the Change Log (Release Number 2023). | E08¦E0857
| Choroidal Detachment |  | E0858 | A detachment of the choroid from the underlying sclera often due to the accumulation of fluid (blood for hemorrhagic, other fluids for serous detachment) in the suprachoroidal space | MedDRA:10008783:Choroidal detachment |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 7 of Change Log. Editorial change to term for Release Number 2026. See no. 266 of Change Log. | E08¦E0858
| Choroidal Effusion |  | E0859 | Fluid accumulation in the suprachoroidal space, can lead to detachment | MedDRA:10063970:Choroidal effusion |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 8 of Change Log. Editorial change to term for Release Number 2026. See no. 267 of Change Log. | E08¦E0859
| Eye Strain/Fatigue |  | E0860 | A feeling of general eye weakness/fatigue. (e.g. asthenopia). | MedDRA:10049755:Eye Strain |  |  |  | Term added for Release Number 2025. See no. 100 of Change Request Log | E08¦E0860
| Retinal Vein Occlusion |  | E0861 | Retinal vein occlusion | MedDRA:10038907:Retinal vein occlusion |  |  | New | Term added for Release Number 2026. See no. 87 of Change Request Log | E08¦E0861
| Maculopathy |  | E0862 | Any pathological condition of the macula of the eye. | MedDRA:10025425:Maculopathy |  |  | New | Term added for Release Number 2026. See no. 77 of Change Request Log | E08¦E0862
| Ocular Procedural Complication |  | E2135 | Any complication arising from ocular surgical procedures involving the conjunctiva or adjacent tissues, including but not limited to subconjunctival bleb formation (e.g., trabeculectomy, glaucoma filtration procedures). This may encompass issues such as bleb absence, infection, leakage, perforation, structural thinning, scarring, or other adverse changes affecting the surgical site or surrounding ocular structures. | MedDRA:10080964:Ocular procedural complication | Procedural Complications | Eye | New | Term added for Release Number 2026. See no. 381 of Change Log. | E08¦E2135
Ear and Labyrinth |  |  | E09 | Ear and Labyrinth |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E09
| Acoustic Shock |  | E0901 | Neurophysiological and psychological symptoms caused by sudden and unexpected loud noise. | MedDRA:10068150:Acoustic shock | Ear and Labyrinth | Nervous System |  |  | E09¦E0901
| Acoustic Trauma |  | E0902 | Injury to the hearing mechanisms within the inner ear, caused by excessively loud noise. | MedDRA:10000528:Acoustic trauma |  |  |  |  | E09¦E0902
| Hearing Impairment |  | E0903 | Partial or complete loss of the ability to detect or understand sounds resulting from damage to the outer, middle, or inner ear structures. | MedDRA:10019245:Hearing impaired |  |  |  |  | E09¦E0903
|  | Partial Hearing Loss | E090301 | A condition in which a person partially loses the ability to hear sounds in one or both ears. | MedDRA:10050564:Partial hearing loss |  |  |  |  | E09¦E0903¦E090301
|  | Total Hearing Loss | E090302 | A condition in which a person completely loses the hearing in one or both ears. | MedDRA:10074290:Complete hearing loss |  |  |  |  | E09¦E0903¦E090302
| Tinnitus |  | E0904 | A noise in the ears, such as ringing, buzzing, roaring or clicking in the absence of auditory stimulation. | MedDRA:10043882:Tinnitus | Ear and Labyrinth | Nervous System |  |  | E09¦E0904
| Vertigo |  | E0905 | A feeling of movement, a sensation as if the external world were revolving around the person or as if they were revolving in space (subjective vertigo). Vertigo is medically distinct from dizziness, lightheadedness, and unsteadiness. | MedDRA:10047340:Vertigo |  |  |  |  | E09¦E0905
| Unspecified Ear or Labyrinth Problem |  | E0906 | The report describes a non-specific problem with the ear or labyrinth. Note: Please use E2402 if the report describes a specific problem but the relevant term does not exist. |  |  |  | Modified (editorial) | Term was added on 18 March 2020 for cases where the category of a health effect was reported but with no further details. Editorial change to definition for Release Number 2026. See no. 204 of Release Log. | E09¦E0906
| Ear Pain |  | E0907 | Earache or ear pain. | MedDRA:10014020:Ear pain | Ear and Labyrinth | Generalized Disorders |  | Term was added on 27 January 2022. For details, see comment No. 234 / 235 of the Change Log (Release Number 2022). | E09¦E0907
| Ear Infection |  | E0908 | Infection of the ear. | MedDRA:10014011:Ear infection | Ear and Labyrinth | Infections | Modified (editorial) | Term was added on 27 January 2022. For details, see comment No. 223 of the Change Log (Release Number 2022). Editorial change to term for Release Number 2026. See no. 268 of Release Log. | E09¦E0908
|  | Infectious Otitis Externa | E090801 | An infectious process involving the outer ear and ear canal. | MedDRA:10033072:Otitis externa | Ear and Labyrinth | Infections |  | Editorial change to the term on 31 January 2023. | E09¦E0908¦E090801
|  | Infectious Otitis Media | E090802 | An infectious process involving the middle ear, includes acute otitis media, chronic otitis media, chronic suppurative otitis media, otitis media with effusion (OME), etc. | MedDRA:10033078:Otitis media | Ear and Labyrinth | Infections |  | Editorial change to the term on 31 January 2023. | E09¦E0908¦E090802
| Middle Ear Inflammation |  | E0909 | Inflammation (swelling and redness) to the middle ear. | MedDRA:10065838:Middle ear inflammation |  |  |  | Term was added on 27 January 2022. For details, see comment No. 232 of the Change Log (Release Number 2022). | E09¦E0909
| Labyrinthitis |  | E0910 | An inflammation of the labyrinth in the inner ear that causes dizziness and reduced or distorted hearing. | MedDRA:10023567:Labyrinthitis | Ear and Labyrinth | Infections |  | Term was added on 27 January 2022. For details, see comment No. 233 of the Change Log (Release Number 2022). | E09¦E0910
| Cholesteatoma |  | E0911 | An abnormal growth of skin cells in the middle ear, with or without infection. May be invasive and can erode neighboring structures including bone. | MedDRA:10008642:Cholesteatoma |  |  |  | Term added for Release Number 2025. See no. 16 of Change Request Log | E09¦E0911
| Ear Injury |  | E0912 | Injury to the outer, middle or inner ear. | MedDRA:10057446:Ear injury | Ear and Labyrinth | Injury | Modified (editorial) | Term added for Release Number 2025. See no. 76 of Change Request Log. Editorial change to term for Release Number 2026. See no. 269 of Change Log | E09¦E0912
| Cerumen Impaction |  | E0913 | A buildup and/or blockage of earwax in the ear canal. | MedDRA:10050337:Cerumen impaction |  |  | Modified (editorial) | Term added for Release Number 2025. See no. 105 of Change Request Log. Editorial change to term for Release Number 2026. See no. 270 of Change Log | E09¦E0913
| Neonatal Hearing Impairment |  | E1512 | An abnormality that affects hearing within the first month after birth. | MedDRA:10080921:Neonatal hearing impairment | Pregnancy, Childbirth and the Puerperium | Ear and Labyrinth |  |  | E09¦E1512
|  | Complete Neonatal Hearing Loss | E151201 | Complete loss of the ability to detect or understand sounds present in an infant within its first month after birth. | MedDRA:10082499:Complete neonatal hearing loss | Pregnancy, Childbirth and the Puerperium | Ear and Labyrinth |  |  | E09¦E1512¦E151201
|  | Partial Neonatal Hearing loss | E151202 | Partial loss of the ability to detect or understand sounds present in an infant within its first month after birth. | MedDRA:10082496:Partial neonatal hearing loss | Pregnancy, Childbirth and the Puerperium | Ear and Labyrinth |  | Editorial change to term for Release Number 2026. See no.274 of Change Log. | E09¦E1512¦E151202
Gastrointestinal System |  |  | E10 | Gastrointestinal System |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E10
| Abdominal Distention |  | E1001 | The condition of the abdomen being enlarged or swollen from internal pressure. | MedDRA:10000060:Abdominal distension |  |  |  |  | E10¦E1001
| Abdominal Pain |  | E1002 | Painful sensation in the abdominal region. | MedDRA:10000081:Abdominal pain | Gastrointestinal System | Generalized Disorders |  |  | E10¦E1002
| Achalasia |  | E1003 | Failure of the smooth muscle fibers of the gastrointestinal tract to relax at any one point of junction of one part with another. | MedDRA:10000420:Achalasia |  |  |  |  | E10¦E1003
| Ascites |  | E1004 | Accumulation of fluid in the peritoneal cavity. The fluid may be serous, hemorrhagic, or the result of tumor metastasis to the peritoneum. | MedDRA:10003445:Ascites |  |  |  |  | E10¦E1004
| Bowel Burn |  | E1005 | A thermal injury to the bowels. | MedDRA:10082501:Intestinal burn | Gastrointestinal System | Injury |  |  | E10¦E1005
| Bowel Perforation |  | E1006 | A rupture in the wall of the small or large intestine due to traumatic or pathologic processes. | MedDRA:10006056:Bowel perforation | Gastrointestinal System | Procedural Complications |  |  | E10¦E1006
| Constipation |  | E1007 | Irregular and infrequent or difficult evacuation of the bowels. | MedDRA:10010774:Constipation |  |  |  |  | E10¦E1007
| Diarrhea |  | E1008 | Passage of loose or watery stools, often occurring with increased frequency. | MedDRA:10012727:Diarrhea |  |  | Modified (technical) | Technical change to definition for Release Number 2026. See no. 32 of Change Log. | E10¦E1008
| Dysphagia/ Odynophagia |  | E1009 | A symptom referring to difficult and or painful swallowing. | MedDRA:10013950:Dysphagia |  |  |  |  | E10¦E1009
| Eructate |  | E1010 | To belch, the casting of upwind from the stomach. | MedDRA:10015137:Eructation |  |  |  |  | E10¦E1010
| Flatus |  | E1011 | Gas or air in the gastrointestinal tract. | MedDRA:10016769:Flatus |  |  |  |  | E10¦E1011
| Gastritis |  | E1012 | Inflammation of the stomach. | MedDRA:10017853:Gastritis |  |  |  |  | E10¦E1012
| Gastroesophageal Burn |  | E1013 | Thermal, chemical or radiation injury to stomach or esophagus. | MedDRA:10082500:Gastroesophageal burn | Gastrointestinal System | Injury |  |  | E10¦E1013
| Gastrointestinal Hemorrhage |  | E1014 | Bleeding originating from any part of the gastrointestinal tract. | MedDRA:10017960:Gastrointestinal hemorrhage |  |  |  |  | E10¦E1014
| Gastrointestinal Regurgitation |  | E1015 | Flow in the opposite direction from normal, as the casting up of undigested food or gas from the stomach, or the backward flowing of blood into the heart, or between the chambers of the heart when a valve is incompetent. | MedDRA:10067171:Regurgitation |  |  |  |  | E10¦E1015
| Hematemesis |  | E1016 | Vomiting blood. | MedDRA:10019418:Hematemesis |  |  |  |  | E10¦E1016
| Intra-Abdominal Hemorrhage |  | E1017 | Bleeding within the abdominal cavity. | MedDRA:10055291:Intra-abdominal hemorrhage |  |  |  |  | E10¦E1017
| Laceration(s) of Esophagus |  | E1018 | A torn, ragged, mangled wound, or an accidental cut of esophagus. | MedDRA:10068164:Esophageal laceration |  |  |  |  | E10¦E1018
| Melena |  | E1019 | Abnormally dark tarry feces containing blood, usually from gastrointestinal bleeding. | MedDRA:10027162:Melena |  |  |  |  | E10¦E1019
| Nausea |  | E1020 | Upper abdominal discomfort associated with an urge to vomit. | MedDRA:10028813:Nausea |  |  |  |  | E10¦E1020
| Pancreatitis |  | E1021 | Inflammation of the pancreas. | MedDRA:10033645:Pancreatitis | Gastrointestinal System | Infections |  |  | E10¦E1021
| Perforation of Esophagus |  | E1022 | The presence of a hole or other type of opening in the esophageal wall through which the contents of the esophagus can pass into the mediastinum. | MedDRA:10034419:Perforation of esophagus | Gastrointestinal System | Procedural Complications |  |  | E10¦E1022
| Peritoneal Laceration(s) |  | E1023 | A cut or tear of the peritoneum. | MedDRA:10068140:Peritoneal laceration | Gastrointestinal System | Injury |  |  | E10¦E1023
| Peritonitis |  | E1024 | Inflammation of the peritoneum due to infection by bacteria or fungi. | MedDRA:10034674:Peritonitis | Gastrointestinal System | Infections |  |  | E10¦E1024
| Heartburn/Indigestion |  | E1025 | Abdominal discomfort with retrosternal pain usually associated with gastroesophageal reflux, pyrosis or dyspepsia. | MedDRA:10013946:Dyspepsia |  |  |  | Term/Definition was modified on 27 January 2022 (Old term/definition: Pyrosis/Heartburn: Abdominal discomfort with retrosternal pain usually associated with gastroesophageal reflux.) MedDRA mapping also updated from Heartburn(10019326) to Dyspepsia(10013946). For details, see comment No. 110 of the Change Log (Release Number 2022). | E10¦E1025
| Rectal Anastomotic Leakage |  | E1026 | Leakage due to breakdown of a rectal anastomosis. | MedDRA:10065894:Rectal anastomotic leak | Gastrointestinal System | Procedural Complications | Retired (2024) | Term retired for Release Number 2024. Use E2108 - Failure to Anastomose. See no. 81 of Change Log | E10¦E1026
| Retroperitoneal Hemorrhage |  | E1027 | Bleeding originating in the retroperitoneal area. | MedDRA:10038981:Retroperitoneal hemorrhage |  |  |  |  | E10¦E1027
| Salivary Gland Problem |  | E1028 | A non-neoplastic or neoplastic (benign or malignant) disorder involving a salivary gland. | MedDRA:10061935:Salivary gland disorder |  |  |  |  | E10¦E1028
|  | Dry Mouth | E102801 | The condition of not having enough saliva to keep the mouth moist. | MedDRA:10013781:Dry mouth |  |  |  |  | E10¦E1028¦E102801
|  | Salivary Hypersecretion | E102802 | Excessive saliva production. | MedDRA:10039424:Salivary hypersecretion |  |  |  |  | E10¦E1028¦E102802
| Stenosis of the Esophagus |  | E1029 | Narrowing or tightening of the esophagus | MedDRA:10041978:Stenosis esophageal |  |  | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 271 of Change Log. | E10¦E1029
| Stomach Ulceration |  | E1030 | The formation of sores or a break in the wall of the stomach. | MedDRA:10042116:Stomach ulcer |  |  |  |  | E10¦E1030
| Stomatitis |  | E1031 | Inflammation of the oral mucosa. | MedDRA:10042128:Stomatitis |  |  |  |  | E10¦E1031
|  | Glossitis | E103101 | Inflammation of the tongue. | MedDRA:10018386:Glossitis |  |  |  |  | E10¦E1031¦E103101
| Vomiting |  | E1032 | Forceful ejection of the contents of the stomach through the mouth. | MedDRA:10047700:Vomiting |  |  |  |  | E10¦E1032
| Unspecified Gastrointestinal Problem |  | E1033 | The report describes a non-specific problem with the gastrointestinal system. Note: Please use E2402 if the report describes a specific problem but the relevant term does not exist. |  |  |  | Modified (editorial) | Term was added on 18 March 2020 for cases where the category of a health effect was reported but with no further details. Editorial change to definition for Release Number 2026. See no. 205 of Release Log. | E10¦E1033
| Fecal Urgency |  | E1034 | A continual or recurrent inclination to evacuate the bowels. | MedDRA:10012114:Defecation urgency |  |  |  | Term was added on 27 January 2022. For details, see comment No. 7 of the Change Log (Release Number 2022). | E10¦E1034
| Pneumoperitoneum |  | E1035 | Free air within the peritoneal cavity. | MedDRA:10048299:Pneumoperitoneum |  |  |  | Term was added on 27 January 2022. For details, see comment No. 205 of the Change Log (Release Number 2022). | E10¦E1035
| Colitis |  | E1036 | Inflammation of the large intestine | MedDRA:10009887:Colitis |  |  |  | Term added for Release Number 2024. See no. 102 of Change Log | E10¦E1036
| Ileus |  | E1037 | Reduced or absent intestinal motility resulting in impaired transit of intestinal contents; may be referred to as paralytic ileus, non-obstructive ileus, or functional bowel obstruction | MedDRA:10021328:Ileus |  |  | New | Term added for Release Number 2026. See no. 20 of Change Log | E10¦E1037
| Enteritis |  | E1038 | Inflammation of the small intestine, including duodenitis, jejunitis, or ileitis; may be associated with mucosal edema or intestinal wall thickening | MedDRA:10014866:Enteritis |  |  | New | Term added for Release Number 2026. See no. 21 of Change Log | E10¦E1038
| Retching |  | E1039 | Spasmodic and periodic contractions of the respiratory muscles (diaphragm and muscles of the chest and abdominal wall) without expulsion of the stomach contents | MedDRA:10038776:Retching |  |  | New | Term added for Release Number 2026. See no. 28 of Change Log | E10¦E1039
| Abdominal Compartment Syndrome |  | E1040 | Abdominal compartment syndrome (ACS) is a condition that occurs when bleeding or swelling leads to rises of pressure in the abdominal cavity that causes end-organ damage, dysfunction or failure. | MedDRA:10058808:Abdominal compartment syndrome |  |  | New | Term added for Release Number 2026. See no. 162 of Change Log | E10¦E1040
| Increased Intra-Peritoneal Volume (IIPV) |  | E1303 | Abnormal build up of fluid in the abdomen. | MedDRA:10082492:Increased intraperitoneal volume | Kidney and Urinary Tract | Gastrointestinal System |  |  | E10¦E1303
| Difficulty Chewing |  | E1607 | Problem with mastication due to pain, discomfort or loss of mobility in the jaw. | MedDRA:10067464:Chewing difficulty | Musculoskeletal System | Gastrointestinal System |  |  | E10¦E1607
| Fecal Incontinence |  | E232401 | Inability to control voluntary passage of feces from the body. | MedDRA:10016296:Fecal incontinence | Generalized Disorders | Gastrointestinal System |  | Editorial change to add Gastrointestinal System as Secondary Category. For details, see comment No. 99 of the Change Log (Release Number 2023). Code hierarchy adjusted for secondary category to prevent primary parent term appearing in three places. Release No. 2024 | E10¦E232401
Hepatic and Biliary System |  |  | E11 | Hepatic and Biliary System |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E11
| Biliary Cirrhosis |  | E1101 | Cirrhosis of the liver caused either by destruction of the intrahepatic bile ducts (primary biliary cholangitis) or blockage of the extrahepatic bile ducts (secondary biliary cholangitis). | MedDRA:10004659:Biliary cirrhosis |  |  |  | Definition was modified on 21 January 2021 (Old definition: Biliary Cirrhosis/Cirrhosis of the liver caused either by destruction of the intrahepatic bile ducts (primary biliary cirrhosis) or blockage of the extrahepatic bile ducts (secondary biliary cirrhosis).). For details, refer to comment No. 5 of the Change Log. | E11¦E1101
| Hepatitis |  | E1102 | Inflammation of the liver; usually from a viral infection, but sometimes from toxic agents. | MedDRA:10056522:Hepatic infection | Hepatic and Biliary System | Infections |  |  | E11¦E1102
| Hyperbilirubinemia |  | E1103 | Abnormally high level of bilirubin in the blood. Excess bilirubin is associated with jaundice. | MedDRA:10020582:Hyperbilirubinemia |  |  |  |  | E11¦E1103
| Liver Damage/Dysfunction |  | E1104 | Non traumatic damage to the liver parenchyma resulting in hepatic insufficiency. | MedDRA:10019831:Hepatocellular damage |  |  |  |  | E11¦E1104
| Liver Failure |  | E1105 | Sustained, and severe damage to the liver with significant loss of function. | MedDRA:10024678:Liver failure |  |  |  |  | E11¦E1105
| Liver Laceration(s) |  | E1106 | A cut or tear in the liver. | MedDRA:10068143:Liver laceration |  |  |  |  | E11¦E1106
| Unspecified Hepatic or Biliary Problem |  | E1107 | The report describes a non-specific problem with the liver or biliary system. Note: Please use E2402 if the report describes a specific problem but the relevant term does not exist. |  |  |  | Modified (editorial) | Term was added on 18 March 2020 for cases where the category of a health effect was reported but with no further details. Editorial change to definition for Release Number 2026. See no. 206 of Change Log. | E11¦E1107
| Biliary Leak |  | E1108 | Sudden unexpected leakage of bile from the liver or soft tissue. | MedDRA:10057586:Post procedural bile leak | Hepatic and Biliary System | Procedural Complications |  | Term was added on 27 January 2022. For details, see comment No. 5 of the Change Log (Release Number 2022). | E11¦E1108
| Cholangitis |  | E1109 | Inflammation of the bile duct. | MedDRA:10008604:Cholangitis | Hepatic and Biliary System | Procedural Complications |  | Term was added on 27 January 2022. For details, see comment No. 6 of the Change Log (Release Number 2022). | E11¦E1109
Endocrine, Metabolism and Nutrition |  |  | E12 | Endocrine, Metabolism and Nutrition |  |  |  | Not selectable | Please use more detailed term within the hierarchy. NOTE: Term was modified on 18 March 2020 to include Endocrine. | E12
| Respiratory Acidosis |  | E0739 | A condition in which the blood pH is less than normal, secondary to impaired gas exchange. | MedDRA:10038661:Respiratory acidosis | Respiratory System | Endocrine, Metabolism and Nutrition |  |  | E12¦E0739
| Respiratory Alkalosis |  | E0740 | A condition in which the blood pH is greater than normal, secondary to impaired gas exchange. | MedDRA:10038664:Respiratory alkalosis | Respiratory System | Endocrine, Metabolism and Nutrition |  |  | E12¦E0740
| Dehydration |  | E1201 | A condition resulting from the excessive loss of fluid from the body. | MedDRA:10012174:Dehydration |  |  |  |  | E12¦E1201
| Electrolyte Imbalance |  | E1202 | Higher or lower than normal values for the serum electrolytes; usually affecting Na, K, Cl, Ca, CO2, BUN. | MedDRA:10014418:Electrolyte imbalance | Endocrine, Metabolism and Nutrition | Blood and Lymphatic System |  |  | E12¦E1202
|  | Hypernatremia | E120201 | Higher than normal levels of sodium in the circulating blood. | MedDRA:10020680:Hypernatremia | Endocrine, Metabolism and Nutrition | Blood and Lymphatic System |  |  | E12¦E1202¦E120201
|  | Hyponatremia | E120202 | Lower than normal levels of sodium in the circulating blood. | MedDRA:10021038:Hyponatremia | Endocrine, Metabolism and Nutrition | Blood and Lymphatic System |  |  | E12¦E1202¦E120202
|  | Hypercalcemia | E120203 | Higher than normal levels of calcium in the circulating blood. | MedDRA:10020587:Hypercalcemia |  |  |  | Term was added on 27 January 2022. For details, see comment No. 38 / 90 of the Change Log (Release Number 2022). | E12¦E1202¦E120203
|  | Hypocalcemia | E120204 | Lower than normal levels of calcium in the circulating blood. | MedDRA:10020949:Hypocalcemia |  |  |  | Term was added on 27 January 2022. For details, see comment No. 39 / 91 of the Change Log (Release Number 2022). | E12¦E1202¦E120204
|  | Hyperkalemia | E120205 | Higher than normal levels of potassium in the circulating blood. | MedDRA:10020647:Hyperkalemia |  |  |  | Term was added on 27 January 2022. For details, see comment No. 63 of the Change Log (Release Number 2022). | E12¦E1202¦E120205
|  | Hypokalemia | E120206 | Lower than normal levels of potassium in the circulating blood. | MedDRA:10021018:Hypokalemia |  |  |  | Term was added on 27 January 2022. For details, see comment No. 64 of the Change Log (Release Number 2022). | E12¦E1202¦E120206
|  | Blood Bicarbonate Increased | E120207 | Higher than normal levels of bicarbonate in the blood. | MedDRA:10005360:Blood bicarbonate increased |  |  |  | Term was added on 27 January 2022. For details, see comment No. 88 of the Change Log (Release Number 2022). | E12¦E1202¦E120207
|  | Blood Bicarbonate Decreased | E120208 | Lower than normal levels of bicarbonate in the blood. | MedDRA:10005359:Blood bicarbonate decreased |  |  |  | Term was added on 27 January 2022. For details, see comment No. 89 of the Change Log (Release Number 2022). | E12¦E1202¦E120208
|  | Hyperphosphatemia | E120209 | Higher than normal levels of phosphate in the circulating blood. | MedDRA:10020712:Hyperphosphatemia |  |  |  | Term was added on 27 January 2022. For details, see comment No. 61 of the Change Log (Release Number 2022). | E12¦E1202¦E120209
| Feeding Problem |  | E1203 | Interruption or disruption of the intake of food/nutrition. | MedDRA:10061148:Feeding disorder |  |  |  |  | E12¦E1203
| Hot Flashes/Flushes |  | E1204 | A temporary and/or sudden feeling of intense body warmth, flushing, sometimes accompanied by sweating. | MedDRA:10020407:Hot flashes |  |  |  |  | E12¦E1204
| Hyperglycemia |  | E1205 | Abnormally high level of glucose in the blood. | MedDRA:10020639:Hyperglycemia | Endocrine, Metabolism and Nutrition | Investigations and Diagnostic Tests |  |  | E12¦E1205
|  | Elevated Ketones/Diabetic Ketoacidosis | E120501 | Elevated ketones, including metabolic acidosis produced by accumulation of ketone bodies resulting from uncontrolled diabetes mellitus. | MedDRA:10057594:Blood ketone body increased |  |  | Modified (editorial) | Editorial change on 31 January 2023 to update MedDRA mapping from Diabetic ketoacidosis (10012671) to Blood ketone body increased (10057594). For details, see comment No. 85 / 86 of the Change Log (Release Number 2023). Editorial change to term for Release Number 2026. See no. 272 of Change Log. | E12¦E1205¦E120501
| Hypoglycemia |  | E1206 | Abnormally low level of glucose in the blood. | MedDRA:10021005:Hypoglycemia | Endocrine, Metabolism and Nutrition | Investigations and Diagnostic Tests |  |  | E12¦E1206
| Polydipsia |  | E1207 | Extreme thirst, often accompanied by chronic excessive intake of water | MedDRA:10036067:Polydipsia |  |  | Modified (technical) | Technical change to definition for Release Number 2026. See no. 31 of Change Log. | E12¦E1207
| Weight Changes |  | E1208 | Repetitive loss and regain of body weight. | MedDRA:10049040:Weight fluctuation |  |  |  |  | E12¦E1208
| Hormonal Imbalance |  | E1209 | An imbalance of one or more hormones in the body | MedDRA:10020381:Hormonal imbalance | Endocrine, Metabolism and Nutrition | Investigations and Diagnostic Tests |  | Term was added on 18 March 2020 as an additional level 2 term for "Endocrine, Metabolism and Nutrition". | E12¦E1209
| Premature Menopause |  | E1210 | Menopause occurring spontaneously before 40 years of age | MedDRA:10036601:Premature menopause | Endocrine, Metabolism and Nutrition | Reproductive System and Breast |  | Term was added on 18 March 2020 as an additional level 2 term for "Endocrine, Metabolism and Nutrition". | E12¦E1210
| Vitamin Deficiency |  | E1211 | Low levels of one or more vitamins in the body. | MedDRA:10047628:Vitamin deficiency |  |  |  | Term was added on 18 March 2020 as an additional level 2 term for "Endocrine, Metabolism and Nutrition". | E12¦E1211
| Hypothyroidism |  | E1212 | Abnormally low levels of thyroid hormone | MedDRA:10021114:Hypothyroidism |  |  |  | Term was added on 31 January 2023. For details, see comment No. 87 of the Change Log (Release Number 2023). | E12¦E1212
| Uremia |  | E1306 | Abnormally high concentration of non-protein nitrogen in the blood. | MedDRA:10046369:Uremia | Kidney and Urinary Tract | Endocrine, Metabolism and Nutrition |  |  | E12¦E1306
| Decreased Appetite |  | E2306 | Lack of desire for food. | MedDRA:10061428:Decreased appetite | Generalized Disorders | Endocrine, Metabolism and Nutrition |  |  | E12¦E2306
| Increased Appetite |  | E2307 | Intense desire for food. | MedDRA:10021654:Increased appetite | Generalized Disorders | Endocrine, Metabolism and Nutrition |  |  | E12¦E2307
| Metabolic Acidosis |  | E1213 | An increase in the hydrogen ion concentration in the systemic circulation that results in an abnormally low serum bicarbonate level. | MedDRA:10027417:Metabolic acidosis |  |  |  | Term added for Release Number 2025. See no. 33 of Change Request Log | E12¦E1213
| Hyperthyroidism |  | E1214 | Abnormally high levels of thyroid hormone | MedDRA:10020850:Hyperthyroidism |  |  |  | Term added for Release Number 2025. See no. 73 of Change Request Log | E12¦E1214
| Metabolic Alkalosis |  | E1215 | A disorder characterized by abnormally high alkalinity (low hydrogen-ion concentration) of the blood and other body tissues. | MedDRA:10027423:Metabolic alkalosis |  |  |  | Term added for Release Number 2025. See no. 34 of Change Request Log | E12¦E1215
| Malnutrition |  | E1216 | A state of nutritional deficiency or imbalance | MedDRA:10061273:Malnutrition |  |  | New | Term added for Release Number 2026. See no. 30 of Change Log. | E12¦E1216
| Thyroiditis |  | E1217 | Thyroiditis reflects inflammation of the thyroid gland. | MedDRA:10043778:Thyroiditis |  |  | New | Term added for Release Number 2026. See no. 89 of Change Log. | E12¦E1217
Kidney and Urinary Tract |  |  | E13 | Kidney and Urinary Tract |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E13
| Dysuria |  | E1301 | Difficult or painful urination. | MedDRA:10013990:Dysuria |  |  |  |  | E13¦E1301
| Hematuria |  | E1302 | Blood in the urine. | MedDRA:10019450:Hematuria |  |  |  |  | E13¦E1302
| Increased Intra-Peritoneal Volume (IIPV) |  | E1303 | Abnormal build up of fluid in the abdomen. | MedDRA:10082492:Increased intraperitoneal volume | Kidney and Urinary Tract | Gastrointestinal System |  |  | E13¦E1303
| Micturition Urgency |  | E1304 | A sudden compelling urge to urinate. | MedDRA:10027566:Micturition urgency |  |  |  |  | E13¦E1304
| Renal Impairment |  | E1305 | An acute or chronic condition where the kidneys are not functioning normally. | MedDRA:10062237:Renal impairment |  |  |  | Editorial change to the definition on 16 February 2021. | E13¦E1305
|  | Renal Failure | E130501 | An acute or chronic condition that is characterized by the inability of the kidneys to adequately filter the blood resulting in uremia and electrolyte imbalances. | MedDRA:10038435:Renal failure |  |  |  |  | E13¦E1305¦E130501
| Uremia |  | E1306 | Abnormally high concentration of non-protein nitrogen in the blood. | MedDRA:10046369:Uremia | Kidney and Urinary Tract | Endocrine, Metabolism and Nutrition |  |  | E13¦E1306
| Urethral Stenosis/Stricture |  | E1307 | Narrowing of any part of the urethra. With or without affecting the flow of urine. | MedDRA:10065584:Urethral stenosis |  |  |  |  | E13¦E1307
| Urinary Frequency / Polyuria |  | E1308 | Urination at short intervals or at higher volumes; it may result from increased urine formation, decreased bladder capacity, or lower urinary tract irritation. | MedDRA:10036142:Polyuria |  |  |  | Term/Definition was modified on 31 January 2023 (Old term/definition: Urinary Frequency: Urination at short intervals; it may result from increased urine formation, decreased bladder capacity, or lower urinary tract irritation.) MedDRA mapping also updated from Urinary frequency (10046539) to Polyuria (10036142). For details, see comment No. 88 of the Change Log (Release Number 2023). | E13¦E1308
| Urinary Retention |  | E1309 | Accumulation of urine within the bladder because of the inability to urinate. | MedDRA:10046555:Urinary retention |  |  |  |  | E13¦E1309
| Urinary Tract Infection |  | E1310 | An inflammatory process affecting any part of the urinary tract, most commonly the bladder or the urethra. | MedDRA:10046571:Urinary tract infection | Kidney and Urinary Tract | Infections |  |  | E13¦E1310
|  | Kidney Infection | E131001 | An inflammatory process affecting the kidney. | MedDRA:10023424:Kidney infection | Kidney and Urinary Tract | Infections |  |  | E13¦E1310¦E131001
| Unspecified Kidney or Urinary Problem |  | E1311 | The report describes a non-specific problem with the kidneys or urinary system. Note: Please use E2402 if the report describes a specific problem but the relevant term does not exist. |  |  |  | Modified (editorial) | Term was added on 18 March 2020 for cases where the category of a health effect was reported but with no further details. Editorial change to definition for Release Number 2026. See no. 207 of Change Log. | E13¦E1311
| Urolithiasis |  | E1312 | A condition arising from the formation of renal calculi (e.g. kidney stones, bladder stones, urinary tract stones). | MedDRA:10046707:Urolithiasis |  |  |  | Term was added on 31 January 2023. For details, see comment No. 89 of the Change Log (Release Number 2023). | E13¦E1312
| Interstitial Cystitis |  | E1313 | A chronic bladder condition resulting in recurring discomfort or pain in the bladder or surrounding pelvic region. Also known as Bladder Pain Syndrome (BPS). | MedDRA:10011796:Cystitis interstitial |  |  |  | Term was added on 31 January 2023. For details, see comment No. 90 of the Change Log (Release Number 2023). | E13¦E1313
| Bladder Wall Thickening |  | E1314 | Thickening of the bladder walls, including Bladder Trabeculation. | MedDRA:10079064:Bladder wall thickening |  |  |  | Term was added on 31 January 2023. For details, see comment No. 91 of the Change Log (Release Number 2023). | E13¦E1314
| Hydronephrosis |  | E1315 | A condition of excess urine accumulation in kidney(s) that causes swelling of kidneys. | MedDRA:10020524:Hydronephrosis |  |  |  | Term added for Release Number 2024. See no. 5 of Change Log | E13¦E1315
| Bladder Dysfunction |  | E1316 | Impaired or abnormal functioning of the bladder, including sphincters. | MedDRA:10069632:Bladder dysfunction |  |  |  | Term added for Release Number 2024. See no. 14 of Change Log | E13¦E1316
| Kidney Injury |  | E1317 | Traumatic damage to the kidney(s). | MedDRA:10070510:Traumatic renal injury |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 15 of Change Log. Editorial change to term for Release Number 2026. See no. 273 of Change Log. | E13¦E1317
| Urinary Incontinence |  | E232402 | Inability to control voluntary passage of urine from the body. | MedDRA:10046543:Urinary incontinence | Generalized Disorders | Kidney and Urinary Tract |  | Editorial change to add Kidney and Urinary Tract as Secondary Category. For details, see comment No. 100 of the Change Log (Release Number 2023). Amended in Version 1 on 23 February 2023. Code hierarchy adjusted for secondary category to prevent primary parent term appearing in three places. Release No. 2024 | E13¦E232402
Reproductive System and Breast |  |  | E14 | Reproductive System and Breast |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E14
| Premature Menopause |  | E1210 | Menopause occurring spontaneously before 40 years of age | MedDRA:10036601:Premature menopause | Endocrine, Metabolism and Nutrition | Reproductive System and Breast |  | Term was added on 18 March 2020 as an additional level 2 term for "Endocrine, Metabolism and Nutrition". | E14¦E1210
| Abnormal Vaginal Discharge |  | E1401 | Secretions from the vagina with a smell, color, or texture which is different from normal. | MedDRA:10057094:Vaginal discharge abnormality |  |  | Modified (editorial) | Editorial change to definition for Release Number 2026. See no. 189 of Change Log. | E14¦E1401
| Breast Discomfort/Pain |  | E1402 | Having pain, tenderness or discomfort in the breast. | MedDRA:10006298:Breast pain |  |  |  |  | E14¦E1402
| Breast Mass |  | E1403 | Abnormality in the breast parenchyma, identified either by palpation or imaging. These abnormalities include calcifications, breast tissue distortion, presence of areas of density, and presence of benign or malignant tumors. | MedDRA:10006272:Breast mass |  |  |  |  | E14¦E1403
| Cervical Changes |  | E1404 | Changes such as abnormal, pathological, benign, malignant, etc. | MedDRA:10008351:Cervix disorder | Reproductive System and Breast | Investigations and Diagnostic Tests |  |  | E14¦E1404
| Dyspareunia |  | E1405 | Painful or difficult coitus. | MedDRA:10013941:Dyspareunia |  |  |  |  | E14¦E1405
| Fallopian Tube Perforation |  | E1406 | A rupture in the fallopian tube due to traumatic or pathologic processes. | MedDRA:10065790:Fallopian tube perforation | Reproductive System and Breast | Procedural Complications |  |  | E14¦E1406
| Genital Bleeding |  | E1407 | Abnormal bleeding from the genitalia. | MedDRA:10071812:Genital bleeding |  |  |  |  | E14¦E1407
|  | Intermenstrual Bleeding | E140701 | Uterine bleeding that occurs during a time not associated with the normal menstrual period. | MedDRA:10022559:Intermenstrual bleeding |  |  |  |  | E14¦E1407¦E140701
|  | Menorrhagia | E140702 | Heavy or large blood loss during menses. | MedDRA:10027313:Menorrhagia |  |  |  |  | E14¦E1407¦E140702
| Menstrual Irregularities |  | E1408 | Deviations from the normal process; e.g. delayed, difficult, profuse, scanty, unusual bleeding, etc. | MedDRA:10027333:Menstrual irregularity |  |  |  |  | E14¦E1408
| Nipple Sensation Changes |  | E1409 | Excessive sensitivity or changes in sensitivity in the nipple(s) | MedDRA:10082498:Nipple sensation changes |  |  |  | Editorial change to the definition on 16 February 2021. | E14¦E1409
| Nipple Ulceration |  | E1410 | Ulcer that develops in the nipple. | MedDRA:10029424:Nipple ulceration |  |  |  |  | E14¦E1410
| Pelvic Inflammatory Disease |  | E1411 | Any pelvic infection involving the upper female genital tract beyond the cervix. | MedDRA:10034254:Pelvic inflammatory disease |  |  |  |  | E14¦E1411
| Rectovaginal Fistula |  | E1412 | Abnormal connection between rectum and vagina. | MedDRA:10051097:Rectovaginal fistula |  |  |  | Editorial change to the definition on 16 February 2021. | E14¦E1412
| Sexual Dysfunction |  | E1413 | Difficulty, lack of desire, or pain associated with coitus. | MedDRA:10040477:Sexual dysfunction |  |  |  |  | E14¦E1413
|  | Erectile Dysfunction | E141301 | Persistent or recurrent inability to achieve or to maintain an erection during sexual activity. | MedDRA:10061461:Erectile dysfunction |  |  |  |  | E14¦E1413¦E141301
| Toxic Shock Syndrome |  | E1414 | A rare acute life-threatening systemic bacterial noncontagious illness caused by staphylococcal exotoxins. | MedDRA:10044248:Toxic shock syndrome | Reproductive System and Breast | Infections |  |  | E14¦E1414
| Uterine Perforation |  | E1415 | A rupture in the uterus due to traumatic or pathologic processes. | MedDRA:10046810:Uterine perforation | Reproductive System and Breast | Procedural Complications |  |  | E14¦E1415
| Vesicovaginal Fistula |  | E1416 | Abnormal connection between vagina and bladder. | MedDRA:10069086:Vesicovaginal fistula |  |  |  |  | E14¦E1416
| Unspecified Reproductive System or Breast Problem |  | E1417 | The report describes a non-specific problem with the breasts or reproductive system. Note: Please use E2402 if the report describes a specific problem but the relevant term does not exist. |  |  |  | Modified (editorial) | Term was added on 18 March 2020 for cases where the category of a health effect was reported but with no further details. Editorial change to definition for Release Number 2026. See no. 208 of Change Log. | E14¦E1417
| Breast Feeding Problem |  | E1418 | Issues related to the ability to breastfeed. |  |  |  |  | Term was added on 27 January 2022. For details, see comment No. 158 / 160 of the Change Log (Release Number 2022). | E14¦E1418
| Vaginal Dryness |  | E1419 | An uncomfortable feeling of itching and burning in the vaginal opening resulting from inadequate vaginal lubrication. | MedDRA:10046904:Vaginal dryness |  |  |  | Term was added on 31 January 2023. For details, see comment No. 92 of the Change Log (Release Number 2023). | E14¦E1419
| Pelvic Pain |  | E1420 | Pain in the pelvic area | MedDRA:10034263:Pelvic pain | Reproductive System and Breast | Generalized Disorders |  | Term was added on 31 January 2023. For details, see comment No. 93 of the Change Log (Release Number 2023). | E14¦E1420
| Symmastia |  | E1421 | Symmastia is a condition where the breasts join together in the middle of the chest, resulting in a lack of cleavage. | MedDRA:10073930:Symmastia |  |  | New | Term added for Release Number 2026. See no. 92 of Change Log. | E14¦E1421
| Ectopic Pregnancy |  | E1506 | Pregnancy condition that happens when there is a developing fetus or embryo outside the uterus. | MedDRA:10014166:Ectopic pregnancy | Pregnancy, Childbirth and the Puerperium | Reproductive System and Breast |  |  | E14¦E1506
| Incompetent Cervix |  | E1509 | Cervix that is abnormally prone to dilate without contractions in the second trimester of pregnancy, resulting in premature expulsion of the fetus. | MedDRA:10021632:Incompetent cervix | Pregnancy, Childbirth and the Puerperium | Reproductive System and Breast |  |  | E14¦E1509
| Vaginal Mucosa Damage |  | E1519 | Injury to vaginal mucosa. | MedDRA:10080926:Vaginal mucosal damage | Pregnancy, Childbirth and the Puerperium | Reproductive System and Breast |  |  | E14¦E1519
| Double Capsule |  | E2341 | The presence of two distinct capsular layers around an implant with an intercapsular space |  | Generalized Disorders | Reproductive System and Breast |  | Term was added on 18 March 2020 as an additional level 2 term for "Generalized Disorders". | E14¦E2341
Pregnancy, Childbirth and the Puerperium |  |  | E15 | Pregnancy, Childbirth and the Puerperium |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E15
| Abortion |  | E1501 | The unintentional or intentional loss of a pregnancy before 22 weeks gestation. | MedDRA:10000210:Abortion |  |  |  |  | E15¦E1501
|  | Complete Induced Abortion | E150101 | Expulsion of all the products of conception following medical or surgical pregnancy termination. | MedDRA:10060928:Abortion induced complete |  |  |  |  | E15¦E1501¦E150101
|  | Incomplete Induced Abortion | E150102 | The uterus is not entirely emptied of products of conception following medical or surgical pregnancy termination. | MedDRA:10053984:Abortion induced incomplete |  |  |  |  | E15¦E1501¦E150102
|  | Induced Abortion | E150103 | Abortion brought on intentionally. | MedDRA:10021716:Induced abortion |  |  |  |  | E15¦E1501¦E150103
|  | Missed Abortion | E150104 | Retention in uterus of an abortus. | MedDRA:10027704:Missed abortion |  |  |  |  | E15¦E1501¦E150104
|  | Spontaneous Abortion | E150105 | Spontaneous expulsion of a human fetus before it is viable. | MedDRA:10041687:Spontaneous abortion |  |  |  |  | E15¦E1501¦E150105
| Amnionitis |  | E1502 | Inflammation of the amnion. | MedDRA:10051242:Amnionitis |  |  |  |  | E15¦E1502
| Cephalohematoma |  | E1503 | A subperiosteal hemorrhage limited to the surface of one cranial bone, a usually benign condition seen in the newborn as a result of bone trauma. | MedDRA:10060531:Cephalohematoma |  |  |  |  | E15¦E1503
| Chorioamnionitis |  | E1504 | An infection involving chorion, amnion and amniotic fluid; generally, there is also involve placental villi and from decidua. | MedDRA:10008755:Chorioamnionitis | Pregnancy, Childbirth and the Puerperium | Infections |  |  | E15¦E1504
| Congenital Defect/Deformity |  | E1505 | Any abnormality either hereditary or due an influence that happen during the pregnancy until the moment of birth. | MedDRA:10048305:Birth defects | Pregnancy, Childbirth and the Puerperium | Musculoskeletal System |  |  | E15¦E1505
| Ectopic Pregnancy |  | E1506 | Pregnancy condition that happens when there is a developing fetus or embryo outside the uterus. | MedDRA:10014166:Ectopic pregnancy | Pregnancy, Childbirth and the Puerperium | Reproductive System and Breast |  |  | E15¦E1506
| Fetal Distress |  | E1507 | Signs or symptoms which may be indicative of distress to a fetus, which may include repetitive variable decelerations, fetal tachycardia or bradycardia, late decelerations, or low biophysical profile. | MedDRA:10016486:Fetal distress | Pregnancy, Childbirth and the Puerperium | Heart |  |  | E15¦E1507
| Hypoxia in Utero |  | E1508 | Caused by conditions such as inadequate placental function (often abruptio placentae), preeclamptic toxicity, prolapse of the umbilical cord, or complications from anesthetic administration. | MedDRA:10022748:Intra-uterine hypoxia |  |  |  |  | E15¦E1508
| Incompetent Cervix |  | E1509 | Cervix that is abnormally prone to dilate without contractions in the second trimester of pregnancy, resulting in premature expulsion of the fetus. | MedDRA:10021632:Incompetent cervix | Pregnancy, Childbirth and the Puerperium | Reproductive System and Breast |  |  | E15¦E1509
| Low Apgar Score |  | E1510 | Any Apgar score less than seven. This may be the initial or a subsequent apgar score. | MedDRA:10002944:Apgar score low | Pregnancy, Childbirth and the Puerperium | Investigations and Diagnostic Tests |  |  | E15¦E1510
| Neonatal Deformities |  | E1511 | Anatomic abnormality manifested during the neonatal period of physical or mechanical origin. | MedDRA:10082494:Neonatal deformity | Pregnancy, Childbirth and the Puerperium | Musculoskeletal System |  |  | E15¦E1511
| Neonatal Hearing Impairment |  | E1512 | An abnormality that affects hearing within the first month after birth. | MedDRA:10080921:Neonatal hearing impairment | Pregnancy, Childbirth and the Puerperium | Ear and Labyrinth |  |  | E15¦E1512
|  | Complete Neonatal Hearing Loss | E151201 | Complete loss of the ability to detect or understand sounds present in an infant within its first month after birth. | MedDRA:10082499:Complete neonatal hearing loss | Pregnancy, Childbirth and the Puerperium | Ear and Labyrinth |  |  | E15¦E1512¦E151201
|  | Partial Neonatal Hearing Loss | E151202 | Partial loss of the ability to detect or understand sounds present in an infant within its first month after birth. | MedDRA:10082496:Partial neonatal hearing loss | Pregnancy, Childbirth and the Puerperium | Ear and Labyrinth | Modified (editorial) | Editorial change to term for Release Number 2026. See no.274 of Change Log. | E15¦E1512¦E151202
| Perinatal Brain Injury |  | E1513 | Injury to the brain of a newborn infant occurring around the time of birth. | MedDRA:10056392:Perinatal brain damage | Pregnancy, Childbirth and the Puerperium | Nervous System |  |  | E15¦E1513
| Pregnancy with a Contraceptive Device |  | E1514 | The woman became pregnant despite using a contraceptive device. | MedDRA:10063130:Pregnancy with contraceptive device |  |  |  |  | E15¦E1514
| Premature Labor |  | E1515 | Delivery of a viable infant before the normal end of gestation. | MedDRA:10036599:Premature labor |  |  |  |  | E15¦E1515
| Respiratory Distress Syndrome of Newborns |  | E1516 | A condition of the newborn marked by dyspnea with cyanosis. | MedDRA:10038692:Respiratory distress syndrome in newborn | Pregnancy, Childbirth and the Puerperium | Respiratory System |  |  | E15¦E1516
| Teratogenic Effects |  | E1517 | Appearance of anomalies and/or developmental defects due to exposure to an agent that can disturb the development of an embryo or fetus. | MedDRA:10043275:Teratogenicity |  |  |  |  | E15¦E1517
| Toxins In Children |  | E1518 | Children affected by toxins may be due to silicone breast implants and their effects on unborn children and from breastfeeding. | MedDRA:10053487:Exposure to toxic agent |  |  |  |  | E15¦E1518
| Vaginal Mucosa Damage |  | E1519 | Injury to vaginal mucosa. | MedDRA:10080926:Vaginal mucosal damage | Pregnancy, Childbirth and the Puerperium | Reproductive System and Breast |  |  | E15¦E1519
| Infertility |  | E1520 | Inability to conceive for at least one year after trying and having unprotected sex. Causes of female infertility include endometriosis, fallopian tubes obstruction, and polycystic ovary syndrome. Causes of male infertility include abnormal sperm production or function, blockage of the epididymis, blockage of the ejaculatory ducts and hypospadias. | MedDRA:10021926:Infertility |  |  |  | Term was added on 27 January 2022. For details, see comment No. 159 / 206 of the Change Log (Release Number 2022). | E15¦E1520
| Toxemia |  | E190103 | The condition resulting from the spread of bacterial products (toxins) by the bloodstream. | MedDRA:10059857:Toxemia | Infections | Pregnancy, Childbirth and the Puerperium |  |  | E15¦E190103
Musculoskeletal System |  |  | E16 | Musculoskeletal System |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E16
| Congenital Defect/Deformity |  | E1505 | Any abnormality either hereditary or due an influence that happen during the pregnancy until the moment of birth. | MedDRA:10048305:Birth defects | Pregnancy, Childbirth and the Puerperium | Musculoskeletal System |  |  | E16¦E1505
| Neonatal Deformities |  | E1511 | Anatomic abnormality manifested during the neonatal period of physical or mechanical origin. | MedDRA:10082494:Neonatal deformity | Pregnancy, Childbirth and the Puerperium | Musculoskeletal System |  |  | E16¦E1511
| Arthralgia |  | E1601 | Pain in a joint. | MedDRA:10003239:Arthralgia |  |  |  |  | E16¦E1601
| Arthritis |  | E1602 | An inflammatory process affecting a joint. | MedDRA:10003246:Arthritis |  |  |  |  | E16¦E1602
|  | Rheumatoid Arthritis | E160201 | A chronic systemic disease, primarily of the joints, marked by inflammatory changes in the synovial membranes and articular structures, widespread fibrinoid degeneration of the collagen fibers in mesenchymal tissues, and by atrophy and rarefaction of bony structures. | MedDRA:10039073:Rheumatoid arthritis |  |  |  |  | E16¦E1602¦E160201
| Bone Fracture(s) |  | E1603 | A traumatic or pathological injury to the bone in which the continuity of the bone is broken. | MedDRA:10017076:Fracture | Musculoskeletal System | Injury |  |  | E16¦E1603
|  | Hip Fracture | E160301 | Traumatic or pathological injury to the hip in which the continuity of either the femoral head, femoral neck, intertrochanteric or subtrochanteric regions is broken. | MedDRA:10020100:Hip fracture | Musculoskeletal System | Injury |  |  | E16¦E1603¦E160301
|  | Limb Fracture | E160302 | A traumatic break in one or more of the bones in a limb. | MedDRA:10074551:Limb fracture | Musculoskeletal System | Injury |  |  | E16¦E1603¦E160302
|  | Multiple Fractures | E160303 | Two or more simultaneous traumatic breaks in one or more bones. | MedDRA:10028200:Multiple fractures | Musculoskeletal System | Injury |  |  | E16¦E1603¦E160303
|  | Skull Fracture | E160304 | A traumatic or pathologic injury to the bones of the skull in which the continuity of the bones of the skull is broken. | MedDRA:10061365:Skull fracture | Musculoskeletal System | Injury |  |  | E16¦E1603¦E160304
|  | Vertebral Fracture | E160305 | Traumatic or pathological injury to the spine in which the continuity of a vertebral bone is broken. | MedDRA:10047333:Vertebral fracture | Musculoskeletal System | Injury |  |  | E16¦E1603¦E160305
| Connective Tissue Disease |  | E1604 | A non-neoplastic or neoplastic disorder that affects the connective tissue. | MedDRA:10061087:Connective tissue disorder |  |  |  |  | E16¦E1604
| Cramp(s) /Muscle Spasm(s) |  | E1605 | A sustained, sudden and involuntary contraction of a muscle or group of muscles. | MedDRA:10011285:Cramp |  |  |  |  | E16¦E1605
|  | Abdominal Cramps | E160501 | An involuntary muscular contraction involving a muscle of the abdomen or of a hollow organ within the abdomen. | MedDRA:10000057:Abdominal cramps |  |  |  |  | E16¦E1605¦E160501
| Damage to Ligament(s) |  | E1606 | An injury to any of the fibrous bands of tissue that comprise an articulation (joint). | MedDRA:10061223:Ligament injury | Musculoskeletal System | Injury |  |  | E16¦E1606
| Difficulty Chewing |  | E1607 | Problem with mastication due to pain, discomfort or loss of mobility in the jaw. | MedDRA:10067464:Chewing difficulty | Musculoskeletal System | Gastrointestinal System |  |  | E16¦E1607
| Fasciitis |  | E1608 | Inflammation of fascia. | MedDRA:10016228:Fasciitis |  |  |  |  | E16¦E1608
| Fibromyositis |  | E1609 | Inflammation and fibrous degeneration of a muscle. | MedDRA:10063041:Fibromyositis | Musculoskeletal System | Infections |  |  | E16¦E1609
| Head Injury |  | E1610 | A traumatic injury to the head. | MedDRA:10019196:Head injury | Musculoskeletal System | Injury |  |  | E16¦E1610
| Hyperextension |  | E1611 | Forceful extension of a joint beyond its normal limits. | MedDRA:10062310:Joint hyperextension |  |  |  |  | E16¦E1611
| Inadequate Osseointegration |  | E1612 | Insufficient structural and functional connection between living bone and the surface of a synthetic implant. | MedDRA:10068157:Inadequate osteointegration |  |  |  |  | E16¦E1612
|  | Aseptic Loosening | E161201 | Failure of the fixation of a prosthetic component in the absence of infection. | MedDRA:10090445:Aseptic joint prosthesis loosening |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 69 of Change Log. Editorial change to term for Release Number 2026. See no. 275 of Change Log. | E16¦E1612¦E161201
|  | Septic Loosening | E161202 | Failure of the fixation of a prosthetic component in the presence of infection. | MedDRA:10090446:Septic joint prosthesis loosening |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 70 of Change Log. Editorial change to term for Release Number 2026. See no. 276 of Change Log. | E16¦E1612¦E161202
| Intervertebral Disc Compression or Protrusion |  | E1613 | Degenerative, pathological or traumatic changes in the cartilaginous tissue of the spine leading to unwanted compression or protrusion of the disc. | MedDRA:10055039:Intervertebral disc compression |  |  |  |  | E16¦E1613
| Joint Dislocation |  | E1614 | An abnormal separation where two or more bones meet. | MedDRA:10023204:Joint dislocation | Musculoskeletal System | Injury |  |  | E16¦E1614
|  | Subluxation | E161401 | Partial dislocation of the bones in a joint. | MedDRA:10042402:Subluxation | Musculoskeletal System | Injury |  |  | E16¦E1614¦E161401
| Joint Laxity |  | E1615 | Looseness or instability of a joint. | MedDRA:10070874:Joint laxity |  |  |  |  | E16¦E1615
| Loss of Range of Motion |  | E1616 | A finding that the flexibility of a joint is below the expected range of normal for that individual. | MedDRA:10048706:Joint range of motion decreased |  |  |  |  | E16¦E1616
|  | Ankylosis | E161601 | Immobility of the joint due to the formation of bone, cartilaginous or fibrous tissues around the joints. | MedDRA:10002559:Ankylosis |  |  |  |  | E16¦E1616¦E161601
|  | Joint Contracture | E161602 | Permanent restriction of a joint, usually resulting from prolonged spasticity in a muscle area. | MedDRA:10023201:Joint contracture |  |  |  |  | E16¦E1616¦E161602
|  | Muscular Rigidity | E161603 | An involuntary, persistent state of firm, tense muscles with marked resistance to passive movement. | MedDRA:10028330:Muscle rigidity |  |  |  |  | E16¦E1616¦E161603
|  | Arthrofibrosis | E161604 | Fibrotic joint disorder characterized by excessive collagen production and adhesions that result in restricted joint motion and pain. | MedDRA:10058029:Arthrofibrosis |  |  |  | Term added for Release Number 2024. See no. 71 of Change Log | E16¦E1616¦E161604
|  | Joint Locking | E161605 | Joint is blocked or becomes stuck. For Pseudo-locking use E1616 | MedDRA:10061258:Joint lock |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 72 of Change Log. Editorial change to term and definition for Release Number 2026. See no. 277 of Change Log. | E16¦E1616¦E161605
| Malunion of Bone |  | E1617 | Faulty healing of bone, resulting in improper anatomical alignment. | MedDRA:10026729:Malunion of fracture |  |  |  |  | E16¦E1617
| Metal Related Pathology |  | E1618 | Also known as Metallosis. Includes trunnionosis, aseptic fibrosis or local necrosis secondary to metal corrosion and release of wear debris. | MedDRA:10081986:Periprosthetic metallosis |  |  |  | Definition was modified on 27 January 2022 (Old definition: Aseptic fibrosis or local necrosis secondary to metal corrosion and release of wear debris. It's also known as Metallosis.). For details, see comment No. 114 / 203 of the Change Log (Release Number 2022). | E16¦E1618
| Muscle Hypotonia |  | E1619 | Abnormally low level of muscle tone. | MedDRA:10081180:Muscle hypotonia |  |  |  |  | E16¦E1619
| Muscle/Tendon Damage |  | E1620 | Any injury to muscle or tendon tissue including strains. | MedDRA:10028314:Muscle injury |  |  |  |  | E16¦E1620
| Muscle Weakness/Atrophy |  | E1621 | A reduction in the strength and/or mass of one or more muscles. | MedDRA:10028350:Muscle weakness | Musculoskeletal System | Nervous System |  | Term/Definition was modified on 27 January 2022 (Old term/definition: Muscle Weakness: A reduction in the strength of one or more muscles.) For details, see comment No. 120 of the Change Log (Release Number 2022). | E16¦E1621
| Myalgia |  | E1622 | Pain sensation originating from a muscle or group of muscles. | MedDRA:10028411:Myalgia | Musculoskeletal System | Generalized Disorders |  |  | E16¦E1622
| Neck Pain |  | E1623 | Pain in the neck area. | MedDRA:10028836:Neck pain | Musculoskeletal System | Generalized Disorders |  |  | E16¦E1623
| Neck Stiffness |  | E1624 | Limited mobility of the neck, usually accompanied by pain. | MedDRA:10028843:Neck stiffness |  |  |  |  | E16¦E1624
| Nonunion/Delayed-Union Bone Fracture |  | E1625 | Broken bone that doesn't heal or is delayed in healing. | MedDRA:10017081:Fracture delayed union |  |  | Modified (editorial) | Term/Definition was modified on 27 January 2022 (Old term/definition: Nonunion Bone Fracture: Broken bone that doesn't heal.) MedDRA mapping also updated from Nonunion of fracture(10029746) to Fracture delayed union(10017081). For details, see comment No. 117 of the Change Log (Release Number 2022). Editorial change to term for Release Number 2026. See no. 278 of Change Log. | E16¦E1625
| Extraskeletal Ossification |  | E1626 | The formation of bone or of a bony substance outside the normal skeleton. Includes ectopic or heterotopic ossification, myositis ossificans. | MedDRA:10049811:Extraskeletal ossification |  |  | Modified (editorial) | Term and Definition and MedDRA map modified for Release Number 2024. Previous term was Ossification. Previous MedDRA map was 10065453:Ectopic ossification. See no. 73 of Change Log. Editorial change to term for Release Number 2026. See no. 279 of Change Log. | E16¦E1626
| Osteolysis |  | E1627 | Dissolution of bone; applied especially to the removal or loss of the calcium of bone. | MedDRA:10031248:Osteolysis |  |  |  |  | E16¦E1627
| Osteomyelitis |  | E1628 | Infection of bone or bone marrow. | MedDRA:10031252:Osteomyelitis | Musculoskeletal System | Infections |  |  | E16¦E1628
| Osteopenia/ Osteoporosis |  | E1629 | Decreased calcification or density of bone tissue. | MedDRA:10049088:Osteopenia |  |  |  |  | E16¦E1629
| Polymyositis |  | E1630 | An idiopathic inflammatory disorder affecting the muscles. It presents with symmetrical proximal muscle weakness and elevated skeletal muscle enzymes. | MedDRA:10036102:Polymyositis |  |  |  |  | E16¦E1630
| Spinal Column Injury |  | E1631 | Damage to the bones of the spine. | MedDRA:10069689:Spinal column injury | Musculoskeletal System | Injury |  |  | E16¦E1631
| Synovitis |  | E1632 | Inflammation of a synovial membrane. | MedDRA:10042868:Synovitis | Musculoskeletal System | Infections |  |  | E16¦E1632
| Tooth Fracture |  | E1633 | Breakage of any portion of the tooth. | MedDRA:10062544:Tooth fracture | Musculoskeletal System | Injury |  |  | E16¦E1633
| Unequal Limb Length |  | E1634 | Discrepancy between the lengths of the lower or upper extremities. | MedDRA:10065738:Unequal limb length |  |  |  |  | E16¦E1634
| Unspecified Musculoskeletal Problem |  | E1635 | The report describes a non-specific problem with the musculoskeletal system. Note: Please use E2402 if the report describes a specific problem but the relevant term does not exist. |  |  |  | Modified (editorial) | Term was added on 18 March 2020 for cases where the category of a health effect was reported but with no further details. Editorial change to term and definition for Release Number 2026. See no. 209 and 280 of Change Log. | E16¦E1635
| Tooth Loss |  | E1636 | One or more teeth lost/dislodged from the mouth. | MedDRA:10044044:Tooth loss |  |  |  | Term was added on 27 January 2022. For details, see comment No. 218 of the Change Log (Release Number 2022). | E16¦E1636
| Bursitis |  | E1637 | Inflammation of the fluid-filled pad (bursa). | MedDRA:10006811:Bursitis |  |  |  | Term was added on 27 January 2022. For details, see comment No. 112 of the Change Log (Release Number 2022). | E16¦E1637
| Cartilage Damage |  | E1638 | Damage of cartilage. | MedDRA:10007705:Cartilage damage |  |  | Modified (editorial) | Term was added on 27 January 2022. For details, see comment No. 113 of the Change Log (Release Number 2022). Editorial change to term for Release Number 2026. See no. 281 of Change Log. | E16¦E1638
| Joint Impingement |  | E1639 | Impingement relating to joints. Can be e.g. implant-on-bone, implant-on-soft tissue, or implant-on-implant impingement. | MedDRA:10087218:Joint impingement |  |  |  | Term was added on 27 January 2022. For details, see comment No. 119 of the Change Log (Release Number 2022). | E16¦E1639
| Mastoiditis |  | E1640 | An infection of the bony air cells in the mastoid bone. | MedDRA:10026900:Mastoiditis | Musculoskeletal System | Infections |  | Term was added on 27 January 2022. For details, see comment No. 241 of the Change Log (Release Number 2022). | E16¦E1640
| Limb Pain |  | E1641 | Pain in the arms and/or legs. | MedDRA:10033447:Pain in limb |  |  |  | Term was added on 27 January 2022. For details, see comment No. 57 / 187 of the Change Log (Release Number 2022). | E16¦E1641
| Spinal Deformity |  | E1642 | Abnormal curvature and/or alignment of the spine. | MedDRA:10058907:Spinal deformity |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 18 of Change Log. Editorial change to term for Release Number 2026. See no. 282 of Change Log. | E16¦E1642
| Tendinopathy |  | E1643 | Overuse tendon injuries characterized by a combination of pain, diffuse or localized swelling, and impaired performance. Problems include tendonitis/tendinitis and tendinosis, as well as tendon tears and inflammation of membranes around a tendon | MedDRA:10080130:Tendinopathy |  |  |  | Term added for Release Number 2024. See no. 68 of Change Log | E16¦E1643
| Compartment Syndrome |  | E1644 | Condition in which a closed muscle compartment's pressure increases to such an extent that the circulation is diminished, e.g. due to decrease in a compartment volume or an increase in the contents of a compartment, or both. | MedDRA:10010121:Compartment syndrome | Musculoskeletal System | Vascular System | Modified (editorial) | Term added for Release Number 2024. See no. 74 of Change Log. Editorial change to term for Release Number 2026. See no. 283 of Change Log. | E16¦E1644
| Hemarthrosis |  | E1645 | Bleeding into the joints | MedDRA:10019409:Hemarthrosis | Musculoskeletal System | Vascular System |  | Term added for Release Number 2024. See no. 75 of Change Log | E16¦E1645
| Bone Notching |  | E1646 | An indentation made at the edge of the bone without bone breakage. | MedDRA:10057328:Bone injury | Musculoskeletal System | Procedural Complications |  | Term added for Release Number 2024. See no. 76 of Change Log | E16¦E1646
| Exostosis (bone) |  | E1647 | Benign hypertrophy that projects outward from the surface of bone, often containing a cartilaginous component, such as bone spurs/osteophytes. | MedDRA:10015688:Exostosis |  |  |  | Term added for Release Number 2024. See no. 77 of Change Log | E16¦E1647
| Acetabular Protrusion |  | E1648 | Acetabular defect with a shift of the femoral head through the pelvic bone. | MedDRA:10090179:Acetabular protrusion |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 90 of Change Log. Editorial change to term for Release Number 2026. See no. 284 of Change Log. | E16¦E1648
| Crepitus |  | E1649 | A grating, popping, cracking, or clicking sound or sensation in a joint, typically produced by friction between structures such as bone, cartilage, or fractured parts of a bone. | MedDRA:10051374:Joint crepitation |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 91 of Change Log. Editorial change to definition for Release Number 2026. See no. 145 of Change Log. | E16¦E1649
| Arthropathy |  | E1650 | Diseases and disorders of a joint (e.g. Rotator cuff arthropathy, cuff tear arthropathy (CTA)) More specific codes and terms should be used where they exist e.g. E1602. | MedDRA:10003285:Arthropathy |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 103 of Change Log. Editorial change to definition for Release Number 2026. See no. 417 of Change Log. | E16¦E1650
| Back Pain |  | E1651 | Pain in any part of the back. | MedDRA:10003988:Back pain | Musculoskeletal System | Generalized Disorders | Modified (editorial) | Term added for Release Number 2024. See no. 107 of Change Log. Editorial change to term for Release Number 2026. See no. 285 of Change Log. | E16¦E1651
| Loose Body in Joint |  | E1652 | Presence of fibrous, bony, cartilaginous or osseocartilaginous fragments in a synovial joint. | MedDRA:10024829:Loose body in joint |  |  | Modified (editorial) | Term added for Release Number 2025. See no. 69 of Change Log. Editorial change to term for Release Number 2026. See no. 286 of Change Log. | E16¦E1652
| Patellar Tracking Disorder |  | E1653 | Abnormal alignment or movement of the patella relative to the trochlear groove during knee flexion; may be described as patellar instability or lateral patellar displacement | MedDRA:10077393:Patellar tracking disorder |  |  | New | New term added for Release Number 2026. See no. 19 of Change Log. | E16¦E1653
| Loss of fracture reduction |  | E2126 | Loss of anatomical alignment after fracture reduction. | MedDRA:10069723:Loss of anatomical alignment after fracture reduction | Procedural Complications | Musculoskeletal System |  | Term added for Release Number 2024. See no. 65 of Change Log. Editorial change to term for Release Number 2026. See no. 296 of Change Log. | E16¦E2126
| Periprosthetic fracture |  | E2127 | Fractures around joint replacement prosthetics or implants. These can occur intraoperatively or postoperatively. | MedDRA:10069135:Periprosthetic fracture | Procedural Complications | Musculoskeletal System |  | Term added for Release Number 2024. See no. 67 of Change Log | E16¦E2127
| Ambulation or Postural Difficulties |  | E2302 | Any persistent problem with walking, gait or maintaining posture (sitting, standing). | MedDRA:10047810:Walking difficulty | Generalized Disorders | Musculoskeletal System |  | Term/Definition was modified on 31 January 2023 (Old term/definition: Ambulation Difficulties: Any persistent problem with walking or gait.) For details, see comment No. 98 of the Change Log (Release Number 2023). | E16¦E2302
| Hyporeflexia |  | E0146 | Diminished or absent reflex responses | MedDRA:10021089:Hyporeflexia | Nervous System | Musculoskeletal | New | Term added for Release Number 2026. See no. 14 of Change Log. | E16¦E0146
Skin and Subcutaneous Tissue |  |  | E17 | Skin and Subcutaneous Tissue |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E17
| Abrasion |  | E1701 | Superficial damage to the skin caused by rubbing or scraping. | MedDRA:10080919:Abrasion | Skin and Subcutaneous Tissue | Injury |  |  | E17¦E1701
| Angioedema |  | E1702 | Rapid swelling of the deep layers of the skin due to transient vascular leakage of serous fluid. | MedDRA:10002424:Angioedema | Skin and Subcutaneous Tissue | Generalized Disorders |  |  | E17¦E1702
| Blister |  | E1703 | A collection of fluid underneath the top layer of skin. | MedDRA:10005191:Blister |  |  |  |  | E17¦E1703
| Burn(s) |  | E1704 | An injury or tissue damage caused by exposure to caustic chemicals, extreme heat, extreme cold or excessive radiation. | MedDRA:10006634:Burn | Skin and Subcutaneous Tissue | Injury |  |  | E17¦E1704
|  | Bowel Burn | E1005 | A thermal injury to the bowels. | MedDRA:10082501:Intestinal burn | Gastrointestinal System | Injury |  |  | E17¦E1704¦E1005
|  | Gastroesophageal Burn | E1013 | Thermal, chemical or radiation injury to stomach or esophagus. | MedDRA:10082500:Gastroesophageal burn | Gastrointestinal System | Injury |  |  | E17¦E1704¦E1013
|  | Caustic/Chemical Burns | E170401 | A burn caused by exposure to chemical substances. | MedDRA:10008420:Chemical burn | Skin and Subcutaneous Tissue | Injury |  |  | E17¦E1704¦E170401
|  | Cryogenic Burn | E170402 | Local tissue damage due to exposure to extreme cold. | MedDRA:10081985:Cold burn | Skin and Subcutaneous Tissue | Injury |  |  | E17¦E1704¦E170402
|  | Radiation Burn | E170403 | A burn caused by exposure to x-ray, radium, sunlight, ionizing radiation or any other type of radiant energy. | MedDRA:10063640:Radiation burn | Skin and Subcutaneous Tissue | Injury |  |  | E17¦E1704¦E170403
|  | Superficial (First Degree) Burn | E170404 | A burn that affects the epidermis only, causing erythema without blistering. | MedDRA:10016700:First degree burns | Skin and Subcutaneous Tissue | Injury |  |  | E17¦E1704¦E170404
|  | Partial Thickness (Second Degree) Burn | E170405 | A burn that affects the epidermis and the dermis. | MedDRA:10039798:Second degree burns | Skin and Subcutaneous Tissue | Injury | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 287 of Change Log. | E17¦E1704¦E170405
|  | Full Thickness (Third Degree) Burn | E170406 | A burn to the skin that extends into deeper layers and underlying tissues. | MedDRA:10043441:Third degree burns | Skin and Subcutaneous Tissue | Injury | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 288 of Change Log. | E17¦E1704¦E170406
| Burning Sensation |  | E1705 | A sensation of stinging or heat, not necessarily accompanied by redness or physical signs of irritation. | MedDRA:10006784:Burning sensation | Skin and Subcutaneous Tissue | Generalized Disorders |  |  | E17¦E1705
|  | Skin Burning Sensation | E170501 | A sensation of dermal stinging or heat, not necessarily accompanied by redness or physical signs of skin irritation. | MedDRA:10054786:Skin burning sensation | Skin and Subcutaneous Tissue | Generalized Disorders |  |  | E17¦E1705¦E170501
| Cancer Cells Dissemination |  | E1706 | Spreading benign or malignant tissue during medical procedures. | MedDRA:10082075:Iatrogenic metastasis |  |  |  |  | E17¦E1706
| Itching Sensation |  | E1708 | An irritating feeling that produces the desire to scratch. | MedDRA:10023084:Itching |  |  |  |  | E17¦E1708
| Jaundice |  | E1709 | Yellow pigmentation of the skin, mucous membranes, and the eyes due to hyperbilirubinemia. | MedDRA:10023126:Jaundice |  |  |  |  | E17¦E1709
| Localized Skin Lesion |  | E1710 | A pathologic process that affects the skin and is confined to a specific area. | MedDRA:10040882:Skin lesion |  |  |  |  | E17¦E1710
| Peeling |  | E1711 | A peeling off or loss of epidermis, as in sunburn, postscarlatinal peeling, or toxic epidermal necrolysis. | MedDRA:10034218:Peeling |  |  |  |  | E17¦E1711
| Phototoxicity |  | E1712 | A nonimmunologic, chemically induced type of photosensitivity. | MedDRA:10034977:Phototoxicity |  |  |  |  | E17¦E1712
| Pocket Erosion |  | E1713 | Erosion of patient's tissue pocket which houses a device. | MedDRA:10035771:Pocket erosion | Skin and Subcutaneous Tissue | Injury |  |  | E17¦E1713
| Rash |  | E1714 | An eruption in the skin which effects its appearance and/or texture. | MedDRA:10037844:Rash |  |  |  |  | E17¦E1714
| Scar Tissue |  | E1715 | Formation of new tissue formed in the healing of a wound. | MedDRA:10039580:Scar |  |  |  |  | E17¦E1715
|  | Abnormal Scarring | E171501 | Excessive or unusual scarring related to impaired wound healing (e.g. Keloid or Hypertrophic scar) | MedDRA:10080564:Abnormal scarring |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 108 and 109 of Change Log. Editorial change to term for Release Number 2026. See no. 289 of Change Log. | E17¦E1715¦E171501
| Skin Discoloration |  | E1716 | any abnormal change in skin coloration. | MedDRA:10040828:Skin discoloration |  |  |  |  | E17¦E1716
|  | Erythema | E171601 | Red discoloration of the skin. | MedDRA:10015150:Erythema |  |  |  |  | E17¦E1716¦E171601
|  | Hyperpigmentation | E171602 | Darkening of the skin due to excessive melanin deposition. | MedDRA:10040865:Skin hyperpigmentation |  |  |  | Term was added on 27 January 2022. For details, see comment No. 105 of the Change Log (Release Number 2022). | E17¦E1716¦E171602
|  | Hypopigmentation | E171603 | Abnormal lightening of skin due to decreased melanin production or deposition. | MedDRA:10040868:Skin hypopigmentation |  |  |  | Term was added on 27 January 2022. For details, see comment No. 106 of the Change Log (Release Number 2022). | E17¦E1716¦E171603
| Skin Disorders |  | E1717 | Affecting the skin. | MedDRA:10040831:Skin disorder |  |  |  |  | E17¦E1717
| Skin Erosion |  | E1718 | A gradual breakdown or very shallow ulceration of the skin which involves only the epidermis and heals without scarring. | MedDRA:10040840:Skin erosion | Skin and Subcutaneous Tissue | Injury |  |  | E17¦E1718
| Skin Infection |  | E1719 | An inflammatory process affecting the skin, caused by bacteria, viruses, parasites, or fungi. | MedDRA:10040872:Skin infection | Skin and Subcutaneous Tissue | Infections |  |  | E17¦E1719
| Skin Inflammation/ Irritation |  | E1720 | An inflammatory process affecting the skin. Signs include red rash, itching, and blister formation. Representative examples are contact dermatitis, atopic dermatitis, and seborrheic dermatitis. | MedDRA:10062249:Skin inflammation |  |  |  |  | E17¦E1720
|  | Abscess | E172001 | An inflammatory process characterized by the accumulation of pus within a newly formed tissue cavity. | MedDRA:10000269:Abscess | Skin and Subcutaneous Tissue | Infections |  |  | E17¦E1720¦E172001
|  | Cellulitis | E172002 | Inflammation of the dermis and subcutaneous tissues caused by a bacterial infection. | MedDRA:10007882:Cellulitis | Skin and Subcutaneous Tissue | Infections |  |  | E17¦E1720¦E172002
|  | Contact Dermatitis | E172003 | An inflammatory skin condition caused by direct contact between the skin and either an irritating substance or an allergen. | MedDRA:10056540:Dermatitis irritant contact |  |  |  |  | E17¦E1720¦E172003
|  | Eczema | E172004 | A form of dermatitis characterized by red, itchy, scaly, or crusty patches that can be chronic or intermittent. | MedDRA:10014184:Eczema |  |  |  |  | E17¦E1720¦E172004
|  | Radiodermatitis | E172005 | A cutaneous inflammatory reaction occurring as a result of exposure to biologically effective levels of ionizing radiation. | MedDRA:10061103:Dermatitis radiation |  |  |  |  | E17¦E1720¦E172005
|  | Urticaria | E172006 | A transient, itchy skin eruption characterized by wheals with pale interiors and red margins. | MedDRA:10046735:Urticaria |  |  |  |  | E17¦E1720¦E172006
|  | Wheal(s) | E172007 | A smooth, slightly elevated area on the body surface, which is redder or paler than the surrounding skin. | MedDRA:10047918:Wheals |  |  |  |  | E17¦E1720¦E172007
|  | Granuloma | E2317 | An inflammatory reaction usually caused by infectious organisms, foreign bodies, or cholesterol deposits. It is characterized by the presence of epithelioid histiocytes and chronic inflammation. | MedDRA:10018691:Granuloma | Generalized Disorders | Skin and Subcutaneous Tissue |  |  | E17¦E1720¦E2317
| Skin Tears |  | E1721 | A break in skin tissue caused by friction or shearing. | MedDRA:10048704:Skin tear |  |  |  |  | E17¦E1721
| Subcutaneous Nodule |  | E1722 | A small palpable abnormal mass in the subcutaneous tissues. | MedDRA:10042348:Subcutaneous nodule |  |  |  |  | E17¦E1722
| Wrinkling |  | E1723 | A folding, ridge or creasing of the skin. | MedDRA:10048043:Wrinkling |  |  |  |  | E17¦E1723
| Dry Skin |  | E1724 | Skin that has become dry in texture and flaky. | MedDRA:10013786:Dry skin |  |  |  | Term was added on 27 January 2022. For details, see comment No. 101 of the Change Log (Release Number 2022). | E17¦E1724
| Administration Site Swelling |  | E1725 | Liquid or a pocket of air trapped under the skin (e.g. pooling), typically due to infusion or injection problems. Does not include inflammation. | MedDRA:10075107:Administration site swelling |  |  | Modified (technical) | Term was added on 27 January 2022. For details, see comment No. 200 / 201 of the Change Log (Release Number 2022). Technical change for Release Number 2026. See no. 97 of Change Log. | E17¦E1725
| Skin Maceration |  | E1726 | Softening of tissue by the action of a liquid. Occurs at the wound site/surrounding areas due to excessive moisture from exudate or products applied. | MedDRA:10048625:Skin maceration |  |  |  | Term was added on 27 January 2022. For details, see comment No. 104 of the Change Log (Release Number 2022). | E17¦E1726
| Wound Odor |  | E1727 | Malodor of the wound or skin not otherwise explained. Can include odors due to necrosis, bioburden or other barriers to wound healing. | MedDRA:10069389:Wound odor |  |  |  | Term was added on 31 January 2023. For details, see comment No. 94 of the Change Log (Release Number 2023). | E17¦E1727
| Nail Disorder |  | E1728 | Abnormality of the nail, such as problems with the color, shape, texture, or thickness of the fingernails or toenails. | MedDRA:10028694:Nail disorder |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 12 of Change Log. Editorial change to term for Release Number 2026. See no. 290 of Change Log. | E17¦E1728
| Unspecified Skin and Subcutaneous Tissue Problem |  | E1729 | The report describes a non-specific problem with the skin or subcutaneous tissue. Note: Please use E2402 if the report describes a specific problem but the relevant term does not exist. |  |  |  | Modified (editorial) | Term added for Release Number 2025. See no. 74 of Change Log. Editorial change to definition for Release Number 2026. See no. 210 of Change Log. | E17¦E1729
| Acne |  | E1730 | Clogged (blackheads or whiteheads), inflamed, or infected sebaceous glands in the skin, characterized by painful pus-filled and/or red papules and pustules which may progress to cysts and nodules | MedDRA:10000496:Acne |  |  | New | Term added for Release Number 2026. See no. 96 of Change Log. | E17¦E1730
| Subcutaneous Emphysema |  | E1731 | Presence of gas or air within subcutaneous or deeper soft tissues; may be described as soft tissue emphysema or subdermal gas accumulation | MedDRA:10042344:Subcutaneous emphysema |  |  | New | Term added for Release Number 2026. See no. 16 of Change Log. | E17¦E1731
| Telangiectasia |  | E0526 | Telangiectasia, also known as Spider veins or thread veins, are prominent clusters of small, damaged blood vessels visible in the skin. They appear as thin red, blue, or purple lines. | MedDRA:10043189:Telangiectasia | Vascular System | Skin | New | Term added for Release Number 2026. See no. 67 of Change Log | E17¦E0526
| Easy Bruising |  | E2005 | Appearance of bruises spontaneously or following minimal contact. | MedDRA:10076258:Tendency to bruise easily | Injury | Skin and Subcutaneous Tissue |  |  | E17¦E2005
Neoplasms Benign, Malignant and Unspecified |  |  | E18 | Neoplasms Benign, Malignant and Unspecified |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E18
| Cancer |  | E1801 | A group of diseases in which abnormal cells divide without control, invade nearby tissues and may also spread to other parts of the body through the blood and lymph systems. | MedDRA:10007050:Cancer |  |  |  |  | E18¦E1801
|  | Breast Cancer | E180101 | A malignant neoplasm that develops or arises in breast tissue. | MedDRA:10006187:Breast cancer |  |  |  |  | E18¦E1801¦E180101
|  | Breast Implant Associated Anaplastic Large Cell Lymphoma (BIA ALCL) | E180102 | A non-Hodgkin lymphoma which is associated with breast implants, and consists of large anaplastic cells. The WHO diagnostic criteria for BIA-ALCL includes positive expression of CD30 (CD30+) and negativity for anaplastic lymphoma kinase (ALK-). | MedDRA:10082495:Breast implant-associated anaplastic large cell lymphoma |  |  |  | Definition was modified on 27 January 2022 (Old definition: A peripheral (mature) T-cell lymphoma, consisting of usually large anaplastic, CD30 positive cells. The majority of cases are positive for the anaplastic large cell lymphoma (ALK) protein.). For details, see comment No. 9 of the Change Log (Release Number 2022). | E18¦E1801¦E180102
|  | Leiomyosarcoma | E180103 | A malignant tumor of smooth muscle cells that can arise almost anywhere in the body but is most common in the uterus, abdomen, or pelvis. | MedDRA:10024189:Leiomyosarcoma |  |  |  |  | E18¦E1801¦E180103
|  | Lymphoma | E180104 | A malignant (clonal) proliferation of B- lymphocytes or T- lymphocytes which involves the lymph nodes, bone marrow and/or extranodal sites. This category includes Non-Hodgkin lymphomas and Hodgkin lymphomas. | MedDRA:10025310:Lymphoma |  |  |  |  | E18¦E1801¦E180104
| Cyst(s) |  | E1802 | A sac-like closed membranous structure that may be empty or contain fluid or amorphous material. | MedDRA:10011732:Cyst |  |  |  |  | E18¦E1802
| Nodule |  | E1803 | A small lump, swelling or collection of tissue. | MedDRA:10054107:Nodule |  |  |  |  | E18¦E1803
| Solid Tumor |  | E1804 | Abnormal mass of tissue that usually does not contain cysts or liquid areas. Solid tumors may be benign (not cancerous), or malignant (cancerous). Different types of solid tumors are named for the type of cells that form them. If tumors are associated with those under the term cancer, these terms should be used. | MedDRA:10065252:Solid tumor |  |  | Modified (editorial) | Editorial change to term and definition for Release Number 2026. See no. 190 of Change Log. | E18¦E1804
| Non-Malignant Neoplasm |  | E1805 | A new and abnormal growth of tissue in a part of the body which is not indicative of malignant disorder. Can include benign tumors. | MedDRA:10060999:Benign neoplasm |  |  | Modified (editorial) | Term was added on 27 January 2022. For details, see comment No. 109 of the Change Log (Release Number 2022). Editorial change to term for Release Number 2026. See no. 291 of Change Log. | E18¦E1805
Infections |  |  | E19 | Infections |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E19
| Encephalitis |  | E0114 | An acute inflammatory process affecting the brain parenchyma. | MedDRA:10014581:Encephalitis | Nervous System | Infections |  |  | E19¦E0114
| Meningitis |  | E0121 | A disorder characterized by acute inflammation of the meninges of the brain and/or spinal cord. | MedDRA:10027199:Meningitis | Nervous System | Infections |  |  | E19¦E0121
| Sepsis |  | E0306 | The presence of pathogenic microorganisms in the blood stream causing a rapidly progressing systemic reaction that may lead to shock. Symptoms include fever, chills, tachycardia, and increased respiratory rate. It is a medical emergency that requires urgent medical attention. | MedDRA:10040047:Sepsis | Blood and Lymphatic System | Infections |  |  | E19¦E0306
| Endocarditis |  | E0610 | Inflammation of the endocardium. | MedDRA:10014665:Endocarditis | Heart | Infections |  |  | E19¦E0610
| Pneumonia |  | E0733 | An acute and/or chronic inflammation focally or diffusely affecting the lung parenchyma. | MedDRA:10035664:Pneumonia | Respiratory System | Infections |  |  | E19¦E0733
|  | Bronchopneumonia | E073301 | Acute inflammation of the walls of the terminal bronchioles. | MedDRA:10006469:Bronchopneumonia | Respiratory System | Infections |  |  | E19¦E0733¦E073301
| Respiratory Tract Infection |  | E0744 | An infectious disease affecting the respiratory tract. | MedDRA:10062352:Respiratory tract infection | Respiratory System | Infections |  |  | E19¦E0744
| Chemosis |  | E0803 | Edema of the mucous membrane of the eyeball and eyelid lining. | MedDRA:10008435:Chemosis | Eye | Infections |  |  | E19¦E0803
| Endophthalmitis |  | E0816 | Inflammation of the intraocular fluids or tissues usually with microbial organisms. | MedDRA:10014801:Endophthalmitis | Eye | Infections |  |  | E19¦E0816
| Eye Infections |  | E0818 | An infection of any part of the eye, including the eyelid. | MedDRA:10015929:Eye infection | Eye | Infections |  |  | E19¦E0818
|  | Intraocular Infection | E081801 | Infection within the eye. | MedDRA:10054762:Eye infection intraocular | Eye | Infections |  |  | E19¦E0818¦E081801
| Pancreatitis |  | E1021 | Inflammation of the pancreas. | MedDRA:10033645:Pancreatitis | Gastrointestinal System | Infections |  |  | E19¦E1021
| Peritonitis |  | E1024 | Inflammation of the peritoneum due to infection by bacteria or fungi. | MedDRA:10034674:Peritonitis | Gastrointestinal System | Infections |  |  | E19¦E1024
| Hepatitis |  | E1102 | Inflammation of the liver; usually from a viral infection, but sometimes from toxic agents. | MedDRA:10056522:Hepatic infection | Hepatic and Biliary System | Infections |  |  | E19¦E1102
| Urinary Tract Infection |  | E1310 | An inflammatory process affecting any part of the urinary tract, most commonly the bladder or the urethra. | MedDRA:10046571:Urinary tract infection | Kidney and Urinary Tract | Infections |  |  | E19¦E1310
|  | Kidney Infection | E131001 | An inflammatory process affecting the kidney. | MedDRA:10023424:Kidney infection | Kidney and Urinary Tract | Infections |  |  | E19¦E1310¦E131001
| Toxic Shock Syndrome |  | E1414 | A rare acute life-threatening systemic bacterial noncontagious illness caused by staphylococcal exotoxins. | MedDRA:10044248:Toxic shock syndrome | Reproductive System and Breast | Infections |  |  | E19¦E1414
| Chorioamnionitis |  | E1504 | An infection involving chorion, amnion and amniotic fluid; generally, there is also involve placental villi and from decidua. | MedDRA:10008755:Chorioamnionitis | Pregnancy, Childbirth and the Puerperium | Infections |  |  | E19¦E1504
| Fibromyositis |  | E1609 | Inflammation and fibrous degeneration of a muscle. | MedDRA:10063041:Fibromyositis | Musculoskeletal System | Infections |  |  | E19¦E1609
| Osteomyelitis |  | E1628 | Infection of bone or bone marrow. | MedDRA:10031252:Osteomyelitis | Musculoskeletal System | Infections |  |  | E19¦E1628
| Synovitis |  | E1632 | Inflammation of a synovial membrane. | MedDRA:10042868:Synovitis | Musculoskeletal System | Infections |  |  | E19¦E1632
| Skin Infection |  | E1719 | An inflammatory process affecting the skin, caused by bacteria, viruses, parasites, or fungi. | MedDRA:10040872:Skin infection | Skin and Subcutaneous Tissue | Infections |  |  | E19¦E1719
| Abscess |  | E172001 | An inflammatory process characterized by the accumulation of pus within a newly formed tissue cavity. | MedDRA:10000269:Abscess | Skin and Subcutaneous Tissue | Infections |  |  | E19¦E172001
| Cellulitis |  | E172002 | Inflammation of the dermis and subcutaneous tissues caused by a bacterial infection. | MedDRA:10007882:Cellulitis | Skin and Subcutaneous Tissue | Infections |  |  | E19¦E172002
| Bacterial Infection |  | E1901 | An acute infectious disorder caused by gram positive or gram negative bacteria. | MedDRA:10060945:Bacterial infection |  |  |  |  | E19¦E1901
|  | Drug Resistant Bacterial Infection | E190101 | An infection that is resistant to antibiotics and antimicrobials, for example, MRSA or VRE. | MedDRA:10004052:Bacterial resistance |  |  |  |  | E19¦E1901¦E190101
|  | Pyogenic Infection | E190102 | An infection caused by pus-producing organisms. | MedDRA:10021861:Infection pyogenic |  |  |  |  | E19¦E1901¦E190102
|  | Toxemia | E190103 | The condition resulting from the spread of bacterial products (toxins) by the bloodstream. | MedDRA:10059857:Toxemia | Infections | Pregnancy, Childbirth and the Puerperium |  |  | E19¦E1901¦E190103
| Fungal Infection |  | E1902 | An infection caused by a fungus. | MedDRA:10017533:Fungal infection |  |  |  |  | E19¦E1902
| Sexually Transmitted Infection |  | E1903 | A disease acquired predominantly through sexual contact. | MedDRA:10068720:Sexually transmitted infection |  |  |  |  | E19¦E1903
| Subclinical Infection |  | E1904 | Infection associated with no detectable signs or symptoms but caused by microorganisms capable of producing easily recognizable diseases. | MedDRA:10068163:Subclinical infection |  |  |  | The definition was changed on 21 January 2021 (Old definition: Infection associated with no detectable symptoms but caused by microorganisms capable of producing easily recognizable diseases.). For details, refer to comment No. 8 of the Change Log). | E19¦E1904
| Transmissible Spongiform Encephalopathy(TSE) |  | E1905 | A group of rare degenerative brain disorders associated with prions for example Creutzfeldt-Jakob disease (CJD) or Bovine Spongiform Encephalopathy (BSE). | MedDRA:10080750:Transmissible spongiform encephalopathy |  |  |  | Editorial change to the definition on 16 February 2021. | E19¦E1905
| Unspecified Infection |  | E1906 | The report describes a non-specific or unidentified infection. Note: Please use the appropriate term if the report describes a specific infection. i.e. E1901, E1902, E1907, or a relevant Level 3 term where more detail is available. | MedDRA:10021789:Infection |  |  | Modified (editorial) | Editorial change to the definition on 18 March 2020 for consistency with other "Unspecified ____" definitions. Editorial change to definition for Release Number 2026. See no. 219 of Change Log | E19¦E1906
| Viral Infection |  | E1907 | Any disease caused by a virus. | MedDRA:10047461:Viral infection |  |  |  |  | E19¦E1907
|  | Cytomegalovirus (CMV) Infection | E190701 | A genus of the family herpesviridae, subfamily betaherpesvirinae, infecting the salivary glands, liver, spleen, lungs, eyes, and other organs, in which they produce characteristically enlarged cells with intranuclear inclusions. | MedDRA:10011831:Cytomegalovirus infection |  |  |  |  | E19¦E1907¦E190701
| Wound Infection |  | E2012 | Infection of a wound. This does not include post surgical wound infection. | MedDRA:10048038:Wound Infection | Injury | Infections |  | Term/Definition was modified on 31 January 2023 (Old term/definition: Post Traumatic Wound Infection: Infection of a wound following trauma. This does not include post surgical wound infection.) MedDRA mapping also updated from Post-traumatic wound infection (10036317) to Wound Infection (10048038). For details, see comment No. 95 of the Change Log (Release Number 2023). | E19¦E2012
| Post Operative Wound Infection |  | E2115 | Infection of a surgical skin incision. | MedDRA:10036410:Postoperative wound infection | Procedural Complications | Infections |  |  | E19¦E2115
| Bacteremia |  | E2205 | Condition that relates to the presence of viable bacteria in the bloodstream, whether associated with active disease or not. | MedDRA:10003999:Bacteremia | Investigations and Diagnostic Tests | Infections |  | Term was added on 27 January 2022. For details, see comment No. 49 of the Change Log (Release Number 2022). | E19¦E2205
| Sinusitis |  | E2122 | Inflammation (swelling and redness) of the paranasal sinuses, associated or related to implantation procedures or devices. | MedDRA:10040753:Sinusitis | Procedural Complications | Infections |  | Term was added on 27 January 2022. For details, see comment No. 240 of the Change Log (Release Number 2022). | E19¦E2122
| Medical device site infection |  | E2123 | Infection local to an implanted or invasive medical device. | MedDRA:10075580:Medical device site infection | Procedural Complications | Infections |  | Term was added on 27 January 2022. For details, see comment No. 122 of the Change Log (Release Number 2022). Term and Definition and MedDRA map modified for Release Number 2024. Previous term was Peri-prosthetic Joint Infection. Previous MedDRA map was 10076118:Medical device site joint infection. See no. 10 of Change Log. Editorial change to term for Release Number 2026. See no. 294 of Change Log. | E19¦E2123
| Ear Infection |  | E0908 | Infection of the ear. | MedDRA:10014011:Ear infection | Ear and Labyrinth | Infections | Modified (editorial) | Term was added on 27 January 2022. For details, see comment No. 223 of the Change Log (Release Number 2022). Editorial change to definition for Release Number 2026. See no. 268 of Release Log. | E19¦E0908
|  | Infectious Otitis Externa | E090801 | An infectious process involving the outer ear and ear canal. | MedDRA:10033072:Otitis externa | Ear and Labyrinth | Infections |  | Editorial change to the term on 31 January 2023. | E19¦E0908¦E090801
|  | Infectious Otitis Media | E090802 | An infectious process involving the middle ear, includes acute otitis media, chronic otitis media, chronic suppurative otitis media, otitis media with effusion (OME), etc. | MedDRA:10033078:Otitis media | Ear and Labyrinth | Infections |  | Editorial change to the term on 31 January 2023. | E19¦E0908¦E090802
| Labyrinthitis |  | E0910 | An inflammation of the labyrinth in the inner ear that causes dizziness and reduced or distorted hearing. | MedDRA:10023567:Labyrinthitis | Ear and Labyrinth | Infections |  | Term was added on 27 January 2022. For details, see comment No. 233 of the Change Log (Release Number 2022). | E19¦E0910
| Mastoiditis |  | E1640 | An infection of the bony air cells in the mastoid bone. | MedDRA:10026900:Mastoiditis | Musculoskeletal System | Infections |  | Term was added on 27 January 2022. For details, see comment No. 241 of the Change Log (Release Number 2022). | E19¦E1640
Injury |  |  | E20 | Injury |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E20
| Brain Injury |  | E0102 | Damage to the brain. | MedDRA:10060690:Traumatic brain injury | Nervous System | Injury |  |  | E20¦E0102
|  | Encephalocele | E010201 | Hernia of brain substance and meninges through a congenital or traumatic opening of the skull. | MedDRA:10014617:Encephalocele | Nervous System | Injury |  |  | E20¦E0102¦E010201
| Concussion |  | E0108 | Traumatic brain injury as a result of the action of a mechanical force on the head. | MedDRA:10010254:Concussion | Nervous System | Injury |  |  | E20¦E0108
| Nerve Damage |  | E0123 | Injury to nervous tissue. | MedDRA:10029177:Nerve damage | Nervous System | Injury |  |  | E20¦E0123
|  | Optical Nerve Damage | E012301 | Injury to optical nerve. | MedDRA:10030938:Optic nerve injury | Nervous System | Injury |  |  | E20¦E0123¦E012301
| Nervous System Injury |  | E0124 | Injury to the central or peripheral nervous system. | MedDRA:10080898:Nervous system injury | Nervous System | Injury |  |  | E20¦E0124
|  | Spinal Cord Injury | E012401 | Traumatic damage of the spinal cord. | MedDRA:10041552:Spinal cord injury | Nervous System | Injury |  |  | E20¦E0124¦E012401
| Peripheral Nervous Injury |  | E0128 | Injury to the peripheral nervous system. | MedDRA:10034586:Peripheral nerve injury | Nervous System | Injury |  |  | E20¦E0128
| Choking |  | E0710 | Interference with respiration by obstruction of the larynx or trachea. | MedDRA:10008589:Choking | Respiratory System | Injury |  | Editorial change to the definition on 16 February 2021. | E20¦E0710
| Strangulation |  | E0749 | Interference with respiration by compression of the larynx, laryngopharynx or trachea, as well as of the carotid arteria or jugular veins. | MedDRA:10042169:Strangulation | Respiratory System | Injury |  |  | E20¦E0749
| Eye Injury |  | E0819 | Damage to the eye. | MedDRA:10061128:Eye injury | Eye | Injury |  |  | E20¦E0819
|  | Corneal Abrasion | E081901 | Injury to the epithelium of the cornea. | MedDRA:10010984:Corneal abrasion |  |  |  |  | E20¦E0819¦E081901
|  | Eye Burn | E081902 | Burn injury, including chemical burn, to any part of the eye. | MedDRA:10015911:Eye burns | Eye | Injury |  |  | E20¦E0819¦E081902
|  | Retinal Detachment | E081903 | Separation of the inner retina layers from the underlying pigment epithelium. | MedDRA:10038848:Retinal detachment |  |  |  |  | E20¦E0819¦E081903
|  | Retinal Injury | E081904 | Damage to the retina due to toxicity, exposure to bright lights, or trauma. | MedDRA:10057430:Retinal injury | Eye | Injury |  |  | E20¦E0819¦E081904
|  | Retinal Tear | E081905 | A usually small tearing of the retina occurring when the vitreous separates from the retina. It may lead to retinal detachment. | MedDRA:10038897:Retinal tear |  |  |  |  | E20¦E0819¦E081905
|  | Vitreous Detachment | E081906 | Detachment of the vitreous humor from the retina. | MedDRA:10047650:Vitreous detachment | Eye | Procedural Complications |  |  | E20¦E0819¦E081906
| Ear Injury |  | E0912 | Injury to the outer, middle or inner ear. | MedDRA:10057446:Ear injury | Ear and Labyrinth | Injury |  | Term added for Release Number 2025. See no. 76 of Change Request Log. Editorial change to term for Release Number 2026. See no. 269 of Change Log | E20¦E0912
| Peritoneal Laceration(s) |  | E1023 | A cut or tear of the peritoneum. | MedDRA:10068140:Peritoneal laceration | Gastrointestinal System | Injury |  |  | E20¦E1023
| Bone Fracture(s) |  | E1603 | A traumatic or pathological injury to the bone in which the continuity of the bone is broken. | MedDRA:10017076:Fracture | Musculoskeletal System | Injury |  |  | E20¦E1603
|  | Hip Fracture | E160301 | Traumatic or pathological injury to the hip in which the continuity of either the femoral head, femoral neck, intertrochanteric or subtrochanteric regions is broken. | MedDRA:10020100:Hip fracture | Musculoskeletal System | Injury |  |  | E20¦E1603¦E160301
|  | Limb Fracture | E160302 | A traumatic break in one or more of the bones in a limb. | MedDRA:10074551:Limb fracture | Musculoskeletal System | Injury |  |  | E20¦E1603¦E160302
|  | Multiple Fractures | E160303 | Two or more simultaneous traumatic breaks in one or more bones. | MedDRA:10028200:Multiple fractures | Musculoskeletal System | Injury |  |  | E20¦E1603¦E160303
|  | Skull Fracture | E160304 | A traumatic or pathologic injury to the bones of the skull in which the continuity of the bones of the skull is broken. | MedDRA:10061365:Skull fracture | Musculoskeletal System | Injury |  |  | E20¦E1603¦E160304
|  | Vertebral Fracture | E160305 | Traumatic or pathological injury to the spine in which the continuity of a vertebral bone is broken. | MedDRA:10047333:Vertebral fracture | Musculoskeletal System | Injury |  |  | E20¦E1603¦E160305
| Damage to Ligament(s) |  | E1606 | An injury to any of the fibrous bands of tissue that comprise an articulation (joint). | MedDRA:10061223:Ligament injury | Musculoskeletal System | Injury |  |  | E20¦E1606
| Head Injury |  | E1610 | A traumatic injury to the head. | MedDRA:10019196:Head injury | Musculoskeletal System | Injury |  |  | E20¦E1610
| Joint Dislocation |  | E1614 | An abnormal separation where two or more bones meet. | MedDRA:10023204:Joint dislocation | Musculoskeletal System | Injury |  |  | E20¦E1614
|  | Subluxation | E161401 | Partial dislocation of the bones in a joint. | MedDRA:10042402:Subluxation | Musculoskeletal System | Injury |  |  | E20¦E1614¦E161401
| Spinal Column Injury |  | E1631 | Damage to the bones of the spine. | MedDRA:10069689:Spinal column injury | Musculoskeletal System | Injury |  |  | E20¦E1631
| Tooth Fracture |  | E1633 | Breakage of any portion of the tooth. | MedDRA:10062544:Tooth fracture | Musculoskeletal System | Injury |  |  | E20¦E1633
| Abrasion |  | E1701 | Superficial damage to the skin caused by rubbing or scraping. | MedDRA:10080919:Abrasion | Skin and Subcutaneous Tissue | Injury |  |  | E20¦E1701
| Burn(s) |  | E1704 | An injury or tissue damage caused by exposure to caustic chemicals, extreme heat, extreme cold or excessive radiation. | MedDRA:10006634:Burn | Skin and Subcutaneous Tissue | Injury |  |  | E20¦E1704
|  | Bowel Burn | E1005 | A thermal injury to the bowels. | MedDRA:10082501:Intestinal burn | Gastrointestinal System | Injury |  |  | E20¦E1704¦E1005
|  | Gastroesophageal Burn | E1013 | Thermal, chemical or radiation injury to stomach or esophagus. | MedDRA:10082500:Gastroesophageal burn | Gastrointestinal System | Injury |  |  | E20¦E1704¦E1013
|  | Caustic/Chemical Burns | E170401 | A burn caused by exposure to chemical substances. | MedDRA:10008420:Chemical burn | Skin and Subcutaneous Tissue | Injury |  |  | E20¦E1704¦E170401
|  | Cryogenic Burn | E170402 | Local tissue damage due to exposure to extreme cold. | MedDRA:10081985:Cold burn | Skin and Subcutaneous Tissue | Injury |  |  | E20¦E1704¦E170402
|  | Radiation Burn | E170403 | A burn caused by exposure to x-ray, radium, sunlight, ionizing radiation or any other type of radiant energy. | MedDRA:10063640:Radiation burn | Skin and Subcutaneous Tissue | Injury |  |  | E20¦E1704¦E170403
|  | Superficial (First Degree) Burn | E170404 | A burn that affects the epidermis only, causing erythema without blistering. | MedDRA:10016700:First degree burns | Skin and Subcutaneous Tissue | Injury |  |  | E20¦E1704¦E170404
|  | Partial thickness (Second Degree) Burn | E170405 | A burn that affects the epidermis and the dermis. | MedDRA:10039798:Second degree burns | Skin and Subcutaneous Tissue | Injury |  | Editorial change to term for Release Number 2026. See no. 287 of Change Log. | E20¦E1704¦E170405
|  | Full thickness (Third Degree) Burn | E170406 | A burn to the skin that extends into deeper layers and underlying tissues. | MedDRA:10043441:Third degree burns | Skin and Subcutaneous Tissue | Injury |  | Editorial change to term for Release Number 2026. See no. 288 of Change Log. | E20¦E1704¦E170406
| Bone Shedding Debris |  | E2001 | Shedding of accumulated bone fragments. | MedDRA:10006011:Bone shedding debris |  |  |  |  | E20¦E2001
| Bruise/Contusion |  | E2002 | Injury of the soft tissues without breaking the skin that is characterized by blood leakage into the surrounding area. | MedDRA:10006502:Bruise |  |  |  |  | E20¦E2002
|  | Myocardial Contusion | E0615 | A bruise to the heart. | MedDRA:10068147:Myocardial contusion | Heart | Procedural Complications |  |  | E20¦E2002¦E0615
| Chemical Exposure |  | E2003 | Contact with a chemical substance through touch, inhalation, or ingestion. | MedDRA:10068154:Chemical exposure |  |  |  |  | E20¦E2003
| Crushing Injury |  | E2004 | Traumatic compression of the body or a part of the body. | MedDRA:10061097:Crush injury |  |  |  |  | E20¦E2004
| Easy Bruising |  | E2005 | Appearance of bruises spontaneously or following minimal contact. | MedDRA:10076258:Tendency to bruise easily | Injury | Skin and Subcutaneous Tissue |  |  | E20¦E2005
| Erosion |  | E2006 | An eating away or breakdown of any type of external or internal human tissue including but not limited to skin, teeth, mucosa, or somatic, which involves only the outer tissue layer. When tissue surrounds an implanted device, the tissue breakdown may result in migration and loss of the implant material and may result in further complications such as infection or abscess. | MedDRA:10068527:Tissue erosion associated with device |  |  |  |  | E20¦E2006
|  | Pocket Erosion | E1713 | Erosion of patient's tissue pocket which houses a device. | MedDRA:10035771:Pocket erosion | Skin and Subcutaneous Tissue | Injury |  |  | E20¦E2006¦E1713
|  | Skin Erosion | E1718 | A gradual breakdown or very shallow ulceration of the skin which involves only the epidermis and heals without scarring. | MedDRA:10040840:Skin erosion | Skin and Subcutaneous Tissue | Injury |  |  | E20¦E2006¦E1718
| Fall |  | E2007 | A sudden movement downward, usually resulting in injury. | MedDRA:10016173:Fall |  |  |  |  | E20¦E2007
| Foreign Body In Patient |  | E2008 | An occurrence involving any object including device and/or fragments of device being left unintentionally in the body. This includes devices left in the body as part of risk mitigation or surgical bailout procedures since their presence or location is not as originally intended. | MedDRA:10070245:Foreign body | Injury | Procedural Complications |  | Definition modified for Release Number 2025. See no. 25 of Change Request Log | E20¦E2008
|  | Device Embedded In Tissue or Plaque | E200801 | Device and/or fragments of device are embedded in patient's tissue and/or plaque. | MedDRA:10074498:Embedded device | Injury | Procedural Complications |  |  | E20¦E2008¦E200801
| Laceration(s) |  | E2009 | A cut or tear in any tissue. | MedDRA:10023572:Laceration |  |  |  |  | E20¦E2009
| Needle Stick/Puncture |  | E2010 | A penetrating stab wound from a needle (or other sharp object) that may result in exposure to blood or other body fluids. | MedDRA:10028895:Needle stick/puncture |  |  |  |  | E20¦E2010
| Physical Entrapment |  | E2011 | Patient becomes entangled or trapped in a device. | MedDRA:10082491:Patient entrapped in device |  |  |  |  | E20¦E2011
| Wound Infection |  | E2012 | Infection of a wound. This does not include post surgical wound infection. | MedDRA:10048038:Wound Infection | Injury | Infections |  | Term/Definition was modified on 31 January 2023 (Old term/definition: Post Traumatic Wound Infection: Infection of a wound following trauma. This does not include post surgical wound infection.) MedDRA mapping also updated from Post-traumatic wound infection (10036317) to Wound Infection (10048038). For details, see comment No. 95 of the Change Log (Release Number 2023). | E20¦E2012
| Rupture |  | E2013 | Forcible tearing or disruption of tissue. | MedDRA:10080972:Tissue rupture |  |  |  |  | E20¦E2013
| Tissue Breakdown |  | E2014 | A process caused by injury or disease or aging and characterized by a partial/complete disruption of tissue. The disruption may be temporary or permanent and may be associated with either (1) impairment or (2) loss of function of the affected organ(s) or system(s). | MedDRA:10068521:Tissue breakdown associated with device |  |  |  |  | E20¦E2014
|  | Pressure Sore/Ulcer | E201401 | Damage to tissue due as a result of external pressure. | MedDRA:10011985:Decubitus ulcer |  |  |  | Term and Definition and MedDRA map modified for Release Number 2025. Previous term was "Pressure Sores". Previous MedDRA map was 10036651:Pressure sore. See no. 70 of Change RequestLog. | E20¦E2014¦E201401
| Unspecified Tissue Injury |  | E2015 | The report describes a non-specific injury to the skin or subcutaneous tissue. Note: Please use E2402 if the report describes a specific injury but the relevant term does not exist. | MedDRA:10080903:Tissue injury |  |  | Modified (editorial) | Editorial change to the definition on 18 March 2020 for consistency with other "Unspecified ____" definitions. Editorial change to definition for Release Number 2026. See no. 211 of Change Log | E20¦E2015
| Perineal Injury |  | E2016 | Traumatic damage and/or injury to the perineum. | MedDRA:10073329:Perineal injury |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 17 of Change Log. Editorial change to term for Release Number 2026. See no. 292 of Change Log. | E20¦E2016
| Electric Shock |  | E2104 | A shock caused by a brief, unintentional exposure to an electric current from a device. | MedDRA:10014357:Electric shock | Procedural Complications | Injury |  |  | E20¦E2104
|  | Shock from Patient Lead(s) | E210401 | Unintended electric shock from patient leads. This includes an electric shock to a healthcare professional or a patient receiving a shock to an incorrect area of their body. | MedDRA:10075322:Device inappropriate shock delivery | Procedural Complications | Injury |  |  | E20¦E2104¦E210401
| Exposure to Body Fluids |  | E2105 | Exposed to or contacted with blood or body fluids of another person. | MedDRA:10068528:Exposure to device contaminated with body fluid | Procedural Complications | Injury |  |  | E20¦E2105
| Radiation Sickness Syndrome |  | E2116 | The complex of symptoms characterizing the disease known as radiation injury, resulting from excessive exposure of the whole body (or large part) to ionizing radiation. | MedDRA:10068142:Radiation sickness syndrome | Procedural Complications | Injury |  |  | E20¦E2116
| Airway Burns |  | E0762 | Burn to upper respiratory tract from inhalation of heated gases. | MedDRA:10067204:Airway Burns | Respiratory System | Injury | New | Term added for Release Number 2026. See no. 68 of Change Request Log | E20¦E0762
Procedural Complications |  |  | E21 | Procedural Complications |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E21
| Dural Tear |  | E0144 | Injury to the dura mater, with or without damage to any other layer of the meninges. | MedDRA:10063395:Dural tear | Nervous System | Procedural Complications |  | Term added for Release Number 2024. See no. 22 of Change Log | E21¦E0144
| Vitreous Detachment |  | E081906 | Detachment of the vitreous humor from the retina. | MedDRA:10047650:Vitreous detachment | Eye | Procedural Complications |  |  | E21¦E081906
| Rectal Anastomotic Leakage |  | E1026 | Leakage due to breakdown of a rectal anastomosis. | MedDRA:10065894:Rectal anastomotic leak | Gastrointestinal System | Procedural Complications | Retired (2024) | Term retired for Release Number 2024. Use E2108 - Failure to Anastomose. See no. 81 of Change Log | E21¦E1026
| Bone Notching |  | E1646 | An indentation made at the edge of the bone without bone breakage. | MedDRA:10057328:Bone injury | Musculoskeletal System | Procedural Complications |  | Term added for Release Number 2024. See no. 76 of Change Log | E21¦E1646
| Foreign Body In Patient |  | E2008 | An occurrence involving any object including device and/or fragments of device being left unintentionally in the body. This includes devices left in the body as part of risk mitigation or surgical bailout procedures since their presence or location is not as originally intended. | MedDRA:10070245:Foreign body | Injury | Procedural Complications |  | Definition modified for Release Number 2025. See no. 25 of Change Request Log | E21¦E2008
|  | Device Embedded In Tissue or Plaque | E200801 | Device and/or fragments of device are embedded in patient's tissue and/or plaque. | MedDRA:10074498:Embedded device | Injury | Procedural Complications |  |  | E21¦E2008¦E200801
| Adhesion(s) |  | E2101 | A fibrous band of tissue that connects normally separate body regions. | MedDRA:10059837:Adhesion | Procedural Complications | Generalized Disorders |  |  | E21¦E2101
| Awareness During Anesthesia |  | E2102 | A state of consciousness that permits the formation of explicit and implicit memories while under general anaesthesia. | MedDRA:10054330:Awareness during anesthesia |  |  | Modified (editorial) | Editorial change to term and definition for Release Number 2026. See no. 293 of Change Log. | E21¦E2102
| Device Overstimulation of Tissue |  | E2103 | A level of stimulation that is greater than the intended level. | MedDRA:10068141:Device overstimulation of tissue |  |  |  | Editorial change on 27 January 2022 to the MedDRA mapping to correct the MedDRA code. | E21¦E2103
| Electric Shock |  | E2104 | A shock caused by a brief, unintentional exposure to an electric current from a device. | MedDRA:10014357:Electric shock | Procedural Complications | Injury |  |  | E21¦E2104
|  | Shock from Patient Lead(s) | E210401 | Unintended electric shock from patient leads. This includes an electric shock to a healthcare professional or a patient receiving a shock to an incorrect area of their body. | MedDRA:10075322:Device inappropriate shock delivery | Procedural Complications | Injury |  |  | E21¦E2104¦E210401
| Exposure to Body Fluids |  | E2105 | Exposed to or contacted with blood or body fluids of another person. | MedDRA:10068528:Exposure to device contaminated with body fluid | Procedural Complications | Injury |  |  | E21¦E2105
| External Prosthetic Device Pain |  | E2106 | Pain that is associated with an external artificial replacement device. | MedDRA:10059057:Medical device pain | Procedural Complications | Generalized Disorders |  | Definition was modified on 21 January 2021 to correct a typo in "artificial" (Old definition: Pain that is associated with an external ariticial replacement device.) | E21¦E2106
| Failure of Implant |  | E2107 | Malfunction of a medical implant. | MedDRA:10016153:Failure of implant |  |  | Retired (2025) | Term retired for Release Number 2025. When there are no overt clinical signs and symptoms use E2403 - No Clinical Signs, Symptoms or Conditions, otherwise the appropriate E codes should be selected. See no. 61 of Change Request Log | E21¦E2107
| Failure to Anastomose |  | E2108 | Failure of a surgically-induced connection between tubular structures in the body. Includes Anastomotic Leakage | MedDRA:10016164:Failure to anastomose |  |  |  | Definition modified for Release Number 2024. See no. 81 of Change Log | E21¦E2108
| Implant Pain |  | E2109 | Pain localized to the site of the implanted device. | MedDRA:10063782:Implant site pain | Procedural Complications | Generalized Disorders |  |  | E21¦E2109
| Inadequate Pain Relief |  | E2110 | Continuing discomfort due to a lack or failure of a medical approach to treat pain. | MedDRA:10079788:Inadequate pain relief | Procedural Complications | Generalized Disorders |  |  | E21¦E2110
| Administration Site Reaction |  | E2111 | An intense reaction (usually immunologic) developing at the site of infusion or injection | MedDRA:10001315:Administration site reaction |  |  | Modified (technical) | Term and definition modified for Release Number 2026. See no. 101 of Change Log. | E21¦E2111
| Intraoperative Pain |  | E2112 | Pain occurring during a surgical operation. | MedDRA:10064882:Procedural pain | Procedural Complications | Generalized Disorders |  |  | E21¦E2112
| Oversedation |  | E2113 | A level of sedation that is greater than the intended level. | MedDRA:10033305:Oversedation |  |  |  |  | E21¦E2113
| Perforation |  | E2114 | A hole or opening made through a membrane or other tissue or material. | MedDRA:10076705:Perforation |  |  |  |  | E21¦E2114
|  | Perforation of Vessels | E0511 | Damage to a blood vessel resulting from a full thickness disruption in the integrity of the vessel wall. | MedDRA:10068149:Vessel perforation | Vascular System | Procedural Complications |  |  | E21¦E2114¦E0511
|  | Great Vessel Perforation | E051101 | Damage to the venae cavae, pulmonary artery, pulmonary veins or aorta resulting from a full thickness disruption in the integrity of the vessel wall. | MedDRA:10080896:Perforation of great vessels | Vascular System | Procedural Complications |  |  | E21¦E2114¦E051101
|  | Cardiac Perforation | E0604 | The presence of an acquired hole in the heart. | MedDRA:10058039:Cardiac perforation | Heart | Procedural Complications |  |  | E21¦E2114¦E0604
|  | Mitral Perforation | E0614 | A rupture in the mitral valve tissue due to traumatic or pathologic processes. | MedDRA:10068138:Mitral perforation | Heart | Procedural Complications |  |  | E21¦E2114¦E0614
|  | Sinus Perforation | E0745 | A rupture in the sinus tissue due to traumatic or pathologic processes. | MedDRA:10040748:Sinus perforation | Respiratory System | Procedural Complications |  |  | E21¦E2114¦E0745
|  | Corneal Perforation | E0811 | A rupture in the corneal tissue due to traumatic or pathologic processes. | MedDRA:10011039:Corneal perforation | Eye | Procedural Complications |  |  | E21¦E2114¦E0811
|  | Bowel Perforation | E1006 | A rupture in the wall of the small or large intestine due to traumatic or pathologic processes. | MedDRA:10006056:Bowel perforation | Gastrointestinal System | Procedural Complications |  |  | E21¦E2114¦E1006
|  | Perforation of Esophagus | E1022 | The presence of a hole or other type of opening in the esophageal wall through which the contents of the esophagus can pass into the mediastinum. | MedDRA:10034419:Perforation of esophagus | Gastrointestinal System | Procedural Complications |  |  | E21¦E2114¦E1022
|  | Fallopian Tube Perforation | E1406 | A rupture in the fallopian tube due to traumatic or pathologic processes. | MedDRA:10065790:Fallopian tube perforation | Reproductive System and Breast | Procedural Complications |  |  | E21¦E2114¦E1406
|  | Uterine Perforation | E1415 | A rupture in the uterus due to traumatic or pathologic processes. | MedDRA:10046810:Uterine perforation | Reproductive System and Breast | Procedural Complications |  |  | E21¦E2114¦E1415
|  | Internal Organ Perforation | E211401 | A rupture in the organ(s) due to traumatic or pathologic processes. | MedDRA:10080922:Perforation of organ |  |  |  |  | E21¦E2114¦E211401
| Post Operative Wound Infection |  | E2115 | Infection of a surgical skin incision. | MedDRA:10036410:Postoperative wound infection | Procedural Complications | Infections |  |  | E21¦E2115
| Radiation Sickness Syndrome |  | E2116 | The complex of symptoms characterizing the disease known as radiation injury, resulting from excessive exposure of the whole body (or large part) to ionizing radiation. | MedDRA:10068142:Radiation sickness syndrome | Procedural Complications | Injury |  |  | E21¦E2116
| Twiddlers Syndrome |  | E2117 | Dislodgement, breakdown, or other malfunction of an artificial cardiac pacemaker, chemotherapy port, drip infusion valve, or similar implanted diagnostic or therapeutic device as a result of unconscious or habitual manipulation by the patient. | MedDRA:10045184:Twiddler's syndrome |  |  |  |  | E21¦E2117
| Unintended Extubation |  | E2118 | The removal of a previously inserted tube from a hollow organ or passageway, often from the airway. | MedDRA:10066823:Unintended endotracheal extubation |  |  |  |  | E21¦E2118
| Unintended Radiation Exposure |  | E2119 | The unintentional exposure to radiation of a person, or a part of a persons body. This might include exposure of a healthcare professional, or the exposure of the wrong part of the body of a patient who was intended to receive therapeutic or diagnostic radiation. This includes both ionizing and non ionizing radiation. | MedDRA:10063155:Radiation exposure |  |  |  | Editorial change to the definition on 16 February 2021. | E21¦E2119
| Bone Cement Implantation Syndrome |  | E2120 | A rare and potentially fatal perioperative complication of cemented bone surgery, characterised by hypotension, hypoxia, cardiac dysrhythmias, and in severe cases cardiac arrest. | MedDRA:10085596:Bone cement implantation syndrome |  |  |  | Term was added on 21 January 2021. For details, refer to comment No. 33 of the Change Log. | E21¦E2120
| Clinical Endoleak |  | E2121 | The observed patient condition is persistent blood flow/pressure in the aneurysm sac following an endovascular aneurysm repair (EVAR) procedure. Where other patient signs and symptoms indicate the presence of an endoleak those signs and symptoms should be coded. | MedDRA:10063824:Endoleak | Procedural Complications | Vascular System |  | Term was added on 27 January 2022. For details, see comment No. 181 / 207 of the Change Log (Release Number 2022). Term and Definition modified for Release Number 2025. Previous term was "Endoleaks". See no. 117 of Change Request Log. | E21¦E2121
| Sinusitis |  | E2122 | Inflammation (swelling and redness) of the paranasal sinuses, associated or related to implantation procedures or devices. | MedDRA:10040753:Sinusitis | Procedural Complications | Infections |  | Term was added on 27 January 2022. For details, see comment No. 240 of the Change Log (Release Number 2022). | E21¦E2122
| Medical Device Site Infection |  | E2123 | Infection local to an implanted or invasive medical device. | MedDRA:10075580:Medical device site infection | Procedural Complications | Infections | Modified (editorial) | Term was added on 27 January 2022. For details, see comment No. 122 of the Change Log (Release Number 2022). Term and Definition and MedDRA map modified for Release Number 2024. Previous term was Peri-prosthetic Joint Infection. Previous MedDRA map was 10076118:Medical device site joint infection. See no. 10 of Change Log. Editorial change to term for Release Number 2026. See no. 294 of Change Log. | E21¦E2123
| Unintended Loss of Vascular Access |  | E2124 | The accidental removal of a previously inserted line from a blood vessel. | MedDRA:10070944:Vascular access failure |  |  |  | Term was added on 31 January 2023. For details, see comment No. 97 of the Change Log (Release Number 2023). | E21¦E2124
| Complex Regional Pain Syndrome |  | E2125 | Condition where a person experiences persistent severe and debilitating pain, involving an extremity or other body region, hyperesthesia, and localized autonomic dysfunction following injury to soft tissue or nerve | MedDRA:10064332:Complex regional pain syndrome | Procedural Complications | Nervous System | Modified (editorial) | Term added for Release Number 2024. See no. 56 of Change Log. Editorial change to term for Release Number 2026. See no. 295 of Change Log. | E21¦E2125
| Loss of Fracture Reduction |  | E2126 | Loss of anatomical alignment after fracture reduction. | MedDRA:10069723:Loss of anatomical alignment after fracture reduction | Procedural Complications | Musculoskeletal System | Modified (editorial) | Term added for Release Number 2024. See no. 65 of Change Log. Editorial change to term for Release Number 2026. See no. 296 of Change Log. | E21¦E2126
| Periprosthetic Fracture |  | E2127 | Fractures around joint replacement prosthetics or implants. These can occur intraoperatively or postoperatively. | MedDRA:10069135:Periprosthetic fracture | Procedural Complications | Musculoskeletal System | Modified (editorial) | Term added for Release Number 2024. See no. 67 of Change Log. Editorial change to term for Release Number 2026. See no. 297 of Change Log. | E21¦E2127
| Implant Site Sclerosis |  | E2128 | Pathological process consisting of hardening/thickening of tissue or structures. | MedDRA:10063876:Implant site sclerosis |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 94 of Change Log. Editorial change to term for Release Number 2026. See no. 298 of Change Log. | E21¦E2128
| Peritoneal Effluent Abnormal |  | E2129 | A change in the appearance or other characteristics of peritoneal dialysis effluent (e.g. cloudy or discolored). | MedDRA:10069638:Peritoneal effluent abnormal |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 99 of Change Log. Editorial change to term for Release Number 2026. See no. 299 of Change Log. | E21¦E2129
| Graft Failure |  | E2130 | Failure of a transplanted graft. | MedDRA:10069030:Graft failure |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 104 of Change Log. Editorial change to term for Release Number 2026. See no. 300 of Change Log. | E21¦E2130
| Inappropriate Device Stimulation of Tissue |  | E2131 | Unintended reaction in a tissue (e.g. nerve, muscle) caused by electrical stimulation | MedDRA:10064414:Inappropriate device stimulation of tissue |  |  |  | Term added for Release Number 2025. See no. 17 of Change Request Log | E21¦E2131
| Medical Device Site Inflammation |  | E2132 | Inflammation local to an implanted or invasive medical device. | MedDRA:10075581:Medical device site Inflammation |  |  |  | Term added for Release Number 2025. See no. 68 of Change Request Log | E21¦E2132
| Local Anesthetic Systemic Toxicity (LAST) |  | E2133 | A systemic adverse reaction resulting from elevated plasma concentrations of local anesthetics | MedDRA:10082539:Local anesthetic systemic toxicity |  |  | New | Term added for Release Number 2026. See no. 23 of Change Log. | E21¦E2133
| Administration Site Pain |  | E2134 | Pain at the infusion or injection site. | MedDRA:10058049:Administration site pain |  |  | New | Term added for Release Number 2026. See no. 106 of Change Log. | E21¦E2134
| Ocular Procedural Complication |  | E2135 | Any complication arising from ocular surgical procedures involving the conjunctiva or adjacent tissues, including but not limited to subconjunctival bleb formation (e.g., trabeculectomy, glaucoma filtration procedures). This may encompass issues such as bleb absence, infection, leakage, perforation, structural thinning, scarring, or other adverse changes affecting the surgical site or surrounding ocular structures. | MedDRA:10080964:Ocular procedural complication | Procedural Complications | Eye | New | Term added for Release Number 2026. See no. 381 of Change Log. | E21¦E2135
| Gingival Recession |  | E2136 | Retraction or loss of gingival tissue resulting in exposure of the root surface | MedDRA:10018290:Gingival recession |  |  | New | Term added for Release Number 2026. See no. 24 of Change Log. | E21¦E2136
| Physical Asymmetry |  | E2332 | Lack or absence of balanced proportions between parts or organs of the body. |  | Generalized Disorders | Procedural Complications |  |  | E21¦E2332
| Wound Dehiscence |  | E2340 | A complication in which the wound splits or separates after being surgically closed. | MedDRA:10048031:Wound dehiscence | Generalized Disorders | Procedural Complications |  |  | E21¦E2340
| Pigment Dispersion Syndrome |  | E0847 | Condition where iris pigment is released into the anterior chamber and deposited in the trabecular meshwork, due to a procedure or device implantation. | MedDRA:10052134:Pigment dispersion syndrome | Eye | Procedural Complications |  | Term was added on 27 January 2022. For details, see comment No. 127 of the Change Log (Release Number 2022). | E21¦E0847
| Corneal Endothelial Cell Loss |  | E0848 | Decrease in corneal endothelial cell density. | MedDRA:10071681:Corneal endothelial cell loss | Eye | Procedural Complications |  | Term was added on 27 January 2022. For details, see comment No. 129 of the Change Log (Release Number 2022). | E21¦E0848
| Posterior Capsule Opacification |  | E0849 | Remnant lens epithelial cell proliferation and migration to the posterior capsule. | MedDRA:10036346:Posterior capsule opacification | Eye | Procedural Complications |  | Term was added on 27 January 2022. For details, see comment No. 131 of the Change Log (Release Number 2022). | E21¦E0849
| Biliary Leak |  | E1108 | Sudden unexpected leakage of bile from the liver or soft tissue. | MedDRA:10057586:Post procedural bile leak | Hepatic and Biliary System | Procedural Complications |  | Term was added on 27 January 2022. For details, see comment No. 5 of the Change Log (Release Number 2022). | E21¦E1108
| Cholangitis |  | E1109 | Inflammation of the bile duct. | MedDRA:10008604:Cholangitis | Hepatic and Biliary System | Procedural Complications |  | Term was added on 27 January 2022. For details, see comment No. 6 of the Change Log (Release Number 2022). | E21¦E1109
| Intraoperative Cardiac Valve Injury |  | E0624 | An intraoperative event that results in damage to, or abnormal function of, a heart valve. E.g. a tear in the leaflet or the chordae tendinae. | MedDRA:10088741:Intraoperative cardiac valve injury | Heart | Procedural Complications |  | Term was added on 31 January 2023. For details, see comment No. 80 / 114 of the Change Log (Release Number 2023). Definition modified for Release Number 2025. See no. 62 of Change Request Log | E21¦E0624
Investigations and Diagnostic Tests |  |  | E22 | Investigations and Diagnostic Tests |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E22
| High Pulmonary Arterial Wedge Pressure |  | E0507 | The measurement of the mean left atrial pressure, as measured by a catheter introduced into the distal pulmonary artery, is high. | MedDRA:10037328:Pulmonary arterial wedge pressure high | Vascular System | Investigations and Diagnostic Tests |  |  | E22¦E0507
| Low Pulmonary Arterial Wedge Pressure |  | E0508 | The measurement of the mean left atrial pressure, as measured by a catheter introduced into the distal pulmonary artery, is low. | MedDRA:10037330:Pulmonary arterial wedge pressure low | Vascular System | Investigations and Diagnostic Tests |  | Definition was modified on 21 January 2021 (Old definition: The measurement of the mean left pulmonary atrial pressure, as measured by a catheter introduced into the distal pulmonary artery, is low.). For details, see comment No. 3 of the Change Log. | E22¦E0508
| Cardiac Enzyme Elevation |  | E0603 | Increase from the normal levels of cardiac enzymes. | MedDRA:10007548:Cardiac enzymes increased | Heart | Investigations and Diagnostic Tests |  |  | E22¦E0603
| Decreased Peak Expiratory Flow rate |  | E0713 | Decrease in the maximum flow rate generated during a forceful exhalation, starting from full lung inflation. | MedDRA:10034195:Peak expiratory flow rate decreased | Respiratory System | Investigations and Diagnostic Tests |  | Editorial change to term for Release Number 2026. See no. 262 of Change Log. | E22¦E0713
| Increased Peak Expiratory Flow rate |  | E0714 | An increase of the greatest rate of flow that can be achieved during forced expiration beginning with the lungs fully inflated. | MedDRA:10034196:Peak expiratory flow rate increased | Respiratory System | Investigations and Diagnostic Tests |  | Editorial change to term for Release Number 2026. See no. 263 of Change Log. | E22¦E0714
| Decreased Respiratory Rate |  | E0715 | Abnormal decrease of rate of breathing. | MedDRA:10038710:Respiratory rate decreased | Respiratory System | Investigations and Diagnostic Tests |  |  | E22¦E0715
| Increased Respiratory Rate |  | E0716 | Abnormal increase of rate of breathing. | MedDRA:10038712:Respiratory rate increased | Respiratory System | Investigations and Diagnostic Tests |  |  | E22¦E0716
| Forced Expiratory Volume Decreased |  | E0719 | A decrease in the fraction of the forced vital capacity that is exhaled in a specific number of seconds. | MedDRA:10016987:Forced expiratory volume decreased | Respiratory System | Investigations and Diagnostic Tests |  |  | E22¦E0719
| Forced Expiratory Volume Increased |  | E0720 | An increase in the fraction of the forced vital capacity that is exhaled in a specific number of seconds. | MedDRA:10016989:Forced expiratory volume increased | Respiratory System | Investigations and Diagnostic Tests |  |  | E22¦E0720
| Hyperglycemia |  | E1205 | Abnormally high level of glucose in the blood. | MedDRA:10020639:Hyperglycemia | Endocrine, Metabolism and Nutrition | Investigations and Diagnostic Tests |  |  | E22¦E1205
|  | Elevated Ketones/ Diabetic Ketoacidosis | E120501 | Elevated ketones, including metabolic acidosis produced by accumulation of ketone bodies resulting from uncontrolled diabetes mellitus. | MedDRA:10057594:Blood ketone body increased |  |  | Modified (editorial) | Editorial change on 31 January 2023 to align E120501 term. (Old term: Diabetic Ketoacidosis). MedDRA mapping is updated from Diabetic ketoacidosis (10012671) to Blood ketone body increased (10057594). For details, see comment No. 85 / 86 of the Change Log (Release Number 2023). Minor editorial change to term for Release Number 2026. See no. 272 of Change Log. | E22¦E1205¦E120501
| Hypoglycemia |  | E1206 | Abnormally low level of glucose in the blood. | MedDRA:10021005:Hypoglycemia | Endocrine, Metabolism and Nutrition | Investigations and Diagnostic Tests |  |  | E22¦E1206
| Hormonal Imbalance |  | E1209 | An imbalance of one or more hormones in the body | MedDRA:10020381:Hormonal imbalance | Endocrine, Metabolism and Nutrition | Investigations and Diagnostic Tests |  | Editorial change on 27 January 2022 to the MedDRA mapping to include the MedDRA term. | E22¦E1209
| Cervical Changes |  | E1404 | Changes such as abnormal, pathological, benign, malignant, etc. | MedDRA:10008351:Cervix disorder | Reproductive System and Breast | Investigations and Diagnostic Tests |  |  | E22¦E1404
| Low Apgar Score |  | E1510 | Any Apgar score less than seven. This may be the initial or a subsequent apgar score. | MedDRA:10002944:Apgar score low | Pregnancy, Childbirth and the Puerperium | Investigations and Diagnostic Tests |  |  | E22¦E1510
| Abnormal Blood Gases |  | E2201 | Abnormal Blood Gas Measurements. | MedDRA:10005539:Blood gases abnormal |  |  |  |  | E22¦E2201
| High Oxygen Saturation |  | E2202 | A high level of the degree to which oxygen is bound to hemoglobin given as a percentage calculated by dividing the maximum oxygen capacity into the actual oxygen content and multiplying by 100. Oxygen saturation usually is measured using pulse oximetry. | MedDRA:10033319:Oxygen saturation high | Investigations and Diagnostic Tests | Respiratory System |  |  | E22¦E2202
| Low Oxygen Saturation |  | E2203 | A low level of the degree to which oxygen is bound to hemoglobin given as a percentage calculated by dividing the maximum oxygen capacity into the actual oxygen content and multiplying by 100. Oxygen saturation usually is measured using pulse oximetry. | MedDRA:10033321:Oxygen saturation low | Investigations and Diagnostic Tests | Respiratory System |  |  | E22¦E2203
| Lactate Dehydrogenase Increased |  | E2204 | A laboratory test result indicating an increase in the lactate dehydrogenase concentration of a biological sample. | MedDRA:10023654:Lactate dehydrogenase increased |  |  |  |  | E22¦E2204
| Bacteremia |  | E2205 | Condition that relates to the presence of viable bacteria in the bloodstream, whether associated with active disease or not. | MedDRA:10003999:Bacteremia | Investigations and Diagnostic Tests | Infections |  | Term was added on 27 January 2022. For details, see comment No. 49 of the Change Log (Release Number 2022). | E22¦E2205
| High Metal Ion Levels |  | E2206 | High blood or serum levels of ions which are attributed to one or more medical devices used, such as but not limited to Co, Cr, Ti, Ni, Mo. | MedDRA:10086050:Heavy metal increased |  |  | Modified (editorial) | Term was added on 27 January 2022. For details, see comment No. 114 / 154 / 204 of the Change Log (Release Number 2022). Editorial change to term for Release Number 2026. See No. 301 of Change Log. | E22¦E2206
| Radiolucent Lines |  | E2207 | Radiolucencies observed on x-ray. |  |  |  |  | Term was added on 27 January 2022. For details, see comment No. 123 of the Change Log (Release Number 2022). | E22¦E2207
Generalized Disorders |  |  | E23 | Generalized Disorders |  |  |  | Not selectable | Please use more detailed term within the hierarchy | E23
| Anaphylactoid / Anaphylaxis |  | E040202 | Acute immune response that occurs from exposure to an allergen. For conditions which have progressed to systemic shock, use E040201 - Anaphylactic Shock. | MedDRA:10002218:Anaphylaxis | Immune System | Generalized Disorders | Modified (editorial) | Definition was modified on 31 January 2023 (Old definition: Resembling anaphylaxis.). MedDRA mapping also updated from Anaphylactoid reaction (10002216) to Anaphylaxis (10002218). For details, see comment No. 85 of the Change Log (Release Number 2023). | E23¦E040202
| Phosphene Visualization |  | E083802 | An objective visual sensation that appears with the eyes closed and in the absence of visual light. | MedDRA:10034936:Phosphenes | Eye | Generalized Disorders |  |  | E23¦E083802
| Burning Sensation |  | E1705 | A sensation of stinging or heat, not necessarily accompanied by redness or physical signs of irritation. | MedDRA:10006784:Burning sensation | Skin and Subcutaneous Tissue | Generalized Disorders |  |  | E23¦E1705
|  | Skin Burning Sensation | E170501 | A sensation of dermal stinging or heat, not necessarily accompanied by redness or physical signs of skin irritation. | MedDRA:10054786:Skin burning sensation | Skin and Subcutaneous Tissue | Generalized Disorders |  |  | E23¦E1705¦E170501
| Impaired Healing |  | E1707 | Improper repair of any body tissue due to factors affecting one or more of the phases of wound healing. Includes impaired healing of any body tissue. | MedDRA:10021519:Impaired healing | Generalized Disorders |  |  | Term moved and Definition modified for Release Number 2024. See no. 84 of Change Log. Please be aware that the code does not represent the current hierarchy. | E23¦E1707
| Alteration in Body Temperature |  | E2301 | A Change or alteration of Body Temperature to a temperature outside the range of normal body temperature. | MedDRA:10063488:Body temperature fluctuation |  |  |  |  | E23¦E2301
|  | Fever | E230101 | Elevation of body temperature above normal due to the bodies own responses. | MedDRA:10016558:Fever |  |  |  |  | E23¦E2301¦E230101
|  | Hyperthermia | E230102 | An abnormally high body temperature. Not fever. | MedDRA:10020843:Hyperthermia |  |  |  |  | E23¦E2301¦E230102
|  | Hypothermia | E230103 | Abnormally low body temperature. | MedDRA:10021113:Hypothermia |  |  |  |  | E23¦E2301¦E230103
| Ambulation or Postural Difficulties |  | E2302 | Any persistent problem with walking, gait or maintaining posture (sitting, standing). | MedDRA:10047810:Walking difficulty | Generalized Disorders | Musculoskeletal System |  | Term/Definition was modified on 31 January 2023 (Old term/definition: Ambulation Difficulties: Any persistent problem with walking or gait.) For details, see comment No. 98 of the Change Log (Release Number 2023). | E23¦E2302
| Capsular Contracture |  | E2303 | The tightening of scar tissue that forms around the implant. | MedDRA:10080895:Capsular contracture associated with implant |  |  |  |  | E23¦E2303
| Chills |  | E2304 | Shivering or moderate tremors of the body often accompanied by a cold sensation. | MedDRA:10008531:Chills |  |  |  |  | E23¦E2304
| Cyanosis |  | E2305 | A bluish or purplish discoloration of the skin and mucous membranes resulting from a reduced amount of oxygenated hemoglobin in the blood. | MedDRA:10011703:Cyanosis |  |  |  |  | E23¦E2305
| Decreased Appetite |  | E2306 | Lack of desire for food. | MedDRA:10061428:Decreased appetite | Generalized Disorders | Endocrine, Metabolism and Nutrition |  |  | E23¦E2306
| Increased Appetite |  | E2307 | Intense desire for food. | MedDRA:10021654:Increased appetite | Generalized Disorders | Endocrine, Metabolism and Nutrition |  |  | E23¦E2307
| Deformity/ Disfigurement |  | E2308 | An alteration of appearance, shape, size or alignment. | MedDRA:10061619:Deformity |  |  |  |  | E23¦E2308
| Deposits |  | E2309 | Accumulation of extraneous inorganic matter in tissues or cavities. | MedDRA:10070692:Device deposit issue |  |  |  |  | E23¦E2309
|  | Calcium Deposits/Calcification | E230901 | Deposition of calcium salt in the tissues. | MedDRA:10006958:Calcium deposits |  |  |  |  | E23¦E2309¦E230901
| Diaphoresis |  | E2310 | Profuse sweating. | MedDRA:10012703:Diaphoresis |  |  |  |  | E23¦E2310
| Discomfort |  | E2311 | A feeling of mental or physical uneasiness, pain, or distress. | MedDRA:10013082:Discomfort |  |  |  |  | E23¦E2311
| Fatigue |  | E2312 | A state of generalized weakness with a pronounced inability to summon sufficient energy to accomplish daily activities. | MedDRA:10016256:Fatigue |  |  |  |  | E23¦E2312
|  | Extreme Exhaustion | E231201 | Extreme fatigue; the inability to respond to stimuli. | MedDRA:10015667:Exhaustion |  |  |  | Editorial change to the definition on 16 February 2021. | E23¦E2312¦E231201
| Fibrosis |  | E2313 | The formation of fibrous tissue. | MedDRA:10016642:Fibrosis |  |  |  |  | E23¦E2313
|  | Adhesion(s) | E2101 | A fibrous band of tissue that connects normally separate body regions. | MedDRA:10059837:Adhesion | Procedural Complications | Generalized Disorders |  |  | E23¦E2313¦E2101
| Fistula |  | E2314 | Abnormal epithelial-lined communication between two anatomical structures. | MedDRA:10016717:Fistula | Generalized Disorders | Musculoskeletal System |  |  | E23¦E2314
| Fluid Discharge |  | E2315 | Flow or leakage of fluid other than blood, from an orifice, wound or tissue. | MedDRA:10080911:Fluid discharge |  |  |  |  | E23¦E2315
|  | Purulent Discharge | E231501 | A flow or draining of pus or exudate from an orifice or wound. | MedDRA:10037569:Purulent discharge |  |  |  | Editorial change to the definition on 16 February 2021. | E23¦E2315¦E231501
| Foreign Body Reaction |  | E2316 | An inflammatory reaction evoked by the presence of an exogenous material in the tissues. | MedDRA:10068159:Foreign body reaction |  |  |  |  | E23¦E2316
| Granuloma |  | E2317 | An inflammatory reaction usually caused by infectious organisms, foreign bodies, or cholesterol deposits. It is characterized by the presence of epithelioid histiocytes and chronic inflammation. | MedDRA:10018691:Granuloma | Generalized Disorders | Skin and Subcutaneous Tissue |  |  | E23¦E2317
| Hair Loss |  | E2318 | Loss of hair usually from the scalp that may result in bald spots or spread to the entire scalp or epidermis. | MedDRA:10019045:Hair loss |  |  |  |  | E23¦E2318
| Hernia |  | E2319 | The protrusion of part of an organ or fibroadipose tissue through an opening or weakened area. | MedDRA:10019909:Hernia |  |  |  |  | E23¦E2319
| High Blood Pressure/ Hypertension |  | E2320 | Blood pressure that is abnormally high. | MedDRA:10020772:Hypertension |  |  |  |  | E23¦E2320
| Low Blood Pressure/ Hypotension |  | E2321 | Blood pressure that is abnormally low. | MedDRA:10021097:Hypotension |  |  |  |  | E23¦E2321
| Hyperemia |  | E2322 | The presence of an increased amount of blood in a part or organ; engorgement or bloodshot | MedDRA:10020619:Hyperemia |  |  |  | Definition was modified on 31 January 2023 to add "bloodshot". For details, see comment No. 132 of the Change Log (Release Number 2023). | E23¦E2322
| Hyperplasia |  | E2323 | An abnormal increase in the number of cells in an organ or a tissue with consequent enlargement. | MedDRA:10020718:Hyperplasia |  |  |  |  | E23¦E2323
| Incontinence |  | E2324 | Involuntary passage of feces and/or urine from the body. | MedDRA:10021639:Incontinence |  |  |  |  | E23¦E2324
|  | Fecal Incontinence | E232401 | Inability to control voluntary passage of feces from the body. | MedDRA:10016296:Fecal incontinence | Generalized Disorders | Gastrointestinal System |  | Editorial change to add Gastrointestinal System as Secondary Category. For details, see comment No. 99 of the Change Log (Release Number 2023). Code hierarchy adjusted for secondary category to prevent primary parent term appearing in three places. Release No. 2024 | E23¦E2324¦E232401
|  | Urinary Incontinence | E232402 | Inability to control voluntary passage of urine from the body. | MedDRA:10046543:Urinary incontinence | Generalized Disorders | Kidney and Urinary Tract |  | Editorial change to add Kidney and Urinary Tract as Secondary Category. For details, see comment No. 100 of the Change Log (Release Number 2023). Amended in Version 1 on 23 February 2023. Code hierarchy adjusted for secondary category to prevent primary parent term appearing in three places. Release No. 2024 | E23¦E2324¦E232402
| Infiltration into Tissue |  | E2325 | Permeation of foreign substances into tissues or cells. | MedDRA:10080901:Tissue infiltration |  |  |  |  | E23¦E2325
| Inflammation |  | E2326 | A localized physical condition in which part of the body becomes reddened, swollen, or hot and often painful. | MedDRA:10061218:Inflammation |  |  |  |  | E23¦E2326
| Necrosis |  | E2327 | The pathologic death of cells or of a portion of tissue or organ resulting from irreversible damage. | MedDRA:10028851:Necrosis |  |  |  |  | E23¦E2327
| Obstruction/Occlusion |  | E2328 | Blockage or closing of the normal flow of the contents of an anatomical passageway. | MedDRA:10061876:Obstruction |  |  |  |  | E23¦E2328
| Organ Dehiscence |  | E2329 | A splitting open or rupture of an internal organ with the exposure or discharge of its content. | MedDRA:10080904:Dehiscence |  |  |  |  | E23¦E2329
| Pain |  | E2330 | An unpleasant sensory and emotional experience that is associated with actual or potential tissue damage or described in such terms. | MedDRA:10033371:Pain |  |  |  |  | E23¦E2330
|  | Headache | E0116 | Pain in various parts of the head, not confined to the area of distribution of any nerve. | MedDRA:10019211:Headache | Nervous System | Generalized Disorders |  |  | E23¦E2330¦E0116
|  | Eye Pain | E0820 | Painful sensation in the eye. | MedDRA:10015958:Eye pain | Eye | Generalized Disorders |  |  | E23¦E2330¦E0820
|  | Ear Pain | E0907 | Earache or ear pain. | MedDRA:10014020:Ear pain | Ear and Labyrinth | Generalized Disorders |  | Term was added on 27 January 2022. For details, see comment No. 234 / 235 of the Change Log (Release Number 2022). | E23¦E2330¦E0907
|  | Abdominal Pain | E1002 | Painful sensation in the abdominal region. | MedDRA:10000081:Abdominal pain | Gastrointestinal System | Generalized Disorders |  |  | E23¦E2330¦E1002
|  | Myalgia | E1622 | Pain sensation originating from a muscle or group of muscles. | MedDRA:10028411:Myalgia | Musculoskeletal System | Generalized Disorders |  |  | E23¦E2330¦E1622
|  | Neck Pain | E1623 | Pain in the neck area. | MedDRA:10028836:Neck pain | Musculoskeletal System | Generalized Disorders |  |  | E23¦E2330¦E1623
|  | Back pain | E1651 | Pain in any part of the back. | MedDRA:10003988:Back pain | Musculoskeletal System | Generalized Disorders |  | Term added for Release Number 2024. See no. 107 of Change Log. Editorial change to term for Release Number 2026. See no. 285 of Change Log. | E23¦E2330¦E1651
|  | External Prosthetic Device Pain | E2106 | Pain that is associated with an external artificial replacement device. | MedDRA:10059057:Medical device pain | Procedural Complications | Generalized Disorders |  | Definition was modified on 21 January 2021 to correct a typo in "artificial" (Old definition: Pain that is associated with an external ariticial replacement device.) | E23¦E2330¦E2106
|  | Implant Pain | E2109 | Pain localized to the site of the implanted device. | MedDRA:10063782:Implant site pain | Procedural Complications | Generalized Disorders |  |  | E23¦E2330¦E2109
|  | Inadequate Pain Relief | E2110 | Continuing discomfort due to a lack or failure of a medical approach to treat pain. | MedDRA:10079788:Inadequate pain relief | Procedural Complications | Generalized Disorders |  |  | E23¦E2330¦E2110
|  | Intraoperative Pain | E2112 | Pain occurring during a surgical operation. | MedDRA:10064882:Procedural pain | Procedural Complications | Generalized Disorders |  |  | E23¦E2330¦E2112
|  | Chest Pain | E233001 | Pain in the chest. | MedDRA:10008479:Chest pain |  |  |  |  | E23¦E2330¦E233001
| Pallor |  | E2331 | Unhealthy, pale or white appearance of the face. | MedDRA:10033546:Pallor |  |  |  |  | E23¦E2331
| Physical Asymmetry |  | E2332 | Lack or absence of balanced proportions between parts or organs of the body. |  | Generalized Disorders | Procedural Complications |  |  | E23¦E2332
| Prolapse |  | E2333 | A condition in which an organ or body tissue drops or bulges out of place. | MedDRA:10076708:Prolapse |  |  |  |  | E23¦E2333
| Ptosis |  | E2334 | The drooping or downward displacement of a body part. | MedDRA:10037272:Ptosis |  |  |  |  | E23¦E2334
| Reaction to Medicinal Component of Device |  | E2335 | Any local or systemic adverse reaction to the medicinal component of a combination product. | MedDRA:10067484:Adverse reaction |  |  |  |  | E23¦E2335
| Shock |  | E2336 | A life-threatening condition that requires immediate medical intervention. It is characterized by reduced blood flow that may result in damage of multiple organs. | MedDRA:10040560:Shock |  |  |  | Editorial change to the definition on 16 February 2021. | E23¦E2336
|  | Anaphylactic Shock | E040201 | Progression of anaphylaxis into systemic shock, which is characterized by significant reduction in tissue perfusion which leads to hypoxia and hypotension, and can lead to death if untreated. | MedDRA:10002199:Anaphylactic shock | Immune System | Generalized Disorders |  |  | E23¦E2336¦E040201
|  | Cardiogenic Shock | E233601 | Shock resulting from primary failure of the heart in its pumping function, as in myocardial infarction, severe cardiomyopathy, or mechanical obstruction or compression of the heart. | MedDRA:10007625:Cardiogenic shock |  |  |  |  | E23¦E2336¦E233601
|  | Hypoglycemic Shock | E233602 | A hypoglycemic reaction to overdosage of insulin, a skipped meal, or strenuous exercise in an insulin-dependent diabetic. | MedDRA:10060493:Hypoglycemic shock |  |  |  |  | E23¦E2336¦E233602
|  | Hypovolemic Shock | E233603 | Shock resulting from insufficient blood volume for the maintenance of adequate cardiac output, blood pressure and tissue perfusion. | MedDRA:10021140:Hypovolemic shock |  |  |  |  | E23¦E2336¦E233603
|  | Neurogenic Shock | E233604 | Shock resulting from neurogenic vasodilation, which can be produced by cerebral trauma or hemorrhage, spinal cord injury, deep general or spinal anesthesia, or toxic central nervous system depression. | MedDRA:10058119:Neurogenic shock |  |  |  |  | E23¦E2336¦E233604
|  | Septic Shock | E233605 | A state of acute circulatory failure resulting from organ injury or damage in response to infection that results in dangerously low blood pressure. | MedDRA:10040070:Septic shock |  |  |  |  | E23¦E2336¦E233605
| Stenosis |  | E2337 | Narrowing or stricture of a vessel, duct or canal. | MedDRA:10076711:Stenosis | Generalized Disorders | Vascular System |  | Secondary category was added on 27 January 2022. For details, see comment No. 234 / 235 of the Change Log (Release Number 2022). Error in CodeHierarchy corrected for Release No. 2024 | E23¦E2337
|  | Restenosis | E233701 | The reoccurrence of stenosis after it has been treated. | MedDRA:10082493:Restenosis |  |  |  |  | E23¦E2337¦E233701
| Swelling/ Edema |  | E2338 | Enlargement or expansion in size of a body part from injury or inflammation. | MedDRA:10014210:Edema |  |  |  |  | E23¦E2338
|  | Cerebral Edema | E0103 | A swelling in the brain caused by the presence of excessive fluid. | MedDRA:10008107:Cerebral edema | Nervous System | Generalized Disorders |  |  | E23¦E2338¦E0103
|  | Pulmonary Edema | E0736 | Accumulation of fluid in the lung tissues causing disturbance of the gas exchange that may lead to respiratory failure. | MedDRA:10037375:Pulmonary edema | Respiratory System | Generalized Disorders |  |  | E23¦E2338¦E0736
|  | Corneal Edema | E0807 | Hazy, swollen cornea. | MedDRA:10011007:Corneal edema | Eye | Generalized Disorders |  |  | E23¦E2338¦E0807
|  | Corneal Stromal Edema | E080701 | Fluid from the aqueous humor enters the corneal stroma and causes swelling. | MedDRA:10054772:Corneal stromal edema | Eye | Generalized Disorders |  |  | E23¦E2338¦E080701
|  | Microcystic Edema | E080702 | Corneal edema of inflammatory origin associated with contact lens wear and morphologically characterized by presence of epithelial microcysts. | MedDRA:10068529:Corneal microcystic edema | Eye | Generalized Disorders |  |  | E23¦E2338¦E080702
|  | Macular Edema | E0830 | Accumulation of intraretinal fluid and protein in the macula, which may result in swelling and decreased central vision. | MedDRA:10054467:Macular edema | Eye | Generalized Disorders |  |  | E23¦E2338¦E0830
|  | Angioedema | E1702 | Rapid swelling of the deep layers of the skin due to transient vascular leakage of serous fluid. | MedDRA:10002424:Angioedema | Skin and Subcutaneous Tissue | Generalized Disorders |  |  | E23¦E2338¦E1702
|  | Peripheral Edema | E233801 | accumulation of fluid in tissues perfused by the peripheral vascular system. | MedDRA:10034570:Peripheral edema |  |  |  |  | E23¦E2338¦E233801
|  | Joint Effusion | E233802 | An abnormal collection of fluid in hollow spaces or between tissues of the joints. | MedDRA:10023215:Joint effusion |  |  | Modified (editorial) | Term was added on 27 January 2022. For details, see comment No. 115 of the Change Log (Release Number 2022). Editorial change to term for Release Number 2026. See no. 302 of Change Log. | E23¦E2338¦E233802
| Ulcer |  | E2339 | An open sore on an external or internal surface of the body caused by a break in the skin or mucous membrane that fails to heal. | MedDRA:10045285:Ulcer |  |  |  |  | E23¦E2339
| Wound Dehiscence |  | E2340 | A complication in which the wound splits or separates after being surgically closed. | MedDRA:10048031:Wound dehiscence | Generalized Disorders | Procedural Complications |  |  | E23¦E2340
| Double Capsule |  | E2341 | The presence of two distinct capsular layers around an implant with an intercapsular space |  | Generalized Disorders | Reproductive System and Breast |  | Term was added on 18 March 2020 as an additional level 2 term for "Generalized Disorders". | E23¦E2341
| Multiple Organ Dysfunction Syndrome |  | E2342 | Complete impairment of two or more organs or organ systems. Also known as multiple organ failure. | MedDRA:10077361:Multiple organ dysfunction syndrome |  |  |  | Term was added on 21 January 2021. For details, refer to comment No. 34 of the Change Log. | E23¦E2342
| Hemodynamic Instability |  | E2343 | Abnormal or unstable blood pressure | MedDRA:10058270:Hemodynamic instability |  |  | Modified (editorial) | Term was added on 31 January 2023. For details, see comment No. 101 of the Change Log (Release Number 2023). Editorial change to term for Release Number 2026. See no. 303 of Change Log. | E23¦E2343
| Voice Alteration |  | E2344 | Changes in the pitch or quality of the voice, which may sound weak, scratchy, raspy, or otherwise out of the ordinary. | MedDRA:10047681:Voice alteration |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 3 of Change Log. Editorial change to term for Release Number 2026. See no. 304 of Change Log. | E23¦E2344
| Heat Illness |  | E2345 | A set of symptoms brought on by exposure to heat that can include confusion, hot dry skin, seizures, elevated body temperature, altered mental status, etc. Includes heat stroke and heat exhaustion. | MedDRA:10068188:Heat illness |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 4 of Change Log. Editorial change to term for Release Number 2026. See no. 305 of Change Log. | E23¦E2345
| Hypertrophy |  | E2346 | An increase in organ size via increase in cell size rather than cell number | MedDRA:10020880:Hypertrophy |  |  |  | Term added for Release Number 2024. See no. 11 of Change Log | E23¦E2346
| Breath Odor |  | E2347 | Unusual breath odor (e.g., halitosis, fruity/sweet breath) | MedDRA:10006325:Breath odor |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 16 of Change Log. Editorial change to term for Release Number 2026. See no. 306 of Change Log. | E23¦E2347
| Abnormal Mucus Discharge |  | E2348 | Any deviation from the normal quantity, consistency, and/or color of mucus produced by the mucous membranes. | MedDRA:10073765:Mucus discharge |  |  | Modified (editorial) | Term added for Release Number 2024. See no. 20 of Change Log. Editorial change to term for Release Number 2026. See no. 307 of Change Log. | E23¦E2348
| Withdrawal Syndrome |  | E2349 | The substance-specific organic brain syndrome that follows the discontinuation of, or reduction in, administration or intake of an addictive substance. | MedDRA:10048010:Withdrawal syndrome |  |  |  | Term added for Release Number 2024. See no. 24 of Change Log | E23¦E2349
| Tissue Irritation |  | E2350 | Irritation of body tissue(s), other than skin. For skin irritation use E1720 - Skin Inflammation/Irritation | MedDRA:10075582:Medical device site irritation |  |  |  | Term added for Release Number 2024. See no. 89 of Change Log. | E23¦E2350
| Atrophy |  | E2351 | Decrease in size of a body part, cell, organ, or other tissue. For muscular atrophy use E1621 - Muscle Weakness/Atrophy | MedDRA:10003694:Atrophy |  |  |  | Term added for Release Number 2024. See no. 93 of Change Log. | E23¦E2351
| Pelvic Pain |  | E1420 | Pain in the pelvic area | MedDRA:10034263:Pelvic pain | Reproductive System and Breast | Generalized Disorders |  | Term was added on 31 January 2023. For details, see comment No. 93 of the Change Log (Release Number 2023). | E23¦E1420
| Exercise Capacity Decreased |  | E2352 | Reduced ability of the body to perform strenuous activities | MedDRA:10052325:Exercise capacity decreased |  |  | Modified (editorial) | Term added for Release Number 2025. See no. 18 of Change Request Log. Editorial change to term for Release Number 2026. See no. 308 of Change Log. | E23¦E2352
| Scab |  | E2353 | a dry, rough protective crust that forms over a cut or wound during healing. | MedDRA:10039509:Scab |  |  | New | Term added for Release Number 2026. See no. 163 of Change Request Log | E23¦E2353
| Periodontal Disease |  | E2354 | A chronic inflammatory condition affecting the supporting structures of the teeth, including the gingiva, periodontal ligament, and alveolar bone | MedDRA:10034536:Periodontal disease |  |  | New | Term added for Release Number 2026. See no. 25 of Change Request Log | E23¦E2354
Other Terms/Codes Related to Clinical Signs and Symptoms or Conditions |  |  | E24 | Others |  |  |  | Not selectable | Please use more detailed term within the hierarchy. NOTE: Term modified for Release Number 2025. Previous term was "Others". See no. 141 of Change Request Log. | E24
| Insufficient Information |  | E2401 | It is not clear whether any health effects occurred, or health effect appears to have occurred but there is not yet enough information available to classify the clinical signs, symptoms and conditions. |  |  |  |  | Definition was modified on 31 January 2023 (Old definition: Health effect appears to have occurred but there is not yet enough information available to classify the clinical signs, symptoms and conditions.). For details, see comment No. 102 of the Change Log (Release Number 2023). | E24¦E2401
| Appropriate Clinical Signs, Symptoms and Conditions Term/Code not Available |  | E2402 | The clinical signs, symptoms and conditions are not adequately described by any other term. Note: this code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. |  |  |  | Modified (editorial) | Term modified for Release Number 2025. Previous term was "Appropriate Term / Code Not Available". See no. 110 of Change Request Log. Editorial change to term for Release Number 2026. See no. 305 of Change Log. | E24¦E2402
| No Clinical Signs, Symptoms or Conditions |  | E2403 | No patient involvement or, no observable clinical symptoms or a change in symptoms is identified in the patient. | [5 empty] | E24¦E2403

## F
Annex Name: Annex F
Annex Title: Health Effects - Health Impact
Release Number: 2026
Annex Description: Terms/codes for describing the consequences of the medical device adverse event/incident on the person affected.
Annex Instructions: Terms/Codes in this annex describe the resulting consequences of the medical device adverse event/incident on the person affected. The resulting consequences can include final patient outcomes and/or interventions or procedures required as a result of the clinical signs, symptoms and conditions captured using Annex E. It is likely that it will be necessary to use multiple terms. IMDRF Terms and Definitions use American spelling. Wherever appropriate the word 'device' should be interpreted as meaning all or any of the following: medical device, in-vitro diagnostic medical device, accessory to a medical device, medical device parts and/or components. Wherever appropriate the word 'patient' should be interpreted as meaning all or any of the following: user, operator, any other person affected by the incident.
Annex Approval Date: January 2026
NOTE: Maximum field length Term: 100 characters Definition: 700 Status description: 700
Level 1 Term | Level 2 Term | Level 3 Term | Code | Definition | Non-IMDRF Code/Term | Status | Status Description | CodeHierarchy
Change in Therapeutic Response |  |  | F01 | Change in response to treatment or cure of a disorder or disease. |  |  |  | F01
| Therapeutic Response Decreased |  | F0101 | A reduction in or complete loss of the desirable and beneficial effects resulting from a medical treatment. |  | Modified (technical) | Definition modified for Release Number 2026. See no. 114 of Change Request Log | F01¦F0101
| Therapeutic Response Increased |  | F0102 | An increase in the desirable and beneficial effects resulting from a medical treatment. |  |  |  | F01¦F0102
| Unexpected Therapeutic Effects |  | F0103 | Unanticipated desirable and beneficial effects resulting from a medical treatment. |  |  |  | F01¦F0103
| Delay in Therapeutic Response |  | F0104 | An increase in the time taken to achieve the desirable and beneficial effects resulting from a medical treatment |  | New | Term added for Release Number 2026. See no. 33 of Change Request Log | F01¦F0104
Death |  |  | F02 | The cessation of life. |  |  |  | F02
| Intrauterine Fetal Death |  | F0201 | Death in utero; failure of the product of conception to show evidence of respiration, heart beat, or definite movement of a voluntary muscle after expulsion from the uterus, with no possibility of resuscitation. |  |  |  | F02¦F0201
Brain Death |  |  | F03 | The cessation of life due to permanent loss of brain function and brainstem for absence of responsiveness to external stimuli, absence of reflex and apnea. |  |  |  | F03
Delay to Diagnosis |  |  | F04 | Patient diagnosis was clinically significant delayed as a consequence of device performance. |  |  |  | F04
Delay to Treatment/ Therapy |  |  | F05 | Patient treatment was delayed or interrupted as a consequence of device performance. |  | Modified (technical) | Definition modified for Release Number 2026. See no. 178 of Change Request Log | F05
Disruption of Subsequent Medical Procedure |  |  | F06 | Situation in which the use of the device impedes or affects a subsequent medical procedure or use of a medicine or device. The time elapsed between the use of the device and the medical procedure is not a factor. It is not necessary for the device to have broken or malfunctioned. |  |  |  | F06
Exacerbation of Existing Condition |  |  | F07 | Use of the device has lead to worsening of the existing disease or condition, as a direct consequence of device use |  | Modified (technical) | Definition modified for Release Number 2026. See no. 12 of Change Request Log | F07
Hospitalization or Prolonged Hospitalization |  |  | F08 | Hospitalization or prolonged hospitalization due to the health damage accompanying the use of device. |  |  |  | F08
| Intensive Care |  | F0801 | Patient requires admission to or extension of stay in an intensive care unit. |  |  |  | F08¦F0801
| Emergency Hospital Transfer |  | F0802 | Patient required emergency transfer to a different facility because of the need for an additional level of care. |  |  | Term added for Release Number 2025. See no. 28 of Change Request Log | F08¦F0802
Fetal Harm |  |  | F09 | Fetal distress, congenital physical or mental impairment or birth defect, excluding fetal death. |  |  |  | F09
Inadequate/Inappropriate Treatment or Diagnostic Exposure |  |  | F10 | Patient did not receive intended treatment from a device or received inadequate treatment as a consequence of device performance. This includes inadequate/inappropriate electrical stimulation or radiation dosage during procedures. |  |  | Definition modified for Release Number 2025. See no. 19 of Change Request Log | F10
| Absence of Treatment |  | F1001 | Patient treatment was absent as a consequence of device performance. |  |  |  | F10¦F1001
| Incompatible Blood Transfusion |  | F1002 | Receiving incorrect or incompatible whole blood or a blood component. |  |  |  | F10¦F1002
| Missed Dose |  | F1003 | A dose of medicine was not taken or delivered at the prescribed dosing interval. |  |  |  | F10¦F1003
| Underdose |  | F1004 | Exposure to, or intake of, an insufficient amount of a substance. |  |  |  | F10¦F1004
|  | Radiation Underdose | F100401 | A lower than intended dose of radiation to a single field. |  |  |  | F10¦F1004¦F100401
| Overdose |  | F1005 | Exposure to, or intake of, an excessive amount of a substance. |  |  |  | F10¦F1005
|  | Radiation Overdose | F100501 | A higher than intended cumulative dose of radiation to a single field. |  |  |  | F10¦F1005¦F100501
| Inadequate Treatment/Disease Progression |  | F1006 | The worsening of a disease over time due to inadequate treatment. |  | Retired (2025) | Term was added on 27 January 2022. For details, see comment No. 191 of the Change Log (Release Number 2022). Term retired for Release Number 2025. Use F1010 "Inadequate Treatment" or F30 "Disease Progression". See no. 63 of Change Request Log | F10¦F1006
| Exposure to Contaminated Device/Risk of Infection |  | F1007 | Exposure to a non-sterile device or conditions where infection is likely. |  |  | Term was added on 27 January 2022. For details, see comment No. 214 of the Change Log (Release Number 2022). | F10¦F1007
| Accidental Exposure to Substance |  | F1008 | Unintended exposure of the patient and/or healthcare providers to drugs or other substances. This can include any exposure routes such as skin, mucous membranes, and respiratory. |  |  | Term added for Release Number 2024. See no. 25 of Change Log. Formatting change to Term for Release Number 2025. See no. 21 of Change Request Log | F10¦F1008
| Inappropriate Treatment |  | F1009 | Patient received treatment which was unnecessary or excessive. This includes treatment following a false positive diagnosis and excessive stimulation or shock delivery. For drug or radiation overdose use F1005 or child term. |  |  | Term added for Release Number 2025. See no. 20 of Change Request Log | F10¦F1009
| Inadequate Treatment |  | F1010 | Patient received inadequate or insufficient treatment of an appropriate type. This includes inadequate stimulation or shock delivery. For inappropriate treatment use F1009. For drug or radiation underdose use F1004 or child term. |  |  | Term added for Release Number 2025. See no. 63 of Change Request Log | F10¦F1010
Non-Serious Injury/ Illness/ Impairment |  |  | F11 | A non-serious injury, illness or impairment which can be treated with minimal or no intervention, including monitoring only. |  |  | Definition was modified on 31 January 2023 (Old definition: A mild injury, illness or impairment which can be treated with minimal or no intervention.). For details, see comment No. 143 of the Change Log (Release Number 2023). Term and Definition modified for Release Number 2025. Previous term was "Minor Injury/ Illness / Impairment". See no. 79 of Change Request Log. | F11
Serious Injury/ Illness/ Impairment |  |  | F12 | A severe injury, illness or impairment which requires hospitalization or medical intervention. |  |  |  | F12
| Chronic Disease |  | F1201 | Disease that persists over a long period of time, cannot be prevented by vaccines or cured by medication, nor do they just disappear. |  |  |  | F12¦F1201
| Disability |  | F1202 | Any physical or mental impairment that interferes with an individual's ability to perform desired activities. |  |  |  | F12¦F1202
| Life Threatening Illness or Injury |  | F1203 | Patient suffered an illness or injury which if not treated would be fatal. |  |  |  | F12¦F1203
| Permanent Impairment |  | F1204 | Irreversible deterioration of the state of health. |  |  |  | F12¦F1204
| Temporary Impairment |  | F1205 | Reversible deterioration of the state of health. |  |  |  | F12¦F1205
Misdiagnosis/ Misclassification |  |  | F13 | An incorrect diagnosis or classification of an illness or other problem. |  |  |  | F13
Prolonged Episode of Care |  |  | F14 | Clinically significant extended treatment time or procedural time. Not to be used for prolonged hospitalization or prolonged surgery. |  |  | Editorial change to the definition on 16 February 2021. | F14
Recognized Device or Procedural Complication |  |  | F15 | Patient experiencing a complication that is well documented in association with the device or procedure. |  | Modified (editorial) | Editorial change to term for consistency for Release Number 2026. See no. 191 of Change Log. | F15
Reduction in Life Expectancy |  |  | F16 | Decrease in lifetime/ lifespan of an individual. Decline in the normal duration or in average life. |  |  |  | F16
Sedation |  |  | F17 | Patient required sedation or additional sedation. |  |  |  | F17
Rehabilitation |  |  | F18 | Patient requires treatment to facilitate the process of recovery from injury, illness, or disease. |  |  |  | F18
Surgical Intervention |  |  | F19 | One or more surgical procedures, including percutaneous or minimally invasive procedures, were required, or an existing procedure changed. |  |  | Definition was modified on 31 January 2023 (Old definition: One or more surgical procedures was required, or an existing procedure changed). For details, see comment No. 135 /137/138 /139 /140/141 of the Change Log (Release Number 2023). | F19
| Additional Surgery |  | F1901 | Surgical intervention that was not planned and needed to be performed in addition to other interventions including surgical ones. |  |  |  | F19¦F1901
| Amputation |  | F1902 | Patient required surgical removal of limb, extremities or other external body parts. |  |  |  | F19¦F1902
| Device Explantation |  | F1903 | Patient required a corrective invasive procedure to remove a device, or parts/fragments of a device. |  |  |  | F19¦F1903
| Device Repositioning |  | F1904 | Patient required a corrective invasive procedure in which a device was repositioned. This includes implantable, surgically invasive (penetration through the surface of the body) or invasive (penetration though a natural body orifice or permanent artificial opening, e.g. stoma) devices. |  |  |  | F19¦F1904
| Device Revision or Replacement |  | F1905 | Patient required a corrective invasive procedure in which a device was replaced/revised. This includes implantable, surgically invasive (penetration through the surface of the body) or invasive (penetration though a natural body orifice or permanent artificial opening, e.g. stoma) devices. |  |  |  | F19¦F1905
| Modified Surgical Procedure |  | F1906 | A change of surgical procedures from the original plan. |  |  |  | F19¦F1906
| More Complex Surgery |  | F1907 | Surgery that is more complicated, more comprehensive or more extensive than planned (e.g. percutaneous to open surgery). |  |  | Definition was modified on 31 January 2023 (Old definition: Patient required additional or unexpected endoscopic procedure). For details, see comment No. 139 of the Change Log (Release Number 2023). | F19¦F1907
|  | Conversion from Robotic-Assisted Procedure to Other Surgical Approach | F190701 | A surgical procedure initially intended to be completed using robotic assistance that is converted intraoperatively to a different surgical approach. This change may result from robotic system limitations, technical issues, or surgeon preference to complete the procedure safely and effectively. |  | New | Term added for Release Number 2026. See no. 115 of Change Request Log. | F19¦F1907¦F190701
| Prolonged Surgery |  | F1908 | Surgery that takes longer than anticipated. |  |  | Formatting change to Term for Release Number 2025. See no. 21 of Change Request Log | F19¦F1908
| Surgical Procedure Delayed |  | F1909 | Surgical procedure unable to be completed when originally scheduled and put off until some time in the future. |  |  |  | F19¦F1909
| Canceled/Aborted Surgical Procedure |  | F1910 | Surgical procedure aborted, or unable to be completed. This applies to cases where there is no information whether the procedure will be carried out at a later time, or where it is known that surgery will not go ahead. If it is known that the procedure has been, or will be, carried out later then use F1909 - Surgical Procedure Delayed |  | Modified (editorial) | Term added for Release Number 2024. See no. 110 of Change Log. Formatting change to Term for Release Number 2025. See no. 21 of Change Request Log. Editorial change to term for consistency for Release Number 2026. See no. 192 of Change Log. | F19¦F1910
Serious Public Health Threat |  |  | F20 | Any event type, which results in imminent risk of death, serious injury, or serious illness to more than one individual, that requires prompt remedial action. |  |  |  | F20
Unexpected Deterioration |  |  | F21 | Unexpected clinical deterioration of patient. |  |  |  | F21
Unexpected Diagnostic Intervention |  |  | F22 | Diagnostic intervention needed to be completed that was not originally planned. |  |  |  | F22
| Biopsy |  | F2201 | Patient required removal of tissue or fluid specimen to establish a diagnosis, including biopsy and blood samples. |  | Modified (technical) | Technical change to definition for Release Number 2026. See 5 of Change Log | F22¦F2201
| Endoscopic Diagnostic Procedure |  | F2202 | Patient required additional or unexpected endoscopic diagnostic procedure. |  |  | Definition was modified on 31 January 2023 (Old definition: Patient required additional or unexpected endoscopic procedure). For details, see comment No. 137 of the Change Log (Release Number 2023). | F22¦F2202
| Imaging Required |  | F2203 | Requirement for imaging that was not originally planned. |  |  |  | F22¦F2203
| IVD Testing |  | F2204 | Patient required additional or unexpected IVD testing. |  |  | Editorial change to the definition on 16 February 2021. | F22¦F2204
Unexpected Medical Intervention |  |  | F23 | Patient required an unforeseen medical intervention, excluding surgery, which was not on the original treatment plan. |  |  |  | F23
| Additional Device Required |  | F2301 | Use of an additional or alternative device require to achieve optimal outcome. |  |  |  | F23¦F2301
|  | Additional Device of Same Type Required | F230101 | The situation where an additional device of the same type and intended purpose as the original device is required to continue or complete the intended procedure or therapy. |  | New | Term added for Release Number 2026. See no. 35 of Change Request Log | F23¦F2301¦F230101
|  | Additional Device of Different Type Required | F230102 | The situation where an additional device of a different type or intended purpose is required to mitigate risk or prevent harm to the patient following the use or failure of the original device. This may involve introducing a complementary or alternative device to address complications |  | New | Term added for Release Number 2026. See no. 35 of Change Request Log | F23¦F2301¦F230102
| Blood Transfusion |  | F2302 | Patient required an infusion of whole blood or a blood component directly into the bloodstream. |  |  |  | F23¦F2302
| Medication Required |  | F2303 | Patient required new/additional medication or additional dose of existing medication. |  |  |  | F23¦F2303
| Prophylactic Treatment |  | F2304 | A medication or course of action used to prevent a possible disease or condition from occurring. |  |  | Definition modified for Release Number 2025. See no. 99 of Change Request Log | F23¦F2304
| Radiation Therapy |  | F2305 | Patient required additional radiation therapy. |  |  | Editorial change to the definition on 16 February 2021. | F23¦F2305
| Resuscitation |  | F2306 | Patient required resuscitation. |  |  |  | F23¦F2306
| Reprogramming of Device |  | F2307 | Patient required programming changes to be made to an implantable or other programmable device. |  |  | Term was added on 31 January 2023. For details, see comment No. 133 of the Change Log (Release Number 2023). | F23¦F2307
| Medical Treatment Under Anaesthesia |  | F2308 | Patient required non-surgical treatment under anaesthesia. |  | Modified (editorial) | Term added for Release Number 2024. See no. 111 of Change Log. Editorial change to term for consistency for Release Number 2026. See no. 310 of Change Log. | F23¦F2308
| Ventilation |  | F2309 | Patient required ventilation or clinically significant extension of existing ventilation |  |  | Term added for Release Number 2025. See no. 27 of Change Request Log | F23¦F2309
Insufficient Health Impact Information |  |  | F24 | It is not clear whether any health impact occurred, or health impact appears to have occurred, but there is not yet enough information available to classify the health impact. |  |  | Definition was modified on 31 January 2023 (Old definition: There is not yet enough information available to classify the health impact). For details, see comment No. 102 of the Change Log (Release Number 2023). Term and Definition modified for Release Number 2025. Previous term was "Insufficient Information". See no. 122 of Change Request Log. | F24
Unanticipated Adverse Device Effect |  |  | F25 | Complication occurring in the patient that had not previously been identified or expected for this device. |  |  | Editorial change to the definition on 16 February 2021. | F25
No Health Consequences or Impact |  |  | F26 | No apparent harm occurred in relation to the adverse event. |  |  | Editorial change to the definition on 16 February 2021. | F26
| Problem Identified Before Clinical Use/Exposure |  | F2601 | The adverse event occurred during or immediately prior to a clinical procedure, but the problem was identified and corrected before harm could be caused. E.g.1. an IVD gave a false reading but it was recognized as incorrect before treatment decisions were made and without any clinically significant delay to diagnosis; E.g.2. a surgical device malfunctioned during pre-use checks and the malfunction was corrected or an alternative device was available without any clinically significant delay to the procedure. |  | Modified (editorial) | Term added for Release Number 2024. See no. 120 of Change Log. Editorial change to term for consistency for Release Number 2026. See no. 311 of Change Log. | F26¦F2601
Problem Identified During Non-Clinical Procedure |  |  | F27 | The adverse event occurred while the device was being used in a non-clinical setting and/or procedure and could not have any impact on a patient E.g.1. an IVD gave a false reading when testing an External Quality Assessment (EQA) sample; E.g.2. a surgical device malfunctioned during cleaning and/or maintenance. |  | Modified (editorial) | Term and Definition modified for Release Number 2024. Previous term was No Patient Involvement. See no. 121 of Change Log. Editorial change to term for consistency for Release Number 2026. See no. 312 of Change Log. | F27
Appropriate Health Impact Term/Code not Available |  |  | F28 | The Health Impact is not adequately described by any other term. Note: this code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. |  | Modified (editorial) | Term and Definition modified for Release Number 2025. Previous term was "Appropriate Term/Code Not Available". See no. 111 of Change Request Log. Editorial change to term for consistency for Release Number 2026. See no. 313 of Change Log. | F28
Death not Related to Reported Adverse Event |  |  | F29 | A death has occurred, but it was confirmed that the event did not contribute to the death. Other terms may also be selected if it is possible to link the adverse event to other identifiable health impacts. |  | Modified (editorial) | Term added for Release Number 2024. See no. 119 of Change Log. Editorial change to term for consistency for Release Number 2026. See no. 314 of Change Log. | F29
Disease Progression |  |  | F30 | The worsening of a disease over time related to device performance (e.g. insufficient therapeutic effect, delayed action). |  | Modified (editorial) | Definition modified for Release Number 2026. See no. 13 of Change Request Log | F30
Canceled/Aborted Treatment/Therapy |  |  | F31 | Patient treatment aborted, or unable to be completed. This applies to cases where there is no information whether the treatment will be carried out at a later time, or where it is known that treatment will not go ahead. If it is known that the treatment has been, or will be, carried out later then use F05 - Delay to Treatment/ Therapy |  | Modified (editorial) | Term added for Release Number 2025. See no. 84 of Change Request Log. Editorial change to term for consistency for Release Number 2026. See no. 193 of Change Log. | F31

## G
Annex Name: Annex G
Annex Title: Medical Device Component
Release Number: 2026
Annex Description: Terms/codes for describing the parts and components which were involved in, or affected by, the medical device adverse event/incident.
Annex Instructions: These terms allow capturing of the specific part or component of the medical device which was involved in the incident. In the case that a component is also a stand-alone device, the Annex G term should not be used. IMDRF Terms and Definitions use American spelling. Wherever appropriate the word 'patient' should be interpreted as meaning all or any of the following: user, operator, any other person affected by the incident.
Annex Approval Date: January 2026
NOTE: Maximum field length Term: 100 characters Definition: 700 Status description: 700
Level 1 Term | Level 2 Term | Level 3 Term | Code | Definition | Non-IMDRF Code | Status | Status Description | CodeHierarchy
Biological and Chemical |  |  | G01 | Component whose mode of action involves a biological process (ex. test strip which acts on antibodies) and chemical reaction or transformation (ex. activated charcoal absorber). |  | Not selectable | Please use more detailed term within the hierarchy | G01
| Absorber |  | G01001 | A component or material designed to take in or attenuate a substance. |  |  |  | G01¦G01001
| Cautery Tip |  | G01002 | A component designed to coagulate and seal blood vessels or to destroy tissue with heat or electric current. |  |  |  | G01¦G01002
| Device Ingredient or Reagent |  | G01003 | A consumable material that is added to a device and is used to make a finished product or becomes part of a finished product. |  |  |  | G01¦G01003
| Gas Scavenging |  | G01004 | A component designed to remove certain gases from a gas stream or environment. |  |  |  | G01¦G01004
| Monomer Liquid |  | G01005 | The liquid that reacts with the polymer powder to form adhesive, resin, or cement. |  |  |  | G01¦G01005
| Test Strip |  | G01006 | A piece of chemically treated medium designed to react in the presence of specific amounts of a trigger substance. |  |  |  | G01¦G01006
| Polymer Powder |  | G01007 | The powder that reacts with the monomer liquid to form adhesive, resin, or cement. |  |  |  | G01¦G01007
Electrical and Magnetic |  |  | G02 | Components which relates to the force of magnetism or such as microchips and transistors which control and direct electric currents. Also includes components involved in the representation, storage, or transmission of information by electronic systems. |  | Not selectable | Please use more detailed term within the hierarchy | G02
| Antenna |  | G02001 | A component designed to transmit or receive electromagnetic signals. |  |  |  | G02¦G02001
| Battery |  | G02002 | A component designed to produce an electric current through chemical reaction. |  |  |  | G02¦G02002
| Battery Charger |  | G02003 | A device designed to restore the capacity of a battery. |  |  |  | G02¦G02003
| Cable, Electrical |  | G02004 | A long, thin, multistranded electrical wire designed to carry signals or power over a distance. |  |  |  | G02¦G02004
|  | Cable Grip/Tie | G0200401 | Component used for tensioning, pulling, holding or stringing of wires and cables. |  |  | Term/definition was modified on 31 January 2023 (Old term/definition: Cable Grip: Component used for tensioning, pulling or stringing of wires and cables.). For details, see comment No. 151 of the Change Log (Release Number 2023). | G02¦G02004¦G0200401
|  | Cable Sleeve | G0200402 | Component used to protect cables and wires from abrasion, moisture and the elements. |  |  |  | G02¦G02004¦G0200402
| Circuit Board |  | G02005 | A non-conducting board with conductive tracks and electronic components forming a circuit. |  |  |  | G02¦G02005
| Circuit Breaker |  | G02006 | A component designed to open an electrical circuit when it becomes overloaded. |  |  |  | G02¦G02006
| Computer Hardware |  | G02007 | The physical components from which a computer is constructed (electronic circuits and input/output devices). |  |  |  | G02¦G02007
|  | Computer Processor | G0200701 | Component that carries out the instructions of a computer program by performing the basic arithmetic, logic, controlling, and input/output operations specified by the instructions. |  |  |  | G02¦G02007¦G0200701
|  | Memory/Storage | G0200702 | Any component that can hold data in machine-readable format. |  |  |  | G02¦G02007¦G0200702
|  | Network Interface | G0200703 | Point of interconnection between a computer and other computer that are linked each other. |  |  |  | G02¦G02007¦G0200703
| Computer Software |  | G02008 | A collection of data or computer instructions that tell the computer how to work. |  |  |  | G02¦G02008
|  | Driver | G0200801 | A computer interface designed to control the interaction between a CPU and a peripheral device. |  |  |  | G02¦G02008¦G0200801
|  | Software Interface | G0200802 | Languages, codes and messages that programs use to communicate with each other and to the hardware. |  |  |  | G02¦G02008¦G0200802
|  | User Interface | G0200803 | A computer program that controls the interaction between a user and a system. |  |  |  | G02¦G02008¦G0200803
|  | Firmware | G0200804 | Computer instructions and data that reside as read-only software on that device. The computer cannot modify such software during processing. |  |  | Term was added on 31 January 2023. For details, see comment No. 144 of the Change Log (Release Number 2023). | G02¦G02008¦G0200804
| Cooling Module |  | G02009 | A component designed to lower the temperature of a device or system. |  |  |  | G02¦G02009
| Device Programmer |  | G02010 | A piece of hardware for transferring data onto programmable integrated circuits. |  |  |  | G02¦G02010
| Device Reader |  | G02011 | A piece of hardware used to read the memory and the properties of a device. |  |  | Editorial change to the definition on 16 February 2021. | G02¦G02011
| Discrete Electrical Component |  | G02012 | An electrical component which is just one circuit, either passive or active, that is not an integrated circuit. |  |  |  | G02¦G02012
|  | Capacitor | G0201201 | An electrical component designed to store an electric charge. |  |  |  | G02¦G02012¦G0201201
|  | Fuse | G0201202 | An electrical component designed to stop the flow of current when an overload condition exists. |  |  |  | G02¦G02012¦G0201202
|  | Inductor | G0201203 | A component designed to introduce electromotive force to a circuit, usually a coil surrounding a wire. |  |  |  | G02¦G02012¦G0201203
|  | Resistor | G0201204 | An electronic component that opposes the flow of current. |  |  |  | G02¦G02012¦G0201204
|  | Solenoid | G0201205 | An electronic component consisting of a coil surrounding a movable iron core that is designed to act as a switch or relay. |  |  |  | G02¦G02012¦G0201205
|  | Transducer | G0201206 | An electrical component that converts one form of energy into another. |  |  |  | G02¦G02012¦G0201206
|  | Semiconductor | G0201207 | A type of electronic component, including transistors and diodes, that make use of the variable conductivity of certain materials. |  |  |  | G02¦G02012¦G0201207
|  | Feedthrough Capacitor | G0201208 | A component used to carry a current from an enclosing. Ensures that hermeticity is maintained. |  |  | Term added for Release Number 2025. See no. 29 of Change Request Log | G02¦G02012¦G0201208
| IC (Integrated Circuit) Chip |  | G02013 | A microelectronic circuit that incorporates many interconnected transistors and other components. |  |  |  | G02¦G02013
| Display |  | G02014 | A component designed to present information visually. In the case of Touchscreen, use G0204004. |  |  |  | G02¦G02014
|  | Indicator | G0201401 | A component designed to show an operating condition of a system or to attract attention. |  |  |  | G02¦G02014¦G0201401
|  | Screen | G0201402 | A panel or area on an electronic device where images and data are displayed. |  |  |  | G02¦G02014¦G0201402
| Electrical Lead/Wire |  | G02015 | A coated or uncoated wire used to connect two locations electronically. Not to be used for patient connection. |  |  |  | G02¦G02015
|  | Electrode | G0201501 | A small piece of metal or other conductive substance this is used to take an electric current to or from a source of power, a piece of equipment. Not to be used for patient connection. |  |  | Editorial change to the definition on 16 February 2021. | G02¦G02015¦G0201501
|  | Ground Strap/Wire | G0201502 | A wire cable or strap designed to carry current safely away from an electronic device under fault condition |  |  |  | G02¦G02015¦G0201502
|  | Wiring Harness | G0201503 | A collection of grouped wires or cables designed to connect to a specific device. |  |  |  | G02¦G02015¦G0201503
| Electrical Mixer |  | G02016 | An electronic component designed to blend signals. |  |  |  | G02¦G02016
| Electrical Port |  | G02017 | An electronic circuit that acts as a connection to another device or component. |  |  |  | G02¦G02017
| Emitter |  | G02018 | The electron source electrode in a transistor or any source in a system. |  |  |  | G02¦G02018
| Headphone/Headset |  | G02019 | A component that covers the ear through which you can listen to without other people hearing. Some headphones come with a microphone which is used to communicate with other people or the device. |  |  |  | G02¦G02019
| Heater |  | G02020 | A piece of equipment used to raise the temperature of air, gas or water or an object. |  |  | Editorial change to the definition on 16 February 2021. | G02¦G02020
| Hub |  | G02021 | An electronic component designed as a central connection for other devices or components. |  |  |  | G02¦G02021
| Inverter |  | G02022 | An electrical component that converts direct current to alternating current. |  |  |  | G02¦G02022
| Magnet |  | G02023 | A component that attracts iron and produces a magnetic field. |  |  |  | G02¦G02023
| Oscillator |  | G02024 | An electronic component designed to produce a wave signal. |  |  |  | G02¦G02024
| Patient Lead |  | G02025 | An insulated electrical cable designed to connect to an electrical device to a patient. |  |  |  | G02¦G02025
|  | Lead Conductor | G0202501 | A cable designed to conduct electricity from the device to the lead. |  |  |  | G02¦G02025¦G0202501
|  | Patient Electrode | G0202502 | An electrical conductor that is designed to make contact with a patient including defibrillator paddles. |  |  |  | G02¦G02025¦G0202502
| Power Cord |  | G02026 | A flexible cable designed to connect an electrical device to a power outlet. |  |  |  | G02¦G02026
| Power Supply |  | G02027 | Component designed to supply electrical power to devices. |  |  |  | G02¦G02027
| Pressure Transducer Probe |  | G02028 | A probe component designed to convert a change in pressure into a varying electrical signal. |  |  |  | G02¦G02028
| Printer |  | G02029 | An electronic component that is designed to transfer text or images to paper or other substrate. |  |  |  | G02¦G02029
| Receiver |  | G02030 | An electronic component designed to capture an incoming electromagnetic signal and convert it to an audible or visual signal. |  |  |  | G02¦G02030
| Receiver Stimulator Unit |  | G02031 | As part of a cochlear implant, an implanted component designed to receive signal from the external component and then decode the signal and transmit it to the brain. |  |  |  | G02¦G02031
| Scanner |  | G02032 | An electronic component that generates a digital representation of an image for data input to a computer; or a receiver designed to search for a signal within a specified frequency range. |  |  |  | G02¦G02032
| Speaker/Sounder |  | G02033 | A component designed to convert electrical signals to sounds that can be heard. |  |  |  | G02¦G02033
| Switch/Relay |  | G02034 | A mechanical or electronic component designed to break or change the connections in a circuit (e.g. button). |  |  |  | G02¦G02034
|  | Power Switch | G0203401 | A switch designed to regulate the power to a device. |  |  |  | G02¦G02034¦G0203401
|  | Relay | G0203402 | A electronic component designed to break or change the connections in a circuit. |  |  |  | G02¦G02034¦G0203402
| Telemetry |  | G02035 | Component designed to transmit and receive data from a remote source using telecommunications methods. |  |  |  | G02¦G02035
| Temperature Compensator |  | G02036 | A component designed to compensate the temperature of one system in response to temperature changes in another system or the environment. |  |  |  | G02¦G02036
| Thermostat |  | G02037 | A component designed to regulate temperature by controlling the starting and stopping of a heating/cooling system. |  |  |  | G02¦G02037
| Transformer |  | G02038 | A component that either steps up or steps down an alternating electrical current to an output that is suitable for another electrical device or component. |  |  |  | G02¦G02038
| Transmitter |  | G02039 | A component to propagate electromagnetic waves. |  |  |  | G02¦G02039
| User Input Device |  | G02040 | A component that uses a movable handle to create two-axis input to a computer. |  |  |  | G02¦G02040
|  | Joystick | G0204001 | A control component that uses a movable handle to create two-axis input to a computer. |  |  |  | G02¦G02040¦G0204001
|  | Keyboard/Keypad | G0204002 | A component consisting of mechanical keys that are pressed to create input to a computer. |  |  |  | G02¦G02040¦G0204002
|  | Microphone | G0204003 | A component designed to convert sound to an electrical signal. |  |  |  | G02¦G02040¦G0204003
|  | Touchscreen | G0204004 | A control component that operates the device following pressing the display on the screen. |  |  |  | G02¦G02040¦G0204004
|  | Solder Joint | G02041 | Joint between electrical components, often to a printed circuit board |  |  | Term added for Release Number 2025. See no. 30 of Change Request Log | G02¦G02041
Measurement |  |  | G03 | Component capable of measuring something in order to obtain a result. |  | Not selectable | Please use more detailed term within the hierarchy | G03
| Analyzer |  | G03001 | Any component designed to perform an analysis. |  |  |  | G03¦G03001
|  | Oxygen Analyzer | G0300101 | A component designed to measure the concentration of oxygen in a gas mixture. |  |  |  | G03¦G03001¦G0300101
| Aperture |  | G03002 | An instrument designed to measure the size of an opening or one used to increase the diameter of an opening |  |  |  | G03¦G03002
| Calibrator |  | G03003 | A standard or reference material used to set the operating parameters of an instrument. |  |  |  | G03¦G03003
| Clock |  | G03004 | A component designed to indicate the time of day. |  |  |  | G03¦G03004
| Counter |  | G03005 | A component designed to keep track of the number of times something happens. |  |  |  | G03¦G03005
| Cuvette |  | G03006 | A clear container designed to interface with an optical sensor in order to obtain an optical measurement of a contained substance. |  |  | Editorial change to the definition on 16 February 2021. | G03¦G03006
| Gauges/Meters |  | G03007 | A component designed to give a visual indication of the condition of a system. |  |  |  | G03¦G03007
|  | Flowmeter | G0300701 | A component designed to measure the flow rate of a fluid. |  |  |  | G03¦G03007¦G0300701
|  | Manometer | G0300702 | A component designed to measure pressure. |  |  |  | G03¦G03007¦G0300702
|  | Thermometer | G0300703 | A component designed to measure temperature. |  |  |  | G03¦G03007¦G0300703
| Marker |  | G03008 | A visual indicator of position, place or route, including radiopaque markers. |  |  |  | G03¦G03008
| Pipette |  | G03009 | A measuring component, traditionally including a graduated tube, designed for the accurate transfer of liquid volumes. |  |  |  | G03¦G03009
| Pointer |  | G03010 | An indicator component designed to show a position on a scale. |  |  |  | G03¦G03010
| Scale |  | G03011 | A component designed for weighing or an indicator component with a graduated sequence of divisions. |  |  |  | G03¦G03011
| Sensor |  | G03012 | A component designed to respond to a stimulus by generating a signal that can be measured or interpreted. |  |  |  | G03¦G03012
|  | Bubble Sensor | G0301201 | A component designed to signal the presence of bubbles in a system. |  |  |  | G03¦G03012¦G0301201
|  | O2 Sensor | G0301202 | A sensor designed to respond to the presence or level of oxygen in a space or environment. |  |  |  | G03¦G03012¦G0301202
|  | Photodetector | G0301203 | A component designed to detect light. |  |  |  | G03¦G03012¦G0301203
|  | Pressure Sensor | G0301204 | A sensor designed to respond to the level of pressure in a space or pressing on a surface. |  |  |  | G03¦G03012¦G0301204
|  | Sensor Probe | G0301205 | A component designed to reach into a location for sensing. Should be used when the problem involves the probe, not the sensor. |  |  |  | G03¦G03012¦G0301205
|  | Temperature Sensor | G0301206 | A sensor designed to respond to the temperature of a space, surface, or environment. |  |  |  | G03¦G03012¦G0301206
| Timer |  | G03013 | A component designed to measure a time interval. |  |  |  | G03¦G03013
| Control Material |  | G03014 | A standard or reference material used for the purposes of comparative evaluations. |  |  | Term was added on 31 January 2023. For details, see comment No. 145 of the Change Log (Release Number 2023). | G03¦G03014
| Swab |  | G03015 | An absorbent pad or piece of material attached to the end of a stick typically used for applying medicines or taking specimens. |  |  | Term was added on 31 January 2023. For details, see comment No. 146 of the Change Log (Release Number 2023). | G03¦G03015
| Tray |  | G03016 | A flat, shallow container with a raised rim. |  |  | Term was added on 31 January 2023. For details, see comment No. 148 of the Change Log (Release Number 2023). | G03¦G03016
Mechanical |  |  | G04 | Component of a device when it is working, often using power from an engine or from electricity. |  | Not selectable | Please use more detailed term within the hierarchy | G04
| Access Port |  | G04001 | A component designed for the introduction or removal of medical equipment or any substance. |  |  |  | G04¦G04001
| Actuator |  | G04002 | A component which moves or controls a mechanism or system. |  |  |  | G04¦G04002
| Adapter (Adaptor) |  | G04003 | A component designed to make different pieces of apparatus compatible. |  |  |  | G04¦G04003
| Air Eliminator |  | G04004 | A component designed to remove air from a space or fluid. |  |  |  | G04¦G04004
| Anchor |  | G04005 | A component designed to fix the device or a portion of the device in place. |  |  |  | G04¦G04005
|  | Fixation Tines | G0400501 | Projections designed to fix a device inside a cavity of the human body |  |  | Term added for Release Number 2024. See no. 114 of Change Log. | G04¦G04005¦G0400501
| Applicator |  | G04006 | A component designed to transfer a substance to a surface. |  |  |  | G04¦G04006
| Automatic Injection System |  | G04007 | A system designed to pump a gas or liquid into a space or environment without outside intervention. |  |  |  | G04¦G04007
| Bag |  | G04008 | A flexible container, often with a single opening. |  |  |  | G04¦G04008
| Ball |  | G04009 | Any component that is spherical in shape. |  |  |  | G04¦G04009
| Balloon |  | G04010 | A sac designed to be inflated and deflated. |  |  |  | G04¦G04010
| Bearings |  | G04011 | A point of contact between moving parts designed for support or to reduce friction. |  |  |  | G04¦G04011
| Bellows |  | G04012 | A component that expands and contracts to draw in air through a valve or orifice and expels it through a tube. |  |  |  | G04¦G04012
| Belt |  | G04013 | A component consisting of a narrow band of material moving over shafts or pulleys. |  |  |  | G04¦G04013
| Bottle |  | G04014 | A rigid or semi-rigid container used to store liquid. |  |  |  | G04¦G04014
| Breathing Circuit |  | G04015 | In an anesthesia machine, the pathway by which a gaseous anesthetic agent and oxygen are delivered to the patient and carbon dioxide is removed. The circuit can be open or closed. |  |  |  | G04¦G04015
| Brush |  | G04016 | A component consisting of hairs or bristles set into a handle or holder. |  |  |  | G04¦G04016
| Bushing |  | G04017 | A cylindrical metal sleeve designed to reduce the friction of a rotating shaft. |  |  |  | G04¦G04017
| Cable, Mechanical/Structural |  | G04018 | A long, thin, multistranded rope or metallic wire to hold the subject. |  |  |  | G04¦G04018
| Cannula |  | G04019 | A rigid or semi-rigid tube inserted into the body. |  |  |  | G04¦G04019
|  | Cannula Hub | G0401901 | A metal or plastic component that connects to the cannula. |  |  |  | G04¦G04019¦G0401901
| Cap |  | G04020 | A component designed to close an opening of a container or device. |  |  |  | G04¦G04020
| Carrier |  | G04021 | A component designed to facilitate the support, movement, or transport of another device or object. |  |  |  | G04¦G04021
| Caster |  | G04022 | A pivoting roller or wheel designed to attach to an object to make it movable. |  |  |  | G04¦G04022
| Catheter |  | G04023 | A flexible tube inserted into the body designed to permit injection or withdrawal of fluids or to keep passage open. |  |  |  | G04¦G04023
|  | Catheter Hub | G0402301 | A small metal or plastic component that connects to the catheter. |  |  |  | G04¦G04023¦G0402301
| Cell |  | G04024 | A component that is designed as a container to collect and/or transfer materials, reagents or specimens. |  |  |  | G04¦G04024
| Chain |  | G04025 | Assembly of interconnected links, typically made of metal used for connecting other components. |  |  |  | G04¦G04025
| Chamber |  | G04026 | A component designed as a reservoir/storage. |  |  |  | G04¦G04026
| Chassis/Frame |  | G04027 | A supporting frame designed to hold other components or devices such as the internal frame of an electronic device. |  |  |  | G04¦G04027
| Clutch |  | G04028 | A component that engages and disengages power transmission. |  |  |  | G04¦G04028
| Coating Material |  | G04029 | A layer of material covering the surface of a device. |  |  |  | G04¦G04029
| Coil |  | G04030 | A structure consisting of something wrapped in a continuous series of loops. |  |  |  | G04¦G04030
|  | Fixation Helix | G0403001 | A coil designed for fixing or anchoring a device in human tissue. |  |  | Term and Definition modified for Release Number 2024. Previous term was Helifix Coil. See no. 113 of Change Log. | G04¦G04030¦G0403001
| Collimator |  | G04031 | A diaphragm or system of diaphragms made of an absorbing material, designed to define and restrict the dimensions and direction of a beam of radiation. |  |  |  | G04¦G04031
| Concentrator |  | G04032 | A component designed to increase the weight per unit volume of a substance. |  |  |  | G04¦G04032
| Cone |  | G04033 | A three-dimensional part that tapers smoothly from a flat circular shape at one end to a point at the other end. |  |  |  | G04¦G04033
| Connector/Coupler |  | G04034 | A component designed to serve as a link between parts allowing easy disconnection and reconnection when necessary. |  |  |  | G04¦G04034
|  | Connector Pin | G0403401 | A projecting part of a device that allows it to be secured through an opening. |  |  |  | G04¦G04034¦G0403401
|  | Male | G0403402 | A connector with a male design |  | New | New term added for Release Number 2026. See no. 168 of Change Log. | G04¦G04034¦G0403402
|  | Female | G0403403 | A connector with a female design. |  | New | New term added for Release Number 2026. See no. 169 of Change Log. | G04¦G04034¦G0403403
| Controller |  | G04035 | A component designed to control or regulate the operation of another device. |  |  |  | G04¦G04035
| Compressor |  | G04036 | A component that increases the pressure of air or gas. |  |  |  | G04¦G04036
| Cover |  | G04037 | An object designed to conceal, enclose, or protect something. |  |  |  | G04¦G04037
| Cuff |  | G04038 | A bandlike structure that encircles a body part or another component or device. |  |  |  | G04¦G04038
| Cup |  | G04039 | Part or design of a device with a concave hemispherical shape. |  |  |  | G04¦G04039
| Cusp/Leaflet |  | G04040 | A thin blade-like component, typically used as part of a one-way valve. |  |  | Minor editorial change to the term on 25 June 2020. | G04¦G04040
| Cutter/Blade |  | G04041 | A component designed for slicing or cutting. |  |  |  | G04¦G04041
| Cylinder |  | G04042 | A three-dimensional part with flat circular ends and long straight sides. |  |  |  | G04¦G04042
| Device Collapser |  | G04043 | A component designed to fold or collapse something. |  |  |  | G04¦G04043
| Device Deployer |  | G04044 | A component designed to install something or distribute something in a systematic way. |  |  |  | G04¦G04044
| Diaphragm |  | G04045 | A component consisting of a flexible sheet or partition. |  |  |  | G04¦G04045
| Dome |  | G04046 | Part or design of a device with a convex hemispherical shape. |  |  |  | G04¦G04046
| Ejector |  | G04047 | A mechanism that pushes a device or component out. |  |  |  | G04¦G04047
| Equipment Pole |  | G04048 | A structural component designed to hang medical equipment. |  |  |  | G04¦G04048
| Extender |  | G04049 | A component designed to lengthen a structure. |  |  |  | G04¦G04049
| Fabric |  | G04050 | Cloth produced from textile fibers. |  |  |  | G04¦G04050
| Fan/Blower |  | G04051 | A component designed to create an air current through the rotation of a planar surface. |  |  |  | G04¦G04051
| Fastener |  | G04052 | A component designed to hold items in place. |  |  |  | G04¦G04052
|  | Adhesive | G0405201 | Any substance that affixes 2 or more surfaces together. This may be supplied separately, or attached to another item such as tape. |  |  |  | G04¦G04052¦G0405201
|  | Bolt | G0405202 | A cylindrical connector element which may have a thread/nut connection to form a fastener. |  |  | Editorial change to the definition on 16 February 2021. | G04¦G04052¦G0405202
|  | Clamp | G0405203 | A component designed to mechanically hold items firmly together. |  |  |  | G04¦G04052¦G0405203
|  | Clip | G0405204 | A small component designed to hold and attach items together. |  |  |  | G04¦G04052¦G0405204
|  | Fixation Wire | G0405205 | A metal strand designed for structural or other purpose. |  |  |  | G04¦G04052¦G0405205
|  | Latch | G0405206 | A fastening component for a mobile part usually consisting of a bar that is retained in a slot. |  |  |  | G04¦G04052¦G0405206
|  | Nail | G0405207 | A pin-shaped fastener with a sharp point at one end and usually flat on the other end. |  |  |  | G04¦G04052¦G0405207
|  | Nut | G0405208 | A threaded fastener designed to engage a bolt. |  |  |  | G04¦G04052¦G0405208
|  | Pin | G0405209 | A small, slender separate component which is designed to secure another object. |  |  |  | G04¦G04052¦G0405209
|  | Prong | G0405210 | A projecting pointed part of a device, usually one designed to attach a device to something else. |  |  |  | G04¦G04052¦G0405210
|  | Retainer | G0405211 | A physical component designed to hold a fastener in place. |  | Modified (technical) | Technical change for Release Notes 2026. See no. 40 on Change Log. | G04¦G04052¦G0405211
|  | Rivet | G0405212 | A fastening component resembling a heavy pin, having a head on one end and designed to have the other end flattened after passing through the parts to be connected. |  |  |  | G04¦G04052¦G0405212
|  | Screw | G0405213 | A fastening component with a tapered threaded shaft and a head designed to engage with a driving tool. |  |  |  | G04¦G04052¦G0405213
|  | Staple | G0405214 | A fastening component consisting of a bent wire designed to pierce and hold two or more surfaces together. |  |  |  | G04¦G04052¦G0405214
|  | Suture Thread | G0405215 | A monofilament or multifilament material used surgically to close a wound, join tissues, or fasten components of a device (e.g. fabric to wire stent frame). |  |  |  | G04¦G04052¦G0405215
|  | Tape for Fixation | G0405216 | A long, thin, flat, flexible material often used for binding or fastening. |  |  |  | G04¦G04052¦G0405216
| Fiber |  | G04053 | Any component made from a long, slender material. |  |  |  | G04¦G04053
| Filter |  | G04054 | A component designed to remove something from whatever passes through it. |  |  |  | G04¦G04054
| Flange |  | G04055 | A protruding edge designed to strengthen or stabilize a device or facilitate its attachment to a surface. |  |  |  | G04¦G04055
| Foil |  | G04056 | A thin, flexible sheet of metal. |  |  |  | G04¦G04056
| Gas Exchanger |  | G04057 | Component that is used to transfer gasses between two or more locations. |  |  |  | G04¦G04057
| Gasket |  | G04058 | A preformed material designed to form a seal between connecting surfaces. |  |  |  | G04¦G04058
| Gears |  | G04059 | A toothed wheel designed to mesh with another toothed object and transmit motion. |  |  |  | G04¦G04059
| Generator |  | G04060 | A component designed to produce electricity, vapor or gas. |  |  |  | G04¦G04060
| Guide |  | G04061 | A component designed to help direct the passage of another object. |  |  |  | G04¦G04061
| Guidewire |  | G04062 | A flexible wire designed to help position medical devices within the body. |  |  |  | G04¦G04062
| Handpiece |  | G04063 | A part of a device designed to be used while held in the hand. |  |  |  | G04¦G04063
| Header |  | G04064 | The connection point between the leads and the generator. |  |  |  | G04¦G04064
| Sewing Ring |  | G04065 | A ring of supportive material designed to provide a stable surface for attachment to surrounding tissues. |  |  |  | G04¦G04065
| Heat Exchanger |  | G04066 | A component designed to transfer heat between fluids and/or gases across a barrier or to the environment. |  |  |  | G04¦G04066
| Hinge |  | G04067 | A component designed to join two objects and allow them to swing relative to one another. |  |  |  | G04¦G04067
| Holder |  | G04068 | A component, such as a bracket, designed to hold another object |  | Modified (technical) | Technical change to definition for Release Notes 2026. See no. 380 on Change Log. | G04¦G04068
| Hose |  | G04069 | A flexible tube designed to carry a fluid or gas. |  |  |  | G04¦G04069
| Housing |  | G04070 | A rigid casing that encloses and protects a piece of equipment. |  |  |  | G04¦G04070
| Humidifier |  | G04071 | A component used to increase moisture in a gas. |  |  |  | G04¦G04071
| Hydraulic System |  | G04072 | A system designed to use fluid pressure to bring about movement. |  |  |  | G04¦G04072
| Impeller |  | G04073 | The rotating component of a centrifugal pump, compressor, or other machine designed to move a fluid by rotation. |  |  |  | G04¦G04073
| Inserter |  | G04074 | A component whose function is to facilitate the insertion of a particular device. |  |  |  | G04¦G04074
| Insulation |  | G04075 | A material designed to reduce the transmission of heat, sound, or electricity. |  |  |  | G04¦G04075
| Isolator |  | G04076 | Any material or structure designed to limit the interaction between two components. |  |  |  | G04¦G04076
| Jaw |  | G04077 | A component designed to use opposing parts to close on and hold an object. |  |  |  | G04¦G04077
| Joint |  | G04078 | A component designed as the junction between objects; it may be flexible or rigid. |  |  |  | G04¦G04078
| Knob |  | G04079 | A rounded lump or ball used for adjusting or controlling. |  |  |  | G04¦G04079
| Label |  | G04080 | Any written, printed, or graphic matter upon a device to identify its nature, ownership, or other characteristics of the device. |  |  |  | G04¦G04080
| Leaflet |  | G04081 | A device consisting of two thin blades hinged in the center; typically designed to control flow of fluids. |  | Retired (2020) | Term was retired on 25 June 2020. Please use "Cusp/Leaflet" (G04040). | G04¦G04081
| Lever |  | G04082 | A rigid bar that rotates around a fixed point. |  |  |  | G04¦G04082
|  | Foot Pedal | G0408201 | A lever designed to be operated with the foot. |  |  |  | G04¦G04082¦G0408201
| Liner |  | G04083 | A component placed the inside of the walls of a cavity or container for protection or insulation. |  |  |  | G04¦G04083
| Magazine/Cassette |  | G04084 | A compartment in a device designed to house a consumable material for feeding into a mechanism. |  |  |  | G04¦G04084
| Manifold |  | G04085 | A branched system of pipes or channels; typically to funnel the flow of liquids or gases. |  |  |  | G04¦G04085
| Mask |  | G04086 | A flexible, form-shaped component designed to be placed over the nose and/or mouth. |  |  |  | G04¦G04086
| Mechanical Mixer |  | G04087 | A mechanical component designed to blend materials. |  |  |  | G04¦G04087
| Membrane |  | G04088 | A component that is made from or resembles a thin flexible sheet of material acting as a boundary or separating two chambers. |  |  |  | G04¦G04088
| Mesh |  | G04089 | Component made of overlapping strands forming a fine net-like structure. |  |  |  | G04¦G04089
| Motor(s) |  | G04090 | A machine that converts any form of energy to produce or impart motion (kinetic energy). |  |  |  | G04¦G04090
| Mount |  | G04091 | A structural component designed to facilitate the attachment of one object to another. |  |  |  | G04¦G04091
| Needle |  | G04092 | A component with a long, slender, pointed shape. |  |  |  | G04¦G04092
| Nozzle |  | G04093 | A component designed to regulate and direct the flow of a fluid or gas. |  |  |  | G04¦G04093
| Packaging |  | G04094 | The outer wrapping around a device which serves to contain, identify, and protect it prior to use. |  |  |  | G04¦G04094
| Pad |  | G04095 | Typically a soft, cushion-like material used to prevent injury or damage to a person or device. Can also be made of an absorbent material to absorb liquid. |  |  | Editorial change to the definition on 16 February 2021. | G04¦G04095
| Panel |  | G04096 | A rigid sheet that forms a surface of a device or component. |  |  |  | G04¦G04096
| Plate |  | G04097 | A thin, flat sheet or strip used to join, strengthen or to form parts of another structure. For a flat, shallow container with a raised rim, use G03016. |  | Modified (editorial) | Definition was modified on 31 January 2023 (Old definition: A thin, flat sheet or strip used to join, strengthen or to form parts of another structure). For details, see comment No. 148 of the Change Log (Release Number 2023). Minor editorial change to definition for Release Number 2026. See no. 384 of Change Request Log. | G04¦G04097
| Plug |  | G04098 | A component designed to seat into an opening in a device or other object. |  |  |  | G04¦G04098
| Plunger |  | G04099 | A component of a machine, tool or device that pushes or thrusts another object, liquid or gas. |  |  |  | G04¦G04099
| Post |  | G04100 | An upright piece that is fixed firmly that supports, place or aligns other parts. |  |  |  | G04¦G04100
| Potting |  | G04101 | The resin, glue or other substance used to seal other components together. |  |  |  | G04¦G04101
| Probe |  | G04102 | A component designed to reach into a location for manipulating or for measuring. |  |  |  | G04¦G04102
| Processor |  | G04103 | A component designed for preparing or treating a material. Should not be used for computer processors. |  |  |  | G04¦G04103
| Pulley |  | G04104 | A component which changes the direction of a force (e.g. belt or chain). |  |  |  | G04¦G04104
| Pump |  | G04105 | A component designed to facilitate the movement of a fluid. |  |  |  | G04¦G04105
| Pusher |  | G04106 | A component designed to advance something by pushing it. |  |  |  | G04¦G04106
| Rachet |  | G04107 | A part that allows or forces another part's movement in a single direction. |  |  |  | G04¦G04107
| Rail |  | G04108 | A bar designed for support, attachment, guidance, or protection from falling. |  |  |  | G04¦G04108
|  | Side Rail | G0410801 | A supportive or protective rail attached to the side of something. |  |  |  | G04¦G04108¦G0410801
| Regulator |  | G04109 | A component designed to control a process or condition. |  |  |  | G04¦G04109
| Reservoir |  | G04110 | A vessel designed to store a fluid. |  |  |  | G04¦G04110
| Ring |  | G04111 | A circular band-shaped component. |  |  |  | G04¦G04111
| Rod/Shaft |  | G04112 | A long cylindrical bar used to transmit motion or connect other components. |  |  |  | G04¦G04112
| Seal |  | G04113 | A component designed to prevent passage of material through a joint or opening. |  |  |  | G04¦G04113
| Shock Absorber |  | G04114 | A mechanical component designed to dampen or attenuate a force. |  |  |  | G04¦G04114
| Sleeve |  | G04115 | A cylindrical fitting that slides over another part of a device or other object. |  |  |  | G04¦G04115
| Slide |  | G04116 | A flat rectangular piece of glass on which specimens can be mounted for microscopic study. |  |  |  | G04¦G04116
| Socket |  | G04117 | A component designed as an opening into which something else fits. |  |  |  | G04¦G04117
| Spacer |  | G04118 | A component designed to position objects further apart. |  |  |  | G04¦G04118
| Spring |  | G04119 | An elastic component designed to bend under a load and then return to its shape when unloaded. |  |  |  | G04¦G04119
| Stand |  | G04120 | A support component designed to hold an object. |  |  |  | G04¦G04120
| Steering Wire |  | G04121 | A wire designed to enable a device to be maneuvered. |  |  |  | G04¦G04121
| Stent |  | G04122 | Tubular support placed inside a blood vessel, canal, or duct to aid healing or relieve an obstruction. |  |  |  | G04¦G04122
| Stopcock |  | G04123 | A rotating component designed to act like a tap for regulating the flow of a fluid, and to completely stop the flow when closed fully. |  |  |  | G04¦G04123
| Stopper |  | G04124 | Component designed to close an opening. |  |  |  | G04¦G04124
| Strain Relief |  | G04125 | A structure designed to function with a connector to prevent damage to a hose or cable from excess flexing. |  |  |  | G04¦G04125
| Stylet |  | G04126 | A thin metal wire designed to be passed through a needle, catheter, or cannula to stiffen it or clear it of debris. |  |  |  | G04¦G04126
| Syringe |  | G04127 | A component designed as a rigid cylinder with a plunger at one end and a delivery opening at the other. |  |  |  | G04¦G04127
| Table |  | G04128 | A component having a smooth flat surface that is usually supported by one or more vertical legs. |  |  |  | G04¦G04128
| Tip |  | G04129 | Pointed or rounded end of an object. |  |  |  | G04¦G04129
| Tool |  | G04130 | A component that is delivered with a device to support its assembly or operation, and is not part of the device. Use only in the case there are no other terms to describe the component. |  |  | Editorial change to the definition on 16 February 2021. | G04¦G04130
| Translational Motion Component |  | G04131 | A rotating part which is intended to transfer rotational movement or motion to another type of movement (e.g. excenters). |  |  |  | G04¦G04131
| Trap |  | G04132 | A component designed to capture or remove bubbles or fluid. |  |  |  | G04¦G04132
| Trocar |  | G04133 | A sharp, pointed rod designed to pierce the wall of a body cavity to withdraw fluid or allow placement of a catheter or other device. |  |  |  | G04¦G04133
| Tube |  | G04134 | A long hollow cylinder, either rigid or flexible, for holding or transporting liquids or gasses. |  |  |  | G04¦G04134
|  | Capillary Tube | G0413401 | A narrow tube in which a liquid flows up against gravity. |  |  |  | G04¦G04134¦G0413401
| Valve(s) |  | G04135 | A mechanical component designed to control the flow of a fluid or gas. |  |  |  | G04¦G04135
|  | Control Valve | G0413501 | A valve designed to regulate the flow of a fluid or gas. |  |  |  | G04¦G04135¦G0413501
|  | Luer Valve | G0413502 | A valve that incorporates a Luer fitting. |  |  |  | G04¦G04135¦G0413502
|  | One-Way Valve | G0413503 | A valve designed to allow flow in only one direction. |  |  |  | G04¦G04135¦G0413503
|  | Control Clamp | G0413504 | A clamp mechanism designed to occlude flow within a medical device system. This includes roller clamps, slide clamps, and pinch clamps. |  | New | Term added for Release Number 2026. See no. 172-174 of Change Log. | G04¦G04135¦G0413504
| Vaporiser |  | G04136 | A component for gasifying liquids such as drugs. |  |  |  | G04¦G04136
| Vibrator |  | G04137 | A mechanical component designed to create a vibratory motion. |  |  |  | G04¦G04137
| Washer |  | G04138 | A flattened disk used as a mechanical seal between objects. |  |  |  | G04¦G04138
| Weld |  | G04139 | Any joining connection that is the result of welding 2 or more parts. |  |  |  | G04¦G04139
| Wheel |  | G04140 | A mechanical component consisting of a spoked, circular rim or solid disk designed to rotate on an axle or shaft. |  |  |  | G04¦G04140
| Window |  | G04141 | A transparent panel in a device designed for visual inspection or to let light pass. |  |  |  | G04¦G04141
| Shutter |  | G04142 | Aperture that controls or blocks light or radiation passing through. |  |  |  | G04¦G04142
| Spike |  | G04143 | A non-surgical component designed to access or pierce a container (e.g., IV fluid bag or bottle) to allow fluid transfer through an administration set. |  | New | New term added for Release Number 2026. See no. 171 of Change Log. | G04¦G04143
Optical |  |  | G05 | Component which involves or relates to vision, light, or images. |  | Not selectable | Please use more detailed term within the hierarchy | G05
| Camera |  | G05001 | An apparatus for taking photographs, generally consisting of a lightproof enclosure having an aperture with a shuttered lens through which the image of an object is focused and recorded. |  |  |  | G05¦G05001
| Film |  | G05002 | A photographic material designed to capture an image upon exposure to electromagnetic radiation. |  |  |  | G05¦G05002
| Imager |  | G05003 | A component designed to create or record a visual representation. |  |  |  | G05¦G05003
| Laser |  | G05004 | A component designed to emit a monochromatic beam of coherent light. |  |  |  | G05¦G05004
| LED (Light Emitting Diode) |  | G05005 | A type of diode designed to emit light when a current passes through it. |  |  |  | G05¦G05005
| Lenses |  | G05006 | An electric or optical component designed to focus (concentrate) or disperse electromagnetic radiation. |  |  |  | G05¦G05006
| Light Source |  | G05007 | A component that produces visible light. |  |  |  | G05¦G05007
|  | Bulb | G0500701 | A component designed to produce light or heat. |  |  |  | G05¦G05007¦G0500701
| Mirror |  | G05008 | A component consisting of a polished surface designed to reflect light. |  |  |  | G05¦G05008
| Optical Fiber |  | G05009 | A component made with thin glass fibers as a conduit for transmission of light. |  |  |  | G05¦G05009
Safety |  |  | G06 | Safety related component |  | Not selectable | Please use more detailed term within the hierarchy | G06
| Alarm |  | G06001 | A component designed to signal the occurrence of a particular event. |  |  |  | G06¦G06001
|  | Alarm, Audible | G0600101 | A component designed to signal the occurrence of a particular event by making a sound. |  |  |  | G06¦G06001¦G0600101
|  | Alarm, Visual | G0600102 | A component designed to signal the occurrence of a particular event in a way that can be seen. |  |  | Editorial change to the definition on 16 February 2021. | G06¦G06001¦G0600102
| Emergency Button Or Switch |  | G06002 | A button and circuits designed to force the shutdown of a machine or device. |  |  |  | G06¦G06002
| Fail-Safe System |  | G06003 | A component designed to prevent malfunction, unsafe, or unauthorized operation of a device or system. |  |  |  | G06¦G06003
| Locking Mechanism |  | G06004 | A fastening component designed to hold, close, or secure. |  |  |  | G06¦G06004
| Protector/Shield |  | G06005 | A component designed to prevent harm or protect against damage to other components. |  |  |  | G06¦G06005
| Safety Interlock |  | G06006 | A mechanical or electronic component designed to prevent undesired actions due to the changing state of a device, typically to prevent harm to an operator or damage to the device itself. |  |  |  | G06¦G06006
|  | Needle Stick Prevention Mechanism | G0600601 | A mechanism integrated into a device to prevent needle stick injuries. |  |  | Editorial change to the definition on 16 February 2021. | G06¦G06006¦G0600601
| Safety Valve |  | G06007 | A valve designed to automatically open in order to maintain the pressure in a system below a specified pressure. |  |  |  | G06¦G06007
Other Terms/Codes Related to Components |  |  | G07 | Others. |  | Not selectable | Please use more detailed term within the hierarchy. NOTE: Definition changed to update a working note on 16 Feb. 2021 Editorial change to term for Release Number 2025. See no. 123 of Change Log. | G07
| Part/Component/Sub-Assembly Term not Applicable |  | G07001 | The device does not have distinct parts, components, or sub-assemblies, or it would not be appropriate to link the reported incident to a single part, component, or sub-assembly. Use this term if the problem involves or affects the overall device rather than a specific component. |  | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 315 of Change Log. | G07¦G07001
| Appropriate Component Term/Code not Available |  | G07002 | The parts, components, or sub-assemblies are not adequately described by any other term. Note: this code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. |  | Modified (editorial) | Editorial change to term for Release Number 2026. See no. 316 of Change Log. | G07¦G07002
| Insufficient Component Information |  | G07003 | There is not enough information available to classify the medical device component or there is not enough information available to determine whether a medical device component is affected. |  | Modified (technical) | Term was added on 27 January 2022. For details, see comment No. 192 of the Change Log (Release Number 2022). Editorial change to term for Release Number 2025. See no. 123 of Change Log. Technical modification to definition for Release Number 2026. See no. 167 of Change Log. | G07¦G07003