## Slide 1
Day 1 IMDRF Industry Joint Workshop | 14 September 2026

## Slide 2: TGA Adoption of PCCPs
TGA is a member of the IMDRF SaMD WG, and our intent is to adopt an approach consisted with the IMDRF guidance into our pathways
A PCCP describes a plan, proposed by a manufacturer, that states:
  1. the specific planned changes to the medical device software,
  2. the change plan/protocol for implementing and controlling those changes with predefined acceptance criteria/pre-specified performance criteria; and
  3. the assessment of impacts from those changes.
Changes included in a PCCP are limited to changes within the medical device software’s original intended use or intended purpose.
PCCPs are most suitable for predictable, well-characterised modifications. Broad, open-ended statements (e.g. “future improvements”) or reactive/unanticipated changes that are not explicitly described in the PCCP remain outside scope and should be managed via the appropriate regulatory pathway.

## Slide 3: Incorporating PCCP into TGA approval pathways
Two pathways:
1. Conformity Assessment (CA) pathway
The TGA receives an application for a CA certificate that includes a proposed PCCP. The PCCP is assessed within the QMS change control context
2. Comparable Overseas Regulator (COR) reliance pathway
The TGA receives an application for inclusion into the Australian Register of Therapeutic Goods supported by overseas market authorisation where a PCCP has been approved by the COR

## Slide 4: Scope - what changes can be covered by a PCCP?
In-scope changes:
Model re-training / model update using additional, newer, or real-world data to maintain or improve performance within the same intended purpose
New scanner / hardware integration where the intended purpose is unchanged and acceptance criteria confirm performance is maintained.
Including an excessive number of modifications or highly complex modifications within a proposed PCCP may reduce confidence and the PCCP may not be accepted.
Out-of-scope changes:
Changes to intended purpose
Changes to safety-critical attributes
Non-software changes (e.g. material changes) – this guidance is limited to software
Changes triggered by field safety, cyber security events or other unanticipated issues that are not explicitly pre-determined within the PCCP (these should be managed under standard change control and regulatory pathways.)

## Slide 5: Draft TGA guidance
We have drafted guidance on TGA acceptance of PCCPs, aligned to the IMDRF guidance for consultation.
Primary considerations:
Scope – aligned with IMDRF guidance, limited to medical device software
Evidence requirements to support a PCCP
Maintaining and updating a PCCP
Post-market responsibilities
Notification/Evidence of implementation of changes covered by an approved PCCP