## Slide 1
Day 1 IMDRF Industry Joint Workshop | 14 September 2026

## Slide 2: China’s Relevant Work on Cybersecurity for Medical Devices
National Medical Products Administration (NMPA), China
Notes: Distinguished colleagues, good afternoon. I'm from NMPA, China. Today I will share China's Relevant Work on medical device cybersecurity,

## Slide 3: Contents
Regulations, Guidelines, and Standards for Medical Device Cybersecurity
Medical Device Cybersecurity Requirements Analysis
Medical Device Cybersecurity Architecture, Design Review, and Threat Modeling
Coding Rules for Medical Device Cybersecurity
Medical Device Cybersecurity Testing
Summary
Notes: structured around six key modules: regulatory framework, requirements analysis, architecture & threat modeling, secure coding, security testing, and Summary.

## Slide 4: Part 1:Regulations, Guidelines, and Standards
Technical Review Guidelines for Medical Device Cybersecurity Registration(NMPA), YY/T 1843-2022 (Basic Requirements for Cybersecurity of Medical Electrical Equipment), GB/T 44585.1-2024 (IEC 80001-1:2021 Application of risk management for IT-networks incorporating medical devices – Part 1: Safety, effectiveness and security in the implementation and use of connected medical devices or connected health software)
Emphasize Medical Device Cybersecurity Awareness:
  Conduct systematic and continuous cybersecurity training and assessments for product designers, developers, test engineers, document administrators, and other relevant personnel.
  Read books and literature on medical device cybersecurity.
  Conduct comprehensive measurements to ensure that all medical device developers establish cybersecurity awareness.
Notes: China's medical device cybersecurity governance rests on three core documents:
First, the Technical Review Guidelines for Medical Device Cybersecurity Registration, which mandates cybersecurity documentation as part of the pre-market submission;
Second, the sector standard YY/T 1843-2022, specifying Basic Requirements for Cybersecurity of Medical Electrical Equipment;
Third, the national standard GB/T 44585.1-2024, modified adoption of IEC 80001-1:2021, Application of risk management for IT-networks incorporating medical devices – Part 1: Safety, effectiveness and security in the implementation and use of connected medical devices or connected health software, addressing cybersecurity risk management in healthcare facilities.
We emphasize organization-wide cybersecurity awareness across all roles — from R&D to quality to clinical operation. Cybersecurity is not an add-on; it is a mandatory baseline.

## Slide 5: Part 2: Cybersecurity Requirements Analysis
Security Requirements Classification:
  Security Functional Requirements;
  Security Assurance Requirements.
Key Points for Requirements Analysis:
  Analyze the product's vulnerabilities and potential security threats, and describe security requirements alongside each functional requirement;
  Consider not only what the product should do, but also what the product should not do.
Medical device product requirements engineers should not only consider product functionality from the user's perspective, but also evaluate product vulnerabilities from an attacker's perspective.
Notes: Cybersecurity requirements fall into two categories:
Security functional requirements — what defenses the device must have, e.g., authentication, encryption;
Security assurance requirements — how we verify those defenses are effective.
Here we highlight dual-perspective thinking. Requirements engineers must adopt both the user's view — ensuring clinical usability — and the attacker's view — proactively identifying attack surfaces and vulnerabilities, and defining the device's behavioral boundaries under attack.

## Slide 6: Part 3: Architecture, Design Review, and Threat Modeling
Evaluate the medical device software architecture and design to establish a threat model:
  Software architecture refers to the organization or structure of critical software components;
  Critical software components interoperate with other components through data interfaces.
  Medical Device Cybersecurity Threat Modeling Process:
  Identify Objects;
  Identify Threats;
  Assess Threats;
  Mitigate Threats.
Notes: We recommend threat modeling as a mandatory step in design review, focusing on critical software components and their data interfaces.
Our approach follows a four-step process:
Identify assets and system objects to protect;
Identify potential threats;
Assess threat risks by likelihood and impact;
Mitigate high-priority threats.

## Slide 7: Part 3: Architecture, Design Review, and Threat Modeling
Threat | Attributes | Explanation | Example
Spoofing | Identifiability | Impersonate others or entities | Masquerade as a person, file, etc.
Tampering | Integrity | unauthorized data or code modification | Tamper with patient information
Repudiation | Non-repudiation services | Claim not to have executed an action | Deny the modification
Information Disclosure | Confidentiality | Disclose information to unauthorized entities | Patient information and device logs are exposed
Denial of Service | Availability | Denial of Service | Applications and services are unavailable
Elevation of Privilege | Authorization | Unauthorized access | Elevate a standard user to an administrator
Notes: For threat identification, we use the STRIDE model as a systematic tool:
STRIDE enables us to enumerate threats comprehensively and embed security into the design phase.

## Slide 8: Part 4: Coding Rules for Medical Device Cybersecurity
Input Validation:
  Check, validate, escape, and filter all input data;
  Validate upon initial data receipt.
  Internal Interface Security:
  Validate and check data across internal program interfaces.
  Secure Exception Handling:
  Detect exceptions and safely handle all possible execution paths.
  When certain erroneous behaviors or data are detected, they must be handled appropriately to ensure the safe execution of the program.
  Deny service immediately when necessary, and avoid returning detailed error codes.
Notes: To translate architecture into secure products, we enforce strict secure coding rules across several dimensions:
Input validation — rigorously check all external inputs;
Internal interface validation — ensure data integrity between modules;
Robust exception handling — avoid exposing system internals through error messages;

## Slide 9: Part 4: Coding Rules for Medical Device Cybersecurity
Minimizing Feedback:
  Avoid providing untrusted users with excessive information.
  Authentication routines should disclose as little information as possible prior to authentication.
  Never return a password if it is accepted by the program.
Avoiding Race Conditions:
  Ensure that access to shared resources (such as files or variables) is properly controlled.
  Use atomic operations (operations that either fully succeed or have no effect at all).
  Implement locking mechanisms while avoiding deadlocks.
Secure Invocation of Application Components:
  Databases;
  Reusable components (e.g., Spring Boot, Redis, Kafka);
  Network services (e.g., HTTPS, DNS).
Notes: Minimizing Feedback——Avoid providing untrusted users with excessive information；
Race condition avoidance — Ensure that access to shared resources is properly controlled.;
and Secure Invocation of Application Components
These disciplines minimize vulnerabilities at the source-code level.

## Slide 10: Part 5: Medical Device Cybersecurity Testing
Fuzzing
  A method of finding software faults by providing unexpected inputs and monitoring for abnormal results
  Generate a large volume of malformed data as test cases
  Apply these test cases as inputs to the System Under Test
  Monitor and record any crashes or anomalies caused by the inputs;
  Review the test logs to conduct an in-depth analysis of the root causes of the crashes or anomalies.
Notes: Verification is equally vital. We recommend a combination of five testing methodologies:
Fuzz testing — for discovering unknown crash-related vulnerabilities;

## Slide 11: Part 5: Medical Device Cybersecurity Testing
  Penetration Testing:
  An evaluation method that assesses system security by simulating malicious attacks;
  Test whether a software system is secure from an attacker's perspective;
  Use automated tools or manual methods to simulate attacker inputs, in order to identify security vulnerabilities present in the target system at runtime.
Notes: Penetration testing — simulating real-world attacks on deployed systems;

## Slide 12: Part 5: Medical Device Cybersecurity Testing
  White-box Testing:
  Testers possess full internal knowledge and can freely execute destructive actions without the risk of being blocked;
  Suitable for scenarios with tight schedules, or specific penetration testing environments where data collection is out of scope.
  Black-box Testing:
  Relies on the testers' skills to probe and acquire information about the target system. It typically does not aim to uncover all security vulnerabilities;
  Can test the customer organization's incident response procedures, but cannot determine the efficiency of their security defense plans in detecting specific attacks.
Notes: White-box testing — deep coverage based on internal logic;
Black-box testing — external validation of defense-in-depth.

## Slide 13: Part 5: Medical Device Cybersecurity Testing
Static Code Security Testing:
  The security of source code can be audited through either manual reviews or automated tools;
  Generally, automated tools (such as HP Fortify and FindBugs) are utilized to scan the source code. These tools analyze information such as data flow, control flow, and semantics within the program, and match this information against specialized software security rule sets to identify potential security vulnerabilities in the code.
Notes: Static code analysis — scanning source code for defects early;

## Slide 14: Part 5: Medical Device Cybersecurity Testing
  The three main phases of Penetration Testing are: Information Gathering, Vulnerability Discovery, and Vulnerability Exploitation.
  Define the target;
  Analyze risks and obtain authorization;
  Information gathering;
  Vulnerability probing (manual & automated);
  Vulnerability verification;
  Information analysis;
  Exploit vulnerabilities to acquire data;
  Information compilation;
  Generate the report.

## Slide 15: Part 6: Summary
  Cybersecurity issues are a concern that must be addressed throughout the entire lifecycle of a medical device product, and cybersecurity risk control should be integrated throughout this process.
  Cybersecurity issues in medical devices are a concern that requires the attention of all sectors of society, including regulatory authorities, manufacturers, operators, patients, and network service providers. Addressing these challenges requires the joint efforts of all parties involved.
  As a digital technology, medical device cybersecurity is characterized by a rapid pace of iteration, which brings about corresponding risks. This requires us to continuously update our technologies and measures for cybersecurity risk control.
Notes: To conclude, I want to emphasize three points:
First,....
Second, ....
Third, ...