## Annex F
Annex F. Health Impact
Device (bold): For the purpose of this Annex, a device means a medical device including accessories and components
Wherever appropriate "patient" should be taken to include user, operator or any other person affected by the incident.
LEVEL 1 |  |  | LEVEL 2 |  |  | LEVEL 3
Term | Definition | IMDRF Code | Term | Definition | IMDRF Code | Term | Definition | IMDRF Code
Change in Therapeutic Response | Change in response to treatment or cure of a disorder or disease. | F01 | Therapeutic Response Decreased | A reduction in the desirable and beneficial effects resulting from a medical treatment. | F0101
|  |  | Therapeutic Response Increased | An increase in the desirable and beneficial effects resulting from a medical treatment. | F0102
|  |  | Unexpected Therapeutic Effects | Unanticipated desirable and beneficial effects resulting from a medical treatment. | F0103
Death | The cessation of life. | F02 | Intrauterine Fetal Death | Death in utero; failure of the product of conception to show evidence of respiration, heart beat, or definite movement of a voluntary muscle after expulsion from the uterus, with no possibility of resuscitation. | F0201
Brain Death | The cessation of life due to permanent loss of brain function and brainstem for absence of responsiveness to external stimuli, absence of reflex and apnea. | F03
Delay to Diagnosis | Patient diagnosis was clinically significant delayed as a consequence of device performance. | F04
Delay to Treatment/ Therapy | Patient treatment was delayed as a consequence of device performance. | F05
Disruption of Subsequent Medical Procedure | Situation in which the use of the device impedes or affects a subsequent medical procedure or use of a medicine or device. The time elapsed between the use of the device and the medical procedure is not a factor. It is not necessary for the device to have broken or malfunctioned. | F06
Exacerbation of Existing Condition | Use of the device has lead to worsening of the existing disease or condition. | F07
Hospitalization or Prolonged Hospitalization | Hospitalization or prolonged hospitalization due to the health damage accompanying the use of device. | F08 | Intensive Care | Patient requires admission to or extension of stay in an intensive care unit. | F0801
Fetal Harm | Fetal distress, congenital physical or mental impairment or birth defect, excluding fetal death. | F09
Inadequate/Inappropriate Treatment or Diagnostic Exposure | Patient did not receive intended treatment from a device or received inadequate treatment as a consequence of device performance. This includes inadequate/inappropriate radiation dosage during procedures. | F10 | Absence of Treatment | Patient treatment was absent as a consequence of device performance. | F1001
|  |  | Incompatible Blood Transfusion | Receiving incorrect or incompatible whole blood or a blood component. | F1002
|  |  | Missed Dose | A dose of medicine was not taken or delivered at the prescribed dosing interval. | F1003
|  |  | Underdose | Exposure to, or intake of, an insufficient amount of a substance. | F1004 | Radiation Underdose | A lower than intended dose of radiation to a single field. | F100401
|  |  | Overdose | Exposure to, or intake of, an excessive amount of a substance. | F1005 | Radiation Overdose | A higher than intended cumulative dose of radiation to a single field. | F100501
Minor Injury/ Illness / Impairment | A mild injury, illness or impairment which can be treated with minimal or no intervention. | F11
Serious Injury/ Illness/ Impairment | A severe injury, illness or impairment which requires hospitalization or medical intervention. | F12 | Chronic Disease | Disease that persists over a long period of time, cannot be prevented by vaccines or cured by medication, nor do they just disappear. | F1201
|  |  | Disability | Any physical or mental impairment that interferes with an individual's ability to perform desired activities. | F1202
|  |  | Life Threatening Illness or Injury | Patient suffered an illness or injury which if not treated would be fatal. | F1203
|  |  | Permanent Impairment | Irreversible deterioration of the state of health. | F1204
|  |  | Temporary Impairment | Reversible deterioration of the state of health. | F1205
Misdiagnosis/ Misclassification | An incorrect diagnosis or classification of an illness or other problem. | F13
Prolonged Episode of Care | Clinically significant extended treatment time or procedureal time. Not to be used for prolonged hospitalization or prolonged surgery. | F14
Recognised Device or Procedural Complication | Patient experiencing a complication that is well documented in association with the device or procedure. | F15
Reduction in Life Expectancy | Decrease in lifetime/ lifespan of an individual. Decline in the normal duration or in average life. | F16
Sedation | Patient required sedation or additional sedation. | F17
Rehabilitation | Patient requires treatment to facilitate the process of recovery from injury, illness, or disease. | F18
Surgical Intervention | One or more surgical procedures was required, or an existing procedure changed. | F19 | Additional Surgery | Surgical intervention that was not planned and needed to be performed in addition to other interventions including surgical ones. | F1901
|  |  | Amputation | Patient required surgical removal of limb, extremities or other external body parts. | F1902
|  |  | Device Explantation | Patient required a corrective invasive procedure to remove a device, or parts/fragments of a device. | F1903
|  |  | Device Repositioning | Patient required a corrective invasive procedure in which a device was repositioned. This includes implantable, surgically invasive (penetration through the surface of the body) or invasive (penetration though a natural body orifice or permanent artificial opening, e.g. stoma) devices. | F1904
|  |  | Device Revision or Replacement | Patient required a corrective invasive procedure in which a device was replaced/revised. This includes implantable, surgically invasive (penetration through the surface of the body) or invasive (penetration though a natural body orifice or permanent artificial opening, e.g. stoma) devices. | F1905
|  |  | Modified Surgical Procedure | A change of surgical procedures from the original plan. | F1906
|  |  | More Complex Surgery | Surgery that is more complicated, more comprehensive or more extensive than planned. | F1907
|  |  | Prolonged surgery | Surgery that takes longer than anticipated. | F1908
|  |  | Surgical Procedure Delayed | Surgical procedure unable to be completed when originally scheduled and put off until some time in the future. | F1909
Serious Public Health Threat | Any event type, which results in imminent risk of death, serious injury, or serious illness to more than one individual, that requires prompt remedial action. | F20
Unexpected Deterioration | Unexpected clinical deterioration of patient. | F21
Unexpected Diagnostic Intervention | Diagnostic intervention needed to be completed that was not originally planned. | F22 | Biopsy | Patient required removal of tissue or fluid specimen to establish a diagnosis. | F2201
|  |  | Endoscopic Procedure | Patient required additional or unexpected endoscopic procedure. | F2202
|  |  | Imaging Required | Requirement for imaging that was not originally planned. | F2203
|  |  | IVD Testing | Pateint required additional or unexpected IVD testing. | F2204
Unexpected Medical Intervention | Patient required an unforeseen medical intervention, excluding surgery, which was not on the original treatment plan. | F23 | Additional Device Required | Use of an additional or alternative device require to achieve optimal outcome. | F2301
|  |  | Blood Transfusion | Patient required an infusion of whole blood or a blood component directly into the bloodstream. | F2302
|  |  | Medication Required | Patient required additional medication or additional dose of existing medication. | F2303
|  |  | Prophylactic Treatment | A medication or course of action used to prevent a possible disease or condition from occurring. | F2304
|  |  | Radiation Therapy | Patient required addtional radiation therapy. | F2305
|  |  | Resuscitation | Patient required resuscitation. | F2306
Insufficient Information | There is not yet enough information available to classify the health impact. | F24
Unanticipated Adverse Device Effect | Complication occuring in the patient that had not previously been identified or expected for this device. | F25
No Health Consequences or Impact | No apparent harm ocurred in relation to the adverse event. | F26
No Patient Involvement | No patient involvement when the adverse event occurred (for example happened during set-up or cleaning). | F27
Appropriate Term/Code Not Available | The Health Impact is not adequately described by any other term. Note: this code must not be used unless there is no other feasible code. The preferred term should be documented when submitting an adverse event report. This information will be used to determine if a new term should be added to the code table. | F28